## FDA MAUDE Reports
Model Number | Event Date | Event Type | Patient Outcome | Event Description | Manufacturer Narrative | Brand Name | Type of Device | Device Event Key | MDR Report Key | Event Key | Report Number | Device Sequence Number | Product Code | Report Source | Source Type | Reporter Occupation | Type of Report | Report Date | Date FDA Received | Is This An Adverse Event Report? | Is This A Product Problem Report? | Device Operator | Device MODEL Number | Was Device Available For Evaluation? | Is The Reporter A Health Professional? | Date Manufacturer Received | Was Device Evaluated By Manufacturer? | Is The Device Single Use? | Type of Device Usage | Year Reported
IS2000 A60P8 | 2012-02-13 | Injury | Required InterventionHospitalization | On (b)(6) 2013, intuitive surgical, inc. (isi) received a legal complaint, which indicated that a patient allegedly experienced post-operative complications after a da vinci s surgical procedure in (b)(6) 2012. According to the legal complaint, the patient was admitted to the hospital for medical services and care including diagnostic procedures and medical treatment (details were not provided) on or about (b)(6) 2012. On or about (b)(6) 2012, the patient underwent a da vinci si surgical procedure for endometriosis resection, ovarian cystectomy, and extensive lysis of adhesions including enterolysis. The legal complaint alleges that the patient post-operatively experienced an infection resulting from a bowel perforation, which led to the patient undergoing surgery for sigmoid colectomy on an unspecified date. No medical reports, charts, or operative reports were provided. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. Attempts have been made to gather additional information from the risk management group at the site; however, as of the date of this report no response has been received. If additional information is received a follow up medwatch report will be submitted to the fda. Based on the legal complaint, the reported da vinci s surgical procedure occurred on or about (b)(6) 2012. There were no system logs found for (b)(6) 2012. Review of the site's system logs for (b)(6) 2012 revealed there were no system errors that were generated that would have caused or contributed to the post-surgical complications experienced by the patient. No previous complaint was reported related to this event. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si endometriosis resection and ovary cystectomy surgical procedure. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3347300 | 3318237 | 3214832 | 2955842201303308 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-08-06 | 2013-08-30 | Yes | No | Health Professional | IS2000 A60P8 | No | No | 2013-08-06 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-11-17 | Injury | Required InterventionHospitalization | On (b)(6) 2013, intuitive surgical, inc. (isi) received a demand letter and medical records involving a patient who reportedly suffered from post-operative complications after undergoing a da vinci s hysterectomy procedure on (b)(6) 2011. According to the operative report, the patient's pre-operative diagnosis was pelvic pain, abnormal uterine bleeding, and dyspareunia. The patient's cervix and uterus were removed without complications and the patient tolerated the procedure well. The operative reports do not contain any allegation of a malfunction of a da vinci system, instrument, or accessory. According to the demand letter, the patient started to have problems with pelvic pain and vaginal drainage and she stated she felt as if something came loose from my body on an unspecified date. On (b)(6) 2012, approximately 4 months after the da vinci surgical procedure, the patient was diagnosed with a vaginal cuff perforation, cuff dehiscence, and mild peritonitis. The patient underwent diagnostic laparoscopy and laparoscopic repair of the cuff perforation and cuff dehiscence. Per the (b)(6) 2012 operative report, there was a piece of omentum that was protruding through the vaginal perforation. There was also some purulent discharge found in the posterior cul-de-sac consistent with mild peritonitis. The patient's abdomen was well irrigated to clean the leakage due to the peritonitis. The patient was discharged home on (b)(6) 2012. Approximately 2 weeks after her surgery ((b)(6) 2012), the patient contacted her doctor with reported heaviness in her vagina and requested a refill on her pain medications. No additional information was provided to isi about her current status. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. The operative reports do not contain any allegation of a malfunction of a da vinci system, instrument, or accessory. Attempts have been made to gather additional information from the risk management group at the site; however, as of the date of this report no response has been received. Review of the site's system logs for the reported procedure date of (b)(6) 2011 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. In addition, no previous complaint was reported related to this event. If additional information is received a follow up medwatch report will be submitted to the fda. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3347301 | 3318273 | 3214868 | 2955842201303309 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-08-12 | 2013-08-30 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-08-12 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-11-15 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013, intuitive surgical, inc. (isi) received information including the medical records of a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2012 due to recurrent carcinoma in situ of the cervix, fibroids and menorrhagia. There was no report in the patient's operative report (op) that a malfunction of the da vinci surgical system, instruments or accessories occurred and there was no indication that the patient experienced any intraoperative complications. According to the op report, the patient had a previous right salpingectomy for ectopic pregnancy and there was some scar tissue from the right ovary to the uterus. The patient's uterus was found to be too large to remove vaginally, which required that the uterus/cervix be morcellated and then removed. Inspection of the patient's abdomen found that all of the surgical sites were hemostatic. The patient was awakened and transferred to the recovery room in stable condition. The patient reportedly tolerated the procedure well. The medical records received by isi were not complete. Based on the limited information provided, the medical records showed that the patient underwent a cystogram on (b)(6) 2012. The cystogram found no evidence of a leak or fistulous connection with the urinary bladder and a post void image demonstrated near complete emptying of the bladder without evidence of leakage. On (b)(6) 2012, the patient underwent a nephrostomy catheter tube placement procedure, due to an injury of the distal ureter on the lower right side. According to the patient's nephrostogram report, it was discovered that the patient had a obstruction of the ureter several centimeters below the level of the pelvic brim and about 5 cm from the ureteral vesicle junction. It was stated that the ureter appeared to be totally occluded at this level and there was some extravasation of contrast within this region. An attempt to cross the obstruction was not possible. The surgeon made the decision to leave a 10 french nephrostomy catheter in the upper moiety kidney for external drainage and have the patient return in 7 to 10 days to re-attempt to cross the area. No complications were observed and no significant bleeding was noted. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Isi has contacted risk management group at the site to gather additional information; however, as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: the patient sustained an injury during a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3348150 | 3319004 | 3202125 | 2955842201303288 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-08-01 | 2013-08-30 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-08-01 | Device Not Retu | No | Reuse | 2013
| 2012-08-20 | Injury | Required Intervention | On (b)(6) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci hysterectomy procedure on (b)(6) 2012. The legal complaint alleged that during the surgical procedure, the patient's ureter, vagina, and bladder, and other organs not associated with the procedure were negligently injured. The legal complaint also alleged that the injuries resulted in the necessity of continued follow-up care, surgical intervention, and embarrassing side-effects. The details of the patient's injuries and surgical intervention were not provided. | Based on the documents provided by the patient's attorney, it is unknown at what medical center the reported surgical procedure was performed. The name of the surgeon who performed the surgical procedure is also unknown. If additional information is received, a follow-up mdr will be submitted. This complaint is being reported due to the following conclusion: the legal complaint claimed that the patient sustained injuries to her ureter, vagina, and bladder during a da vinci surgical procedure and required surgical intervention. However, at this time, it is unknown if a malfunction of the da vinci system, an instrument, or an accessory caused or contributed to the patient's injuries. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3348441 | 3319372 | 3215967 | 2955842201303313 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-29 | 2013-08-30 | Yes | No | Health Professional |  | No | No | 2013-07-29 | Device Not Retu | No | Unkown | 2013
IS2000 A51P8 | 2011-05-04 | Injury | Required InterventionHospitalization | On (b)(6) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci s total hysterectomy with bilateral salpingo-oophorectomy procedure on (b)(6) 2011. The patient was admitted with the diagnosis of symptomatic uterine fibroids with anemia secondary to abnormal uterine bleeding. According to the operative report, near the conclusion of the surgical procedure, the patient received indigo carmine and a spill from the bladder was observed. The surgeon then placed 180 ml of methylene blue diluted in normal saline into the bladder and no leakages were observed from the bladder or the ureters inside the pelvis. The bladder was also found to be intact during a cystoscopic examination of the bladder. The surgeon noted that the patient tolerated the procedure well and she was taken to the recovery room in stable condition. No intra-operative complications were documented in the operative report. On (b)(6) 2011, the patient was admitted to the hospital complaining of abdominal pain associated with nausea, vomiting, and diarrhea. The patient was also found to have profound dehydration and low-grade fever. A cat scan of the patient's abdomen revealed moderate left hydronephrosis and left hydroureter with obstruction. No free fluid or free air was observed. An interventional radiologist concluded that the patient had a possible infected hematoma instead of a pelvic abscess which was initially suspected. After receiving iv antibiotics, the patient's condition improved. She became afebrile, her diarrhea was completely resolved, and she did not have any complaints of tenderness of the abdomen. The patient was discharged from the hospital on (b)(6) 2011. The discharge summary stated that the patient was doing clinically well and she was to be discharged on oral antibiotics. No additional information was provided after the date of discharge. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. The operative report does not contain any allegation of a malfunction of a da vinci system, instrument, or accessory. No previous complaint was reported relating to this event. Intuitive surgical inc. (isi) has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2011. No system errors were found to have occurred during the reported da vinci s surgical procedure. This complaint is being reported due to the following conclusion: the patient's medical records indicate that the patient developed an infected hematoma and required iv antibiotic treatment after undergoing a da vinci surgical procedure. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3348475 | 3319380 | 3215975 | 2955842201303316 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-08-05 | 2013-08-30 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2013-08-05 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-09-18 | Injury | Hospitalization | As part of a legal mediation effort on (b)(6) 2013, intuitive surgical, inc. (isi) received information including medical records of a patient who underwent a da vinci si partial nephrectomy procedure on (b)(6) 2012 due to a right renal mass and a possible cystic renal cell. According to the patient's operative (op) report the procedure was complex, and took multiple hours to perform. An intra-operative ultrasound was performed and confirmed that the mass was located deep inside of the kidney, which required extensive reconstructive efforts. According to the op report, there was no reported malfunction of the da vinci surgical system, instruments or accessories occurred. No intra-operative adverse event was reported either. The patient was re-admitted due to liver laceration. The medical charts for this hospitalization were not available, but in the documentation dated (b)(6) 2012, this re-admission was stated. On (b)(6), 2012 visit, the patient was reportedly improving. On (b)(6) 2012, the patient visited the doctor's office and he continued to improve. The patient appeared well with no obvious signs of distress. No medical records since this office visit were provided. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. On (b)(4) 2013, isi contacted the hospital risk manager. The risk manager indicated that due to hipaa regulations, she is unable to provide any additional information regarding the patient or the reported event. Review of the site's system logs for the reported procedure date found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: the patient sustained an injury during a da vinci si partial nephrectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3348432 | 3319394 | 3215989 | 2955842201303314 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-08-05 | 2013-08-30 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-08-05 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-01-18 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(6) 2013, intuitive surgical received information including medical records of a patient who underwent a da vinci si hysterectomy procedure due to abnormal uterine bleeding. According to the patient's operative (op) report there was no allegation that a malfunction of the da vinci surgical system, instruments or accessories occurred. According to the information in the op report, during the surgical procedure, a small arterial bleeder from what appeared to be from an ascending cervical branch was appreciated. There were concerns about further continuation of cautery energy in this area and a clip was applied to control the bleeding. Based on the information as provided in the op report there were no other issues encountered during the surgical procedure. After completion of the surgical procedure the patient's urine was found to be clear. A cystoscope was placed and it was confirmed that there was no trauma to the patient's urinary collecting system, as the bladder mucosa was clear. Indigo carmine from both ureters was observed. The patient reportedly tolerated the procedure well and was taken to the recovery room in good condition. On (b)(6) 2013, the patient returned to the hospital complaining of fever and pelvic pain. A ct scan performed showed 5 by 3 cm collection above the closed vaginal cuff. The ct scan also showed stricture of the right distal ureter with significant dilation of the right urinary collecting system. A pelvic examination found that the vaginal cuff was intact without disruption or infection on the vaginal mucosa. On (b)(6) 2013, the patient underwent a cystoscopy, right retrograde pyelography and a right stent placement. Based on the information indicated in the op report, examination of the patient's bladder found no evidence of stones, tumors, or other lesions. Good efflux of urine from the left ureteral orifice was observed; however, there was no efflux of urine from the right ureteral orifice. Retrograde pyelogram performed showed that the distal ureter was narrow and appeared to be compressed from extrinsic compression, but was patent and contrast did flow to the dilated more proximal portion. A stent was installed with good placement confirmed by fluoroscopy. Clear urine was observed to be effluxing after a retrograde pyelography was performed. The contrast drained briskly. After completion of the procedure, the patient was awakened, extubated and taken to the recovery room in stable condition. The patient reportedly tolerated the procedure well. On (b)(6) 2013, the patient underwent a colpotomy with irrigation of abscess procedure. During the procedure there was no indication that the patient suffered any complications. The procedure was successfully completed and the patient was transferred to the recovery room in good condition. The patient was discharged from the hospital on (b)(6) 2013. On (b)(6) 2013, the patient underwent a laparoscopy, lysis of adhesion and left salpingo-oophorectomy procedure. Examination of the patient's vaginal cuff found no actual mass and the cuff was found to be intact with some loosening of suture. Reportedly, the cuff was slightly open due to a previous drainage procedure of an abscess base. Under laparoscopy, the patient's anatomy appeared normal, with some adhesions of the omentum to the pelvis. Minor adhesions to both the left pelvic sidewall and the right pelvic sidewall were noted and the left ovary was enlarged and had multiple cysts. No purulent material was observed during dissection. A lot of dissection was required to locate the patient's left ovary and dissect it free and care was taken to avoid a bowel injury. The patient had a large adnexa and it was also removed. The surgical area was irrigated with copious amounts of bacitracin. Interceed was placed over the dissected cuff area and the area where the ovary had been freed from the sidewall. The procedure was completed and the patient was taken to the recovery room in good condition. The patient was discharged from the hospital on (b)(6) 2013. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Isi has contacted risk management group at the site to gather additional information; however, as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3348456 | 3319414 | 3216009 | 2955842201303317 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-08-05 | 2013-08-30 | No | Yes | Health Professional | IS3000 A60P8 | No | No | 2013-08-05 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-07-31 | Injury | Required Intervention | It was reported that during a da vinci si myomectomy procedure , a fragment from a cable broke off of the fenestrated bipolar forceps instrument and fell into the patient. The fragment was retrieved and the surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The instrument has been returned to intuitive surgical inc. (isi) for evaluation. However, the evaluation of the instrument has not been completed at this time. Therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted post failure analysis evaluation or if additional information is received. Intuitive surgical inc. (isi) has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. This complaint is being reported due to the following conclusion: the initial reporter claimed that a fragment from the fenestrated bipolar forceps instrument broke off and fell into the patient. However, at this time, it is unknown how the fragment was retrieved. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3349323 | 3320150 | 3216745 | 2955842201303332 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-08-02 | 2013-08-30 | Yes | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-08-02 | No | No | Unkown | 2013
IS3000 A60P6 | 2012-11-08 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical received information, including an encounter report and an operative report of a patient who underwent a da vinci si hysterectomy with bilateral salpingo-oophorectomy (bso) procedure on (b)(6) 2012. This legal allegation indicates that the patient suffered post-operative complications due to the da vinci si system. The encounter report indicates that there were no complications during the surgical procedure. The operative report states that there were no complications during this surgical procedure. There is no information relating to a da vinci surgical system, accessory or instrument malfunction. On (b)(6) 2012, the patient was re-admitted with complaints of low back pain and lower abdominal pain. Ct scan of abdomen/ pelvis showed moderate to severe left -sided hydronephrosis with no evidence of any obstructing stone or mass. On (b)(6) 2012 interventional radiology placed a percutaneous left side ureteral stent was placed for moderate hydronephrosis. The stent was removed in 2-13 after resolution of the obstruction. According to post-operative follow up notes, dated (b)(6) , the patient was recovering well with no pelvic pain and no major complaints. No further clinical information has been provided. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation that was provided to intuitive surgical does not contain any report of a malfunction of a da vinci system, instrument or accessory. Intuitive surgical has made an attempt to contact the site to obtain further clinical information; however, as of date of this report no response has been received. Review of the site's system logs for the reported procedure date of (b)(6) 2012 found no system errors were generated that would have caused or contributed to the post-surgical complications after undergoing a da vinci si procedure. This complaint is being reported due to the following conclusion: the patient experienced post-operative complications after undergoing a da vinci si hysterectomy with bilateral salpingo-oophorectomy (bso) procedure. | DA VINCI SI SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3346599 | 3317557 | 3214152 | 2955842201303274 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-08-05 | 2013-08-29 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-08-05 | Device Not Retu | No | Reuse | 2013
42018404 | 2013-07-30 | Injury | Required Intervention | It was reported that during a da vinci si myomectomy procedure, a piece of the permanent cautery spatula instrument broke off and fell into the patient. The broken instrument piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3346783 | 3317764 | 3214366 | 2955842201303279 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-07-30 | 2013-08-29 | Yes | Yes | Health Professional | 42018404 | No | No | 2013-07-30 | Device Not Retu | No | Unkown | 2013
| 2011-12-12 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical received information, including operative report of a patient who underwent a da vinci laparoscopic hysterectomy on (b)(6) 2011. This legal allegation indicates that the patient suffered post-operative complications after a da vinci hysterectomy procedure. The patient underwent a da vinci procedure on (b)(6) 2011 and it was noted in the operative report that the procedure was uneventful. The patient experienced pain during postoperative course which was well controlled by post-operative day 2. The patient was discharged on (b)(6) 2011 (two days after surgery) in a stable condition. There is no indication that the da vinci system, instrument or accessories malfunctioned based on the operative report. As per the medical records, the patient developed left lower quadrant and left flank pain, vaginal leakage and urinary incontinence two weeks post-operatively and was admitted and evaluated on (b)(6) 2011. A ct urogram on (b)(6) 2011 showed a left distal ureteral injury and a ureterovaginal fistula. The patient was taken to the operating room on (b)(6) 2012 for further evaluation and a cystoscopy showed no evidence of vesicovaginal fistula and right retrograde pyelogram was normal. The left pyelogram indicated a left ureterovaginal fistula. Attempted stent placement was unsuccessful and the patient was placed with a left percutaneous nephrostomy tube for percutaneous drainage. This patient was evaluated at a planned 3 month follow up (on (b)(6) 2012) for a further diagnostic evaluation. There was no evidence of vesicovaginal fistula on cystogram. Based on the operative findings the distal left ureteral stricture was deemed appropriate for ureteral reimplantation. The patient tolerated the procedure well and was transferred to a recovery room in stable condition with no noted complications. No further clinical information has been provided about the patient's current condition. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation that was provided to intuitive surgical does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. This complaint is being reported due to the following conclusion: this legal complaint alleges that the patient suffered post-operative complications after a da vinci laparoscopic hysterectomy procedure. | DA VINCI SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3346806 | 3317784 | 3214399 | 2955842201303281 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-08-05 | 2013-08-29 | Yes | No | Health Professional |  | No | No |  | Device Not Retu |  | Invalid Data | 2013
| 2013-07-03 | Death |  | It was reported that post-op a da vinci myomectomy procedure, the patient coded. The site attempted to resuscitate the patient; however, the attempts were unsuccessful and the patient expired. No other information was provided. On (b)(6) 2013, the intuitive surgical, inc. (isi) clinical sales representative who was made aware of this incident, indicated that she was at a gathering when this information was told to her by a scrub technician who was asked to assist with resuscitation of the patient. Reportedly, the da vinci surgical system was undocked from the patient when the patient began to code. The patient had a pulmonary embolism. Reportedly, the patient's demise was unrelated to the da vinci surgical system. On (b)(6) 2013, isi contacted the site to gather additional information. The site's risk manager indicated that they had no record of a patient death related to a da vinci surgical procedure as they do not have a da vinci surgical system and that the initial reporter of the incident was no longer an employee of the hospital. | Based on the limited information provided, intuitive surgical, inc. (isi) was unable to determine the root cause of the patient's demise. If additional information is received a follow up medwatch report will be submitted to the fda. There was no allegation of a malfunction of a da vinci system, instrument or accessory. Isi contacted the site where the reported incident allegedly occurred; however, upon follow up with the site, it was reported that they do not have a da vinci surgical system. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: a patient allegedly expired post a da vinci myomectomy procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3347128 | 3318117 | 3214712 | 2955842201303290 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-07-31 | 2013-08-29 | Yes | No | Service Personnel |  | No | No | 2013-07-31 | Device Not Retu | No | Unkown | 2013
IS3000 A60P8 | 2013-04-10 | Injury | Required InterventionHospitalization | On b4 2013 intuitive surgical inc isi received b4 from the hospital The event details are provided below please note the 5 cases being reported concurrently by our facility have the d vinci robot as a commonality we are not implying that it is the cause Our facility has two d vinci robots We are uncertain which of the two was used in this case On b6 2013 the patient underwent dvinci pelviscopy with excision of endometriosis During procedure it was noticed that patient had adhesions so surgeon was called in to do lysis of abdominal pelvic adhesions The patient was discharged to home b6 2013 On b6 2013 patient was readmitted with abdominal pain Ct of abdomen and pelvis fairly unremarkable The patient was observed in the hospital With no fevers wellcontrolled pain abdomen wound clean and dry Patient was discharged on b6 2013 The patient was readmitted the same day b6 2013 The patient underwent an open laparoscopic exploration with repair of sigmoid colon perforation Patient was discharged on b6 2013 | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The site's uf/importer report does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Isi has contacted the initial reporter of this complaint to gather additional information; however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: the patient sustained a bowel injury during a da vinci si gynecological procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3343491 | 3314289 | 3210884 | 2955842201303248 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2103-07-29 | 2013-08-28 | Yes | No | Health Professional | IS3000 A60P8 | No | Yes | 2103-07-29 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-07-18 | Injury | Required Intervention | It was reported that after completion of the da vinci si hysterectomy procedure, the patient was found to have sustained a bowel injury. According to the initial reporter of this complaint, the patient returned to the hospital several days postop complaining of abdominal pain. An exploratory laparotomy was performed and a small laceration was found on the patient's sigmoid colon. The intuitive surgical inc. (isi) clinical sales representative (csr) assigned to the site contacted the surgeon directly to obtain additional information regarding the reported event. The csr indicated that the surgeon informed him that the anatomy was very distorted due to disease and the surgical procedure was an extremely tough case. | On (b)(4) 2013, isi contacted the initial reporter of this complaint and obtained additional information regarding the reported event. The initial reporter indicated that no malfunction of the da vinci si system, an instrument, or an accessory occurred during the da vinci si hysterectomy procedure. The initial reporter also indicated that there was no malfunction of the da vinci si surgical system that caused or contributed to the patient's injury. The initial reporter did not know the patient's current status. However, she stated that the patient's bowel injury was repaired by another surgeon on an unknown date. On (b)(4) 2013, isi also contacted the csr assigned to the site and obtained additional information regarding the reported event. According to the csr, the surgeon informed him that he used a pk dissecting forceps instrument during the surgical procedure to move the patient's sigmoid colon out of the way so that he could have a better view of the patient's anatomy. The csr stated that there was no allegation that the pk dissecting forceps instrument malfunctioned during the surgical procedure. Isi has attempted to contact the surgeon to obtain additional information regarding the reported event. However, no additional information has been provided by the surgeon as of the date of this report. A follow-up mdr will be submitted if additional information is received. Based on the information provided, this complaint is being reported due to the following conclusion: the patient was found to have a laceration to the sigmoid colon after undergoing a da vinci si hysterectomy procedure. However, at this time, it is unclear how the patient's injury occurred. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3343780 | 3314520 | 3211115 | 2955842201303253 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-07-30 | 2013-08-28 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-30 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-07-23 | Injury | Other | It was reported that during a da vinci s hysterectomy procedure, the surgical staff observed pieces break off the wrist of the permanent cautery hook instrument and fall into the patient. The broken fragments were retrieved with laparoscopic instruments and the surgical procedure was successfully completed. Prior to closure of the patient's vaginal cuff, an x-ray was performed of the patient's abdomen. No broken fragments were found with the x-ray. No patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Based on the information provided, this complaint is being reported due to the following conclusion: an x-ray was performed on the patient's abdomen after fragments from the instrument broke off and fell into the patient. However, the fragments were successfully retrieved laparoscopically and no fragments were found with the x-ray. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3343844 | 3314608 | 3211203 | 2955842201303254 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-07-29 | 2013-08-28 | Yes | Yes | Health Professional | 42018306 | No | No | 2013-07-29 | Device Not Retu | No | Unkown | 2013
|  | Injury | Life Threatening | On (b)(6) 2013, intuitive surgical inc. (isi) received a journal article titled, perioperative outcomes of robotic versus laparoscopic hysterectomy for benign disease. Within the article, the following statement was documented, one patient who underwent robotic hysterectomy sustained a combined bladder and bowel injury diagnosed 14 days postoperatively after presenting with fever and pelvic abscess. No additional information was provided regarding the reported event. | Based on the information provided, intuitive surgical inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. The journal article does not contain any allegation of a malfunction of a da vinci system, instrument, or accessory. On (b)(6) 2013, intuitive surgical inc. (isi) contacted a physician who co-authored the journal article to obtain additional information regarding the reported event. The physician indicated that all surgical procedures reviewed in the journal article occurred at the (b)(6). The physician did not provide the surgeon name or procedure date related to the reported event. On (b)(6) 2013, isi contacted the risk management department at (b)(6). The risk manager was unable to provide any additional information regarding the reported event. This complaint is being reported due to the following conclusion: the journal article stated that a patient sustained a combined bladder and bowel injury after undergoing a robotic hysterectomy procedure. However, at this time, it is unknown how the patient's injuries occurred. It is also unknown if a malfunction of a da vinci system, an instrument, or an accessory caused or contributed to the patient's injury. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3343892 | 3314670 | 3211265 | 2955842201303255 | 1 | NAY | Manufacturer | Literature | Physician | Initial | 2013-07-29 | 2013-08-28 | Yes | No | Health Professional |  | No | Yes | 2013-07-29 | Device Not Retu | No | Reuse | 2013
IS2000 A51P7 | 2012-07-31 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013 intuitive surgical, inc. Received information including a medical records summary of a patient that underwent a da vinci s hysterectomy procedure on (b)(6) 2012. It was reported that the patient underwent the surgical procedure due to recurring pelvic pain, uterine tenderness, bilateral adnexal tenderness and an ovarian cyst. The surgeon encountered filmy adhesions in the bladder area while performing dissection. The bladder was adherent to the uterine segment. In addition, there were peritoneal adhesions around the low uterine segment. After the uterus was removed and closure of the patient's vaginal cuff was completed, a small opening into the dome of the patient's bladder was noted. At that time, a urologist was consulted and the patient's bladder was repaired intraoperatively. A cystoscopy post-bladder closure was performed. Dye was observed flowing through both the right and left ureters. The patient was awakened and returned to the recovery room. The patient was discharged from the hospital on (b)(6) 2012. During a routine follow-up examination on (b)(6) 2012, the patient had suprapubic pains and recurrent urinary tract infections. The patient was admitted into the hospital and underwent cystourethroscopy, bilateral retrograde pyelogram, and cystogram. The area of bladder repair was inspected and it appeared that there was some scarring but no abnormalities were seen. However, there was one area of erythema. The surgeon was unsure of the etiology of the patient's discomfort but opined that she might have a component of interstitial cystitis. No other information regarding the patient's care was provided. | Based on the information provided, intuitive surgical has not determined the root cause of the intra-operative injury subsequent post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: during a da vinci s surgical procedure, the patient sustained a bladder injury and experienced post surgical complications. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335680 | 3306639 | 3203234 | 2955842201303194 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS2000 A51P7 | No | Yes | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2010-08-10 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical received information including medical records, of a patient who underwent a da vinci s prostatectomy procedure (b)(6) 2010. The patient was discharged and then readmitted to the icu and diagnosed with sepsis on (b)(6) 2010 and found to have candida glabrata in his blood stream. The patient's condition was managed with micafungin plus zosyn and followed up by the hospital's infectious disease service. The patient was moved to the medical floor on (b)(6) 2010. The patient's blood glucose was low ranging 45 mg/dl to 80 mg/dl for the past 24 hours, managed with d50, bp 194/100 to 152/100 treated with atenolol and lisinopril. The patient also described some shortness of breath during exertion. A ct angiogram of the patient's chest was performed on (b)(6) 2010 and ruled out a pulmonary embolism. On (b)(6) 2010, the physician notes officially read as no evidence of pulmonary edema. The operative report does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. No further clinical information was provided about the patient's current status. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however, as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)64 )2010, found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci s surgical prostatectomy procedure. | DA VINCI S SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335655 | 3306646 | 3203241 | 2955842201303188 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2007-11-07 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on july 25, 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2007. The documentation received states that the patient sustained an injury to her small bowel during the surgical procedure. In addition, the legal complaint indicated that the patient developed a pelvic abscess post-operatively. The patient's medical records state the patient underwent a robotic assisted hysterectomy, bilateral salpingo-oophorectomy, and a bilateral pelvic and paraaortic lymphadenectomy procedure. After closure and inspection of the vaginal cuff, a small "serotomy" was found in the sigmoid colon. The "serotomy" was repaired with sutures. Upon completion of the surgical procedure, the surgeon stated in the operative report that there were no complications and the patient was taken to the recovery room in stable condition. The patient was discharged home the following day with oral pain medications. At an unspecified time right after the surgery was completed, the patient developed abdominal distension from accumulation of lymphatic fluid. After the first week postop, the patient developed copious drainage of lymphatic fluid from the vagina. She was admitted to the hospital on (b)(6) 2007 with reported complaints of a swollen right lower leg and fever. A week prior to admission, the patient had developed flu-like symptoms with body aches, diarrhea, nausea, and vomiting. A ct scan of the patient's abdomen and pelvis revealed a pelvic abscess in the right lower quadrant. On (b)(6) 2007, the patient underwent a ct-guided drainage of the abscess and a trunculose drain was placed. The patient was discharged home on (b)(6) 2007, with oral antibiotics. According to the documentation provided, the drain was removed approximately 2 months later. No further information was provided after the date of discharge. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient developed a pelvic abscess after undergoing a da vinci hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335696 | 3306668 | 3203263 | 2955842201303190 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2010-03-08 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on july 25, 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si supracervical hysterectomy procedure on (b)(6) 2010. No intra-operative complications were documented in the operative report. After the patient's uterus was resected and morcellated, her pelvis was irrigated. The operative report stated that excellent hemostasis was noted. Upon conclusion of the surgical procedure, the patient's skin incisions were closed. She was then extubated and escorted to recovery in excellent condition. Based on the patient's medical records, the patient experienced vaginitis and vaginal infections on multiple occasions between (b)(6) 2010 and (b)(6), 2012. On (b)(6) 2012, the patient had an office visit and was evaluated for hydronephrosis. The office visit report stated that the patient had an mri that revealed left sided hydronephrosis. It is unknown what date the mri was performed. The operative report also stated that the patient has been told that she has had hydronephrosis in the past. On (b)(6) 2012, a ct scan was performed on the patient's abdomen and pelvis. The ct scan revealed severe left-sided hydronephrosis and hydroureter. No mass or calcification was observed. On (b)(6) 2012, the patient underwent a left retrograde pyelogram which revealed significant narrowing in the left ureter and severe hydroureteronephrosis. The surgeon inserted a left ureteral stent. According to the operative report, the patient tolerated the procedure well. She was extubated in the or and sent to the recovery room in good condition. The patient was discharged home that same day. On (b)(6) 2012, the patient had a post-operative visit. It was stated in the office visit report that the patient was experiencing pain or burning during urination. The office visit report also stated that the patient denied observing blood in her urine and she had good size and strength to her urinary system. In addition, the patient felt that she was emptying her bladder well. On (b)(6) 2013, the patient underwent a robotic-assisted laparoscopic left ureteral reimplant procedure with psoas hitch bladder fixation. No complications were reported. The patient was discharged home from the hospital on (b)(6) 2013 after her jp drain was removed and she was able to urinate without difficulty. On (b)(6) 2013, the patient underwent a cystoscopy and stent removal procedure. During the procedure, the patient's bladder was drained. No complications were documented. The patient claimed that she was told that the ureteral stricture was the result of a thermal injury that occurred during the da vinci s supracervical hysterectomy procedure performed in 2010. Based on the patient's medical records, there is no documentation that a thermal injury occurred. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2010. No system errors were found to have occurred during the reported da vinci si surgical procedure. This complaint is being reported due to the following conclusion: the patient sustained a thermal burn to a ureter during a da vinci s surgical procedure. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335707 | 3306679 | 3203274 | 2955842201303193 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-10-18 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2011. The documentation received states that the patient developed a pelvic abscess after undergoing the surgical procedure. The patient's operative report stated that extreme scar tissue and an extremely large fundal fibroid were observed. Due to the size of the uterus, the uterus was cut in half and removed vaginally. According to the operative report, the patient tolerated the procedure well and she was taken to recovery without sequelae. On (b)(6) 2011, the patient was evaluated in an emergency room (er) for reported complaints of abdominal pain, nausea, vomiting, and fever. A ct scan performed on the patient's abdomen and pelvis revealed a pelvic abscess with stool-like material within. On (b)(6) 2011, the patient underwent percutaneous drainage of the pelvic abscess. A total of 15 ml of purulent fluid was removed. In addition, placement of a drainage catheter was performed. According to the patient's medical records, the patient tolerated the procedure well and no immediate complications were reported. A physician noted on (b)(6) 2011, that the patient was feeling well and did not have any fevers or chills since the drainage was performed on the pelvic abscess. The patient's white blood count was normalized and a urine culture showed no growth. The patient denied having dysuria. On (b)(6) 2011, the patient's medical records indicate that the patient was to be discharged home on oral antibiotics. A physician noted that no definitive bowel injury was present. However, the physician noted possible bladder injury and the patient had a prior procedure that did find a retro vesicular abscess with communication with the bladder at the time of drainage. No additional clinical information was provided after the date of (b)(6) 2011. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient developed a pelvic abscess after undergoing a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335715 | 3306684 | 3203279 | 2955842201303191 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS 2000 A51P8 | 2012-08-22 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013 intuitive surgical received information, including medical records, of a patient who underwent a da vinci s hysterectomy procedure on (b)(6) 2012 for excessive cramping and heavy bleeding. The patient had been experiencing vaginal bleeding since the birth of her child in 2001. In 2008, the patient had an ablation of the uterus to help stop the bleeding but the heavy bleeding persisted. The patient opted for a hysterectomy. The patient was informed that because of her history of 3 cesarean sections, her bladder may possibly tear during the surgical procedure if it was adhering to her uterus. On (b)(6) 2012, the patient underwent a da vinci s hysterectomy, lysis of adhesions and cystoscopy procedure. Adhesions were noted of the omentum to the anterior abdominal wall inferior to the umbilicus. The adhesions were lysed. There were also significant adhesions of the bladder to the underlying cervix and lower uterine segment. It was also noted that there was no entrapment of abdominal contents noted at umbilical defect fascial closure. During the procedure, according to the operative report, the surgeon noted a left inguinal hernia as c02 gas filled the left vulva during the procedure. The swelling resolved following surgery. A cystoscopy did not reveal any bladder injury. The patient woke up happy and was discharged at 6pm the same day. On (b)(6) 2012, the patient reported trouble urinating. The following day, the patient went to the emergency room (er) due to swelling to left lower abdomen near one of the surgical sites and labia. A ct scan of her abdomen/pelvis showed extensive air and edema within the subcutaneous tissue consistent with subcutaneous emphysema. On (b)(6) 2012, a cystogram was performed which showed there was a contained leak from the fundus of the urinary bladder. On (b)(6) 2012, the patient underwent a laparotomy procedure to repair the right posterior bladder injury. On (b)(6) 2012, the patient was discharged. A follow up visit on (b)(6) 2012 indicated that the patient was generally recovering well. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2012 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci s surgical hysterectomy procedure. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335727 | 3306694 | 3203289 | 2955842201303195 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS 2000 A51P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
1200 A43P9 | 2013-08-14 | Other | Other | It was reported that during a da vinci standard prostatectomy procedure, the site experienced an error 212 on the master tool manipulator right (mtmr) of the da vinci standard system on (b)(6) 2013. Due to the alleged error message during the da vinci surgical procedure the surgeon made the decision to convert the surgical procedure to open. No further clinical information was provided about the patient's status post- surgery. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). A system error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. The site contacted intuitive surgical inc technical support on (b)(6) 2013 alleging that the site had a recoverable 10025; however, the system did not start. After restarting several times the issue persisted. Isi ordered replacement parts for the alleged error 212; however, the surgeon decided to cancel the order since the site was thinking about trading the da vinci standard. The site has been instructed to contact intuitive surgical technical support if they decided to have the da vinci standard repaired. Intuitive surgical inc. Attempted to contact the site to obtain additional information concerning the reported event; however, no further information has been provided as of date of this report. | Intuitive surgical inc. Attempted to contact the site to obtain additional information concerning the reported event; however, no further information has been provided as of date of this report. If additional information is received a follow-up medwatch report will be submitted to the fda. Review of the site's system logs for the reported procedure date of (b)(6) 2013 confirmed the alleged error 212. This complaint is being reported due to the following conclusion: during a da vinci standard prostatectomy procedure the site experienced an error 212 and due to the error 212 message the surgeon decided to convert the surgical procedure to open | DA VINCI STANDARD SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335728 | 3306697 | 3203292 | 2955842201303192 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-08-14 | 2013-08-26 | Yes | Yes | Health Professional | 1200 A43P9 | No | Yes | 2013-08-14 | Device Not Retu | No | Reuse | 2013
IS3000 | 2011-01-18 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on july 25th, 2013 intuitive surgical received information including medical records of a patient who underwent a da vinci si hysterectomy with bilateral salpingo-oophorectomy on (b)(6) 2011. According to the surgeon's operative report, there were adhesions of the small bowel to the abdominal wall. Lysis of adhesions was performed. The patient returned to the hospital's emergency room on (b)(6) 2011 with complaints of abdominal distention, nausea and vomiting. A ct scan of the patient's pelvis was performed found that the patient had an organized clot. A needle biopsy was performed and no sign of abscess formation or infection was noted. X-rays performed were consistent with ileus. On (b)(6) 2011 a pelvic exam under anesthesia was performed. The cuff was intact; however, upon opening the cuff, fluid consistent with lysing hematoma was discovered and drained. There was no evidence of fistula or foreign body, thermal energy or damage to the patient's bladder. Following the surgical procedure, the patient did experience improvement. Over the course of the patient's hospitalization, the patient had clinically improved and the patient was discharged from the hospital on (b)(6) 2011. The patient's consultation report indicated that she returned to the hospital's emergency room on (b)(6) 2011. A ct scan performed showed possible diverticulitis, urinary incontinence and urinary retention. The patient was treated with antibiotics. Based on the medical documentation provided the patient continued to be monitored and tested for her condition. There was no indication that the patient underwent any additional surgical procedure for her condition. The date that the patient was discharged from the hospital was not provided. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced post surgical complications after undergoing da vinci si hysterectomy bilateral salpingo-oophorectomyprocedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335685 | 3306710 | 3203305 | 2955842201303196 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS3000 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS 3000 A60P8 | 2011-09-21 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013, intuitive surgical, inc. (isi) received information, including medical records, of a patient who underwent a da vinci si sacrocolpopexy procedure on (b)(6) 2011 for vaginal prolapse. The operative notes indicate that the patient's uterus and right adnexa were absent and the left adnexa and left sigmoid colon were scarred together. The use of monopolar scissors and pk instruments were mentioned in the procedure notes, but there was no report of a da vinci si system or instrument malfunction. The patient tolerated the procedure well and was taken to the recovery room in good condition. According to the medical record, the patient did not have any complications with the da vinci si surgical procedure. The patient had some low back pain that was felt to be related to positioning that kept her from ambulating quickly. She was discharged to return home on hospital day number three. In (b)(6) 2011, the patient was diagnosed with a pelvic abscess, diverticulitis of the colon with perforation, and had a sigmoid colectomy. The patient noted as being hospitalized several times throughout (b)(6) 2011 through (b)(6) 2012, for back and leg issues (osteomyelitis of lumbar spine, diskitis, lumbar spondylosis, and radiculopathy). The patient's current condition was not provided. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however, as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical sacrocolpopexy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335741 | 3306715 | 3203310 | 2955842201303197 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS 3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS1200 A43P9 | 2008-06-11 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on july 25th, 2013 intuitive surgical received information including medical records, of a patient who had a da vinci assisted laparoscopic total hysterectomy, bilateral salpingo-oophorectomy procedure on (b)(6) 2008. The operative report does not contain any allegations of a malfunction of a da vinci system, instrument or accessory. The patient had a ct scan of her abdomen/pelvis on (b)(6) 2008 which revealed fluid collection in the pelvis and rectum, most likely consistent with seroma or hematoma. The patient was seen for similar symptoms and received course of augmentin on (b)(6) 2008. The patient experienced a reaction to augmentin (joint pains) and discontinued antibiotics on (b)(6) 2008. On (b)(6) 2008 the patient experienced orange/blood-tinged watery vaginal bleeding and felt achy. She soaked through several pads over the course of 2-3 hours, but bleeding then subsided. On the morning of (b)(6) 2008, the patient continued to have watery vaginal bleeding and new onset fever to 104. 8 with chills. The patient took motrin and then came to the hospital per her physician's recommendation. The patient was admitted and was started on antibiotics. She was experiencing spontaneous drainage from the vagina and thus no attempt at additional drainage was undertaken. The patient was discharged on (b)(6) 2008. No further information was provided about the patient's current status. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci standard laparoscopic total hysterectomy, bilateral salpingo-oophorectomy procedure. | DA VINCI STANDARD SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335744 | 3306717 | 3203312 | 2955842201303199 | 1 | NAY | Manufacturer | OtherConsumer | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2010-09-15 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on 07/25/2013 intuitive surgical, inc. (isi) received information, including medical records, of a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2010 for dysmenorrhea, menorrhagia, and dyspareunia. According to the operative report for the da vinci si surgical procedure, the surgeon found a retroverted fibroid uterus, no evidence of endometriosis, and normal appearing ovaries and appendix. The report mentioned the use of the bipolar cautery, and endoshears instruments during the surgical procedure. There was no report of any da vinci system or instrument malfunction occurring during the procedure. The patient tolerated the procedure well and was taken the recovery room in stable condition. No complications were noted. The patient was sent home postoperative day 2. On (b)(6) 2010 the patient was admitted to the hospital for abdominal pain. The medical record indicated that the patient was having severe abdominal pain, nausea, vomiting, fever, and bloody stool. A cat scan was done on the pelvis and abdomen showing findings consistent with inflammatory or infectious process within the post-surgical site with fluid collections in the left adnexal region and anterior to the vagina. The patient was treated with iv antibiotics and pain medications and was recommended to wait 6 weeks for the inflammatory changes to heal before repeating the scope. According to the documentation received, the patient was hospitalized for 4 days and was hospitalized again in (b)(6) 2013 for bowel obstruction. The patient's current condition was not provided. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335746 | 3306723 | 3203318 | 2955842201303198 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-11-09 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on july 25, 2013 intuitive surgical, inc. Received information including medical records of a patient who underwent a da vinci si sacrocolpopexy procedure on (b)(6) 2011 due to pelvic prolapse. There was no report in the patient's operative (op) report that a malfunction of the da vinci surgical system, instrument or accessories occurred and there was no indication that the patient experienced any intra-operative complications. It was stated in the op report that after the surgical task of securing the patient's vaginal mesh was performed, excellent hemostasis was noted. The surgical procedure was successfully completed and the patient was taken to the recovery room in stable condition. The patient was discharged from the hospital on (b)(6) 2011. On (b)(6) 2011 the patient returned to the hospital's emergency room complaining of lower abdominal pain and shoulder pain. The patient underwent a diagnostic test and it was discovered that she had a small amount of free air. A chest x-ray was performed and the patient became progressively worse in her symptoms. On examination, the patient was in fairly acute distress. A cbc test showed that the patient's white count was 2900 and her hemoglobin was 10. 7. From (b)(6) 2011 through (b)(6) 2011 the patient remained hospitalized and continued to be treated. On (b)(6) 2011 the patient underwent an exploratory laparotomy. During the laparotomy an area of diverticulum with a 4 to 5 mm perforation was discovered with fecal contamination. There was some evidence of thermal effect in the area of the perforation. The mesh was in the exudate that was present. Extensive lavage was carried out and the rest of the patient's colon and small bowel were inspected. No other injuries were noted. The affected area was debrided and the defect was repaired with sutures. The mesh was removed and a transverse colostomy was performed. A blake drain was placed in the patient's pelvis and irrigation and cleansing of the particulate matter and contaminated liquid material was removed. A colostomy appliance was placed and the patient was closed. The patient tolerated the procedure well and was in stable condition at the end of the procedure. On (b)(6) 2011 the patient was discharged to for rehabilitation. The patient was discharged home from rehabilitation on (b)(6) 2011 in a stable, alert, and oriented condition. | Based on the information provided, intuitive surgical has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained a bowel injury during a da vinci si sacrocolpopexy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335748 | 3306726 | 3203321 | 2955842201303200 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-12-15 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on july 25, 2013 intuitive surgical, inc. Received information including medical records of a patient that underwent a da vinci si hysterectomy with diagnostic cystoscopy on (b)(6) 2011. The patient developed post -surgical complications on (b)(6) 2012, the patient returned to the operating room for an acute right adnexal pain with a pelvic mass. A diagnostic laparoscopy, right salpingo-oophorectomy, and lysis of adhesions were performed for a hemorrhagic right ovarian cyst and adhesions. The surgeon described mild adhesions of the bowel to the vaginal cuff that and some hemorrhagic areas on both sides of the vaginal cuff that were biopsied. The patient's surgical pathology report was not available for review. On (b)(6) 2012 the patient visited the emergency room with reports of several hours of acute onset of abdominal pain that occurred during sexual intercourse. A ct scan showed areas of pneumoperitoneum with fluid in the pelvis. The patient underwent a diagnostic laparoscopy with drainage of a pelvic abscess. It was noted in the diagnostic laparoscopy that the right side of the vaginal cuff was found to be open with the small bowel protruding through the vaginal cuff. The surgeon closed the vaginal cuff dehiscence which showed inflammation on the vaginal cuff as well as the peritoneal surfaces. At the end of the procedure the surgeon documented that the vaginal cuff appeared to be intact. On (b)(6) 2012 the patient underwent a diagnostic laparoscopy for chronic pelvic pain and interstitial cystitis. Upon examination the surgeon found scar tissue, endometriosis, and granulation tissue on the vaginal cuff, which were biopsied and sent for pathological review. The surgeon noted that the left ovary was multi-cystic with an endometrioma, resulting in a left salpingo-oophorectomy. On (b)(6) 2012, the patient underwent a procedure for vaginal cuff dehiscence repair. The surgeon noticed an area of separation of approximately 1 cm in length; the remaining vaginal cuff was intact and appeared strong. The surgeon noticed substantial hemorrhagic material adherent to the small bowel in the posterior cul-de-sac. On (b)(6) 2012 the patient underwent a cystoscopy for chronic pelvic pain with interstitial cystitis. The surgeon noted that the areas of petechiae, which are signs of cystitis in the bladder, were less than previously. On (b)(6) 2012, the patient went in for surgery where the surgeon performed an exploratory laparotomy for chronic pelvic pain and chronic vaginal cuff dehiscence. The surgeon described the vaginal cuff as having hemorrhagic material at the right angle, where her pain was localized; this area was completely removed and sent for pathologic examination. Several areas of endometriosis were also resected off of the left vaginal cuff. All of the above surgeries were tolerated well with no intraoperative complications noted. | Based on the information provided, intuitive surgical has not determined the root cause of the intra-operative injury subsequent post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical procedure. | DA VINCI SI SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335691 | 3306739 | 3203469 | 2955842201303201 | 1 | NAY | Manufacturer | OtherConsumer | Health Professional | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS3000 A60P8 | No | Yes | 2013-07-25 | No | No | Reuse | 2013
IS3000 A60P4 | 2011-12-15 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on july 25, 2013, intuitive surgical received information, including medical records, of a patient who underwent a da vinci si radical prostatectomy procedure on (b)(6) 2011. The patient's operative notes state that during the da vinci si radical prostatectomy procedure the surgeon encountered bleeding of the dorsal vein complex, which was cauterized. However the vein complex started to bleed again. The instrument was changed to a clip applier however access to the port was lost. The da vinci si system was undocked, and all ports were removed and the procedure was converted a laparotomy procedure. The patient was reported to be stable when transferred to the recovery room. During a follow up visit while the patient was in recovery, he was noted to have some respiratory distress and was transferred to icu. Shortly after the patient stabilized and was returned back to the general hospital floor. He was discharged home on (b)(6) 2011. The discharge summary provided states that the patient tolerated the procedure well with minimal complications. The patient reported difficulty urinating and on (b)(6) 2012 the patient had a cystoscopy which showed a very tight urethral stricture and bladder neck contracture. The patient underwent a resection of the bladder neck. No complications were reported during this procedure. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The patient operative reports do not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient had post operative complications after a da vinci si radical prostatectomy procedure. | DA VINCI SI SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335692 | 3306740 | 3203335 | 2955842201303202 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS3000 A60P4 | No | Yes | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-12-12 | Injury | Hospitalization | As part of a legal mediation effort on july 25, 2013 intuitive surgical received information including medical records of a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2011 due to pelvic pain and endometriosis. There was no allegation in the patient's operative report that a malfunction of the da vinci surgical system, instruments or accessories occurred and there was no indication that the patient experienced any intra-operative complications. The planned surgical procedure was completed and the patient was taken to the recovery room in good condition. The patient was discharged from the hospital on (b)(6) 2011. During a post-surgical follow-up examination on (b)(6) 2011 the patient complained that she experienced burning during urination. On (b)(6) 2011 the patient returned to the hospital's emergency room complaining of fever, chills and abdominal pain. Lab tests performed revealed that the patient's had a substantially elevated white blood cell count suggestive of infection. A ct scan showed thickening of the rectum and a ct scan with rectal and oral contrast was also performed that showed a small collection in the pelvis measuring approximately 3cm in dimension. No extravasation was noted. Per the patient's medical records, the patient's healthcare providers made the decision to treat the patient with iv antibiotics and placed the patient on a liquid diet. Throughout the patient's hospitalization the patient continued to be treated with antibiotics and closely monitored. On (b)(6) 2011 the patient was responding well to treatment and was continuing to improve. Reportedly, the patient was afebrile and was discharged from the hospital on (b)(6) 2011. | Based in the information provided, intuitive surgical has not determined the root cause of the infection that the patient developed. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event this complaint is being reported due to the following conclusion: the patient developed an infection post op da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335887 | 3306874 | 3203469 | 2955842201303203 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-05-19 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si pancreatectomy and splenectomy procedure on (b)(6) 2011. No intra-operative complications were documented in the operative report for the da vinci si surgical procedure. The surgeon noted that the patient tolerated the procedure well and without complications. The patient was discharged home on (b)(6) 2011. There were no reports of a malfunction of the da vinci si system, an instrument, or an accessory during the surgical procedure. On (b)(6) 2011, the patient was admitted to a hospital for a reported complaint of gradually worsening abdominal pain. A ct scan revealed that the patient had post-operative and recurrent pancreatitis. The ct scan also found small cystic masses on the pancreas. On (b)(6) 2011, the patient underwent a pancreatic debridement procedure as well as debridement of multiple peripancreatic abscesses. Two drains were placed during the procedure. The operative report indicated that the patient tolerated the procedure well and without complications. According to the patient's medical records, the patient developed respiratory distress during the hospitalization and on (b)(6) 2011, the patient underwent a tracheostomy procedure due to developing respiratory failure. According to the operative report, the patient tolerated the procedure without complications. On (b)(6) 2011, the patient was evaluated for peristomal bleeding and aspiration from the tracheostomy site. The patient subsequently underwent a bronchoscopy, decannulation of the trachea, and direct control of the bleeding vessel. Sutures were placed to control the bleeding and obtain hemostasis. In addition, a revision of the tracheostomy with dilatation of the tracheostomy tract using hegar dilators was performed. Also, a tracheostomy tube was placed. On (b)(6) 2011, the patient received consultation for acute respiratory failure and on (b)(6) 2011, the patient received consultation for polymicrobial pneumonia. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. The operative report does not contain any allegation of a malfunction of a da vinci system, instrument, or accessory. No previous complaint was reported relating to this event. Isi has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. This complaint is being reported due to the following conclusion: the patient's medical records indicate that the patient developed a pancreatic abscesses after undergoing a da vinci si pancreatectomy and splenectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335970 | 3306975 | 3203570 | 2955842201303204 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-26 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-05-08 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(6) 2013, intuitive surgical inc. (isi) received information concerning a patient who underwent a da vinci si hysterectomy with salpingo-oophorectomy on (b)(6) 2012. The documentation received states that the patient sustained vaginal dehiscence and partial evisceration. According to the legal complaint, the patient was found to have adenomyosis and colon adhesions during the surgical procedure. A post-operative cystoscopy was found to be normal. The legal complaint also indicated that the surgeon later noted in an office note that she had taken extra care to take wide bites of tissue with cuff closure due to patient history of heavy smoking. The patient was reportedly discharged home in stable condition on the first post-operative day. The patient was found to have no complications during a post-operative follow-up visit on (b)(6) 2012. On (b)(6) 2012, the patient was reportedly admitted to a hospital for a vaginal cuff dehiscence and partial evisceration. The patient underwent an open exploratory laparotomy with closure of the vaginal cuff. The information received stated the patient has recovered and appears to be doing well. No additional information is available. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Isi has reviewed the system logs of the site with a procedure date of (b)(4) 2012. No system errors were found to have occurred during the surgical procedure. This complaint is being reported due to the following conclusion: the patient sustained vaginal cuff dehiscence and partial evisceration after undergoing a da vinci surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334579 | 3305439 | 3202034 | 2955842201303130 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-08-23 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical, inc. (isi) received a summary outline of the patient's medical records and information regarding a patient that underwent a da vinci si hysterectomy procedure and lysis of adhesions procedure on (b)(6) 2012. No complications during the procedure were noted to have occurred. On (b)(6) 2012 the patient was admitted through the hospital's emergency room with complete vaginal prolapse. On (b)(6) 2013 the patient underwent a diagnostic laparoscopy procedure with the da vinci surgical system and it was found to have multiple intestinal adhesions, a pelvic abscess and vaginal cuff dehiscence. The surgical procedures performed included enterolysis-lysis of the adhesions, evacuation of the abscess, and repair of the vaginal cuff dehiscence. The patient's appendix and small bowel were also inspected by a urologist consultant and at the time of the surgical procedure and it was reportedly determined by the surgeon that no other procedures were required. The procedure was completed and there were no complications. The patient was discharged from the hospital on (b)(6) 2012. It was noted in the patient's discharge summary that the patient's post-operative course was unremarkable. No other information was provided. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced post surgical complications after undergoing a da vinci si hysterectomy and lysis of adhesions procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334584 | 3305452 | 3202047 | 2955842201303131 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
| 2009-08-10 | Injury | HospitalizationLife ThreateningRequired Intervention | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical inc. (isi) received information concerning a patient who underwent a da vinci hysterectomy with bilateral salpingooophorectomy and lymph node dissection procedure on (b)(6) 2009. The documentation received states that the patient sustained a thermal injury to the small bowel causing a large abscess to form and causing a large bowel obstruction. Also stated that the jejunal enterotomy was consistent with a low voltage injury to the small bowel that occurred from an insulation defect in the shaft of the monopolar scissors used during her robotic hysterectomy. As a result of the bowel injury, the patient developed a large right perio-colic abscess and required an extensive operation with removal of the right colon. As a direct consequence of the bowel injury, she had sepsis and later developed a hospital-acquired pneumonia requiring readmission and extensive treatment. The surgery performed on (b)(6) 2009, found that the patient had a normal anatomy with some adhesions from the previous cesarean but there was no other evidence of extra-uterine disease. The patient required exploratory laparotomy for postop bleeding. The patient was discharged home on (b)(6) 2009. On (b)(6) 2009, the patient was evaluated at another hospital for what was reported as increased abdominal distension and discomfort. She underwent an exploratory laparotomy which revealed a small bowel perforation involving the mid jejunum and a large peri-colonic abscess which required a repair of the mid jejunal enterotomy extended right hemi-colectomy, appendectomy, and cholecystectomy. The patient was hospitalized until (b)(6) 2009, according to the legal complaint. On (b)(6) 2009, the patient was evaluated for a complaint of shortness of breath. She was diagnosed with multifocal hospital-acquired pneumonia. The patient received wound-vac therapy for her incision and was treated with triple intravenous antibiotic therapy. She was discharged from the hospital on (b)(6) 2009. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. A follow-up mdr will be submitted if additional information is received. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained a bowel injury during a da vinci hysterectomy procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334596 | 3305471 | 3202066 | 2955842201303132 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | Yes | Health Professional |  | No | No | 2013-07-25 | Device Not Retu | No | Unkown | 2013
IS3000 A60P8 | 2011-06-30 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013 intuitive surgical received information regarding a patient that underwent a da vinci si hysterectomy procedure. The patient had a large uterus and it was noted during the surgical procedure that the patient had fibroids. During the procedure, the surgeon reportedly had to enter the patient's vagina in two spots due to the distortion of the anatomy caused by the fibroids. These spots were repaired vaginally without consequence at the surgery. There were no complications noted at the time of the surgery. The patient was discharged home the next day and on (b)(6) 2011, the patient experienced problems. A ct scan performed on (b)(6) 2011 noted a possible bladder injury. On (b)(6) 2011, the patient underwent a surgical vaginal repair of a vesicovaginal fistula with left ureteral catheter placement, attempted right ureteral catheter placement, and cystoscopy. In (b)(6) 2011 the patient was re-admitted and the patient underwent an exploratory laparotomy with abdominovesical fistula repair on (b)(6) 2011. The patient was discharged after two days. | Based in the information provided, intuitive surgical has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2011 found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: the patient sustained a bladder injury during a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334674 | 3305499 | 3202094 | 2955842201303137 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-07-18 | Injury | Required Intervention | As part of a legal mediation effort, on july 25th, 2013 intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2012. The documentation received states that the patient sustained an injury to her bladder during the surgical procedure. In regards to the da vinci si hysterectomy procedure, the patient's medical records indicate that during creation of the bladder flap, it was noted that there was a 5 mm incision made in the dome of the bladder. The patient's medical records indicate that the cystotomy was inadvertently made in the dome of the bladder. The surgeon requested a urologist to the or, but the urologist declined because he was not familiar with da vinci. The urologist instructed the surgeon to repair the cystotomy. Per instructions given by a urologist via an intraoperative urology consult, the surgeon repaired the cystotomy with 2 layers of sutures. According to the operative report, after the repair was completed, the patient's bladder was filled with 300 ml of diluted indigo carmine solution and the repair was found to be watertight. Towards the conclusion of the surgical procedure, the patient's bladder was checked again and filled with 300 ml of dilute indigo carmine solution. The operative report indicated that a small leak was noted in the previous repair. The small leak was repaired with additional sutures. The repair was noted to be watertight with no further leakage. The operative reporter indicated that the patient tolerated the procedure well. On (b)(6) 2012 the patient complained of severe flank pain. The patient was then seen by a urologist on an unspecified date for a possible left ureteral injury. The urologist's impression was mild nephrosis, probably secondary to a postoperative hematoma, but no obvious extravasation and no obvious ureteral fistula or injury. The urologist report states he believes this could be both chronic changes and acute changes. The patient was placed on an oral antibiotic for 14 days. On (b)(6) 2012, the patient was seen by another urologist due to increasing left flank pain that occurred on and off since her hysterectomy surgery. On (b)(6) 2012 during a follow up appointment the urologist ordered a ct scan of the patient's abdomen and pelvis. The ct scan, x-ray and ultrasound revealed that the patient had a rectal mass. The patient was referred to an oncologist. On (b)(6) 2012, the patient was seen by a gynecologist for progressively worsening pelvic pain and evaluation of the rectal mass. The gynecologist's impression was a possible post-operative stricture of the left ureter and a rectal polyp. On (b)(6) 2012, the patient underwent a colonoscopy. On (b)(6) 2012 the patient underwent a cystourethroscopy and left retrograde pyelogram for a ureteral stent placement by a urologist for persistent left flank discomfort. During the procedure, the urologist found a left ureteral obstruction with ureteral ulceration and necrotic material, but because of the obstruction, he was unable to place a stent. The urologist indicated that it appeared the ulcer site represented the site of cystotomy repair with the original (b)(6) 2012 surgery; it was unclear to him if the repair line might have involved the ureteral opening thus resulting in the ulcer and obstruction. On an unspecified date a radiologist was able to successfully insert a percutaneous nephrostomy tube into and across the left ureteral orifice to effectively bridge the obstructed area. According to the patient's medical records, a urologist indicated that the nephrostomy tube had been converted to an indwelling ureteral stent in (b)(6) 2012. The medical records for the conversion of the nephrostomy tube to an indwelling ureteral stent were not provided. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2012. No system errors were found to have occurred during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the patient sustained an inadvertent cystotomy during the da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334675 | 3305500 | 3202095 | 2955842201303138 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | No | No | Reuse | 2013
IS3000 A60P8 | 2012-08-07 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013 intuitive surgical received information including medical records of a patient who underwent a da vinci si trachelectomy procedure for removal of her cervix and attached uterine tissue on (b)(6) 2012. On (b)(6) 2012 the patient returned to the hospital's emergency room complaining of abdominal pain. The operative report regarding the surgery on (b)(6) 2012 indicated that the patient underwent a da vinci si trachelectomy procedure with extensive lysis of adhesions for the management of endometrial hyperplasia in the patient's residual uterine stump after a previous subtotal hysterectomy. Based on the information indicated in the patient's op report, there was no allegation that a malfunction of the da vinci si surgical system, instruments or accessories occurred during the surgical procedure and there was no indication that the patient experienced any intraoperative complications. Hemostasis was achieved after closure of the vaginal cuff was completed. The initial pathological evaluation of the residual stump suggested that there was evidence of hyperplasia. After the procedure was completed, blood was mopped from the patient's pelvis and hemostasis was achieved. A vaginal examination performed on the patient confirmed that the vaginal cuff was completely closed and dry. The patient's urine was noted to be clear. During the patient's er examination on (b)(6) 2012 the patient noted increasing pain and had multiple episodes of diarrhea and mild tachycardia. During a ct scan it was discovered that the patient had a large amount of free air within the abdomen and pelvis. The ct showed that the collection of fluid posterior to the bladder may have been due to a post-surgical abscess. The patient then underwent an exploratory laparotomy procedure. During the laparotomy procedure it was discovered that the patient had extensive abscesses in the pelvis with free stool and inflamed peritoneum. The patient's pelvis was irrigated and debrided extensively. Inspection of the patient's sigmoid colon found that the patient had a perforation just below the peritoneal reflection and second perforation was more proximal on the sigmoid colon. A sigmoid resection was performed to repair the damage to the patient's colon. After completion of the sigmoid colon resection the patient's pelvis was extensively irrigated with 5 liters of warm normal saline and the remaining necrotic tissue was debrided. Bacitracin irrigation was also performed. An ostomy site was created and upon completion of the ostomy, the patient's subcutaneous tissue was extensively irrigated and then closed. The ostomy was noted to be pink at the end of the procedure. The patient tolerated the procedure well and was taken to the icu in a stable fashion still intubated. The patient was to be monitored over night. On (b)(6) 2012 the patient experienced unexplained vaginal bleeding. A vaginal examination performed on the patient while under anesthesia found that the patient had a small dehiscence on the left side of the vaginal cuff. There was also some oozing from the remaining suture line, which possibly occurred due to recent use of coumadin. The surgical procedure was technically difficult due to the patient's very long and narrow vaginal vault. The patient's vaginal cuff was repaired using suture and the remaining vaginal cuff was oversewn. A small amount of oozing was noted from the right aspect of the incision and the affected area was sutured to control the bleeding. Cautery was also applied to the mucosa where it had been denuded. The vaginal cuff was hemostatic after completion of the procedure. The procedure was successfully completed and the patient tolerated the procedure well. | On (b)(4) 2013, intuitive surgical, inc. , contacted the surgeon who performed the patient's trachelectomy procedure. The surgeon indicated that the patient had cancer and extensive intraabdominal adhesions and that he was very careful while taking down the adhesions. The surgeon stated that there was no malfunction of the da vinci si surgical system, instruments or accessories during the surgical procedure and that it was his belief that the injury sustained by the patient was due to the patient's difficult anatomy. This complaint is being reported due to the following conclusion: the patient sustained an injury during a da vinci surgical si trachelectomy procedure. Based on the information provided, it has been determined that the da vinci si surgical system did not malfunction in a way that caused or contributed to the injury sustained by the patient, but rather occurred while the surgeon was performing the surgical tasking of taking down the patient's adhesions. Upon review of the site's system logs for the reported procedure date of (b)(6) 2012, it was found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334630 | 3305506 | 3202101 | 2955842201303133 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-05-01 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(6) 2013 intuitive surgical, inc. (isi) received information, including medical records, of a patient who underwent a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure on (b)(6) 2012 for endometrial adenocarcinoma. According to the operative notes, the patient had a normal size approximately 6 week uterus with normal bilateral tubes and ovaries, no gross lesions seen. There was a minimal area of adenocarcinoma seen in the endometrium; however, no myometrial invasion was present. All lymph nodes were normal in appearance with no evidence of metastatic disease. No damage to bowel, bladder, or blood vessels was noted throughout the procedure. The patient did not have any complications during the procedure and was taken to the recovery room in good condition. There was no allegation or report of a malfunction of the da vinci system, instruments, or accessories in the operative report. The post operative hospital course was without complication and the patient was discharged on (b)(6) 2012 in stable condition. According to the medical record, approximately 2 weeks after the da vinci surgical procedure ((b)(6) 2012), the patient went to the hospital with complaints of increased abdominal pain and high fever. The ct scan of abdomen and pelvis was performed and revealed there were multiple intra-abdominal fluid collections, consistent with a possible abscess. The patient underwent ir drainage and was treated with pain medications and antibiotics. Over the next few days the patient underwent antibiotic treatment until her symptoms improved. She was discharged on (b)(6) 2012. After going home, she continued to have fevers and significant bouts of pain. On (b)(6) 2012 the patient had an outpatient procedure to attempt to drain the abscess, but this was unsuccessful due to a blood vessel being in the way. The patient reportedly continued to have occasional bouts back and forth visits to the hospital through the summer. On (b)(6) 2012, as per the legal documentation (no medical record information was provided) the patient was admitted to the hospital for significant pelvic and abdominal pain and had a diagnosis of possible abscess. She received iv antibiotics and drainage of mass performed by ct guidance. On (b)(6) 2012, the patient was discharged with a drain and had drain for 11 days until it was removed by the doctor. Patient's current condition was not provided. | Based on the limited information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. Based on the provided information, there is no evidence that a malfunction occurred with the da vinci si system, instruments, or accessories. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334638 | 3305510 | 3202135 | 2955842201303134 | 1 | NAY | Manufacturer | OtherConsumer | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS2000 A5PP8 | 2011-08-19 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013 intuitive surgical received information including limited medical records, of a patient who developed post- surgical complications after having a da vinci s hysterectomy procedure on (b)(6) 2011 for menorrhagia and fibroid uterus. The patient's post operative recovery was uneventful and she was discharged the next day after having the da vinci procedure. No additional operative information was provided. Patient returned to the hospital (b)(6) 2011 after bilateral lower abdominal pain and was found to have pelvic abscess with an abdominal ct. She had a wash out on (b)(6) 2011. She remained afebrile, however continued with leukocytosis. There was no change in abscesses on repeat ct scan on (b)(6) 2011. On (b)(6) 2011, patient had diagnostic laparoscopy converted to an exploratory laparotomy and repeat washout. A wound vac and jp drain was placed after this, which was removed 3 days later, and 5 days later, and the patient's abdominal wound closed. The patient's symptoms continued to improve. The patient was discharged (b)(6) 2011 without additional antibiotics. During her post operative scheduled follow up, the patient was noted to be doing well. The patient was doing well; however had complaints of feeling cold and decreased appetite. The patient was advised to return back to the hospital. Patient was started on levofloxacin. Ct guided drainage was attempted, but unsuccessful. Repeat ct scans show some improvement in patient's pelvic abscesses. As the surgical options were not felt to be optimal, the patient continued on avelox for therapy. The patient reported not to have undergone a follow up ct scan. No further clinical information was provided about the procedure. At this time there is no information that a malfunction of the da vinci system, instrument or accessory occurred at the time of the procedure. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced post- surgical complications after undergoing a da vinci surgical hysterectomy procedure. | DA VINCI S SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334657 | 3305514 | 3202109 | 2955842201303135 | 1 | NAY | Manufacturer | OtherConsumer | Health Professional | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS2000 A5PP8 | No | Yes | 2013-07-25 | No | No | Reuse | 2013
IS3000 A60P8 | 2011-01-31 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013 intuitive surgical received information regarding a patient who experienced post-surgical complications after undergoing a da vinci si hysterectomy for endometrial cancer on (b)(6) 2011. Per the information provided, there were no surgical complications experienced by the patient during da vinci hysterectomy procedure. Intra operative blood loss was reported to be 50cc. Reportedly, the patient's vaginal cuff was closed with running vicryl suture. Post operatively the patient had hyponatremia, which required some intravenous fluid hydration, but otherwise the patient did well and was discharged home in good condition. On (b)(6) 2011 the patient was seen by her doctor, and during her follow -up visit it was indicated that her vaginal cuff was well supported and intact with no tenderness. The patient was asked to follow up in with her doctor in six months. On (b)(6) 2011 the patient complained of vaginal pressure and she went to emergency room. The patient was not diagnosed with any significant issues and was discharged home with a recommendation of follow up. Allegedly, while on vacation, on (b)(6) 2011 while the patient was walking on the beach, she had an acute vaginal cuff dehiscence with prolapse of her small bowel to the knees. The patient was immediately rushed to the emergency room at a local hospital. On (b)(6) 2011 the patient underwent an emergency surgery for her condition. After being placed under anesthesia, the patient was manually reduced back to her intra-abdominal cavity, and the damage to her vaginal cuff was closed with interrupted vicryl sutures. Good hemostasis was noted. No other information was provided. | Based in the information provided, intuitive surgical has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2011 found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334648 | 3305519 | 3202114 | 2955842201303136 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2007-03-05 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si radical prostatectomy procedure on (b)(6) 2007. The documentation provided states that the patient developed a psoas abscess after undergoing the surgical procedure. The operative report related to the da vinci si radical prostatectomy procedure was not provided and no additional details are available at this time. According to the patient's medical records, two weeks postop the patient reported complaints of pain in his left hip which progressively got worse. On (b)(6) 2007, the patient was admitted to a hospital for reported complaints of left pelvic pain, hip pain, and fever. A ct scan revealed that the patient had an iliopsoas abscess. The patient's medical records stated that a fistulogram was performed and a fistula was found between the colon and the abscess. On (b)(6) 2007, the patient underwent a partial colon resection with anastomosis procedure and drainage of the abscess. On (b)(6) 2007, the patient was discharged home with antibiotics, pain pills, and some outpatient therapy. The patient's discharge summary stated that the patient continued to recover very nicely after his surgery. On (b)(6) 2007, a ct scan performed on the patient revealed that the left psoas fluid collection was resolved and there was no evidence of a recurrent or persistent abscess. On (b)(6) 2007, the patient's progress notes stated that the patient had no complaints and was doing well. No further clinical information was provided. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not have any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient developed a psoas abscess after a da vinci prostatectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334668 | 3305525 | 3202120 | 2955842201303139 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-11-13 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013, intuitive surgical received information including medical records of a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2012. The surgical procedure was completed and the patient was transferred to the recovery room in stable condition. There was no allegation in the patient's operative report that a malfunction of the da vinci surgical system, instruments or accessories occurred and there was no indication that the patient was injured or experienced any intra-operative complications. The patient's history and physical report states that the patient returned to the hospital's emergency room on (b)(6) 2012 complaining of fever, severe abdominal pain and distention. A ct scan performed on the patient revealed that the patient had dilated loops of small bowel, intraperitoneal air, abdominal and pelvic free fluid and some high attenuation within the pelvic free fluid. From (b)(6) 2012 through (b)(6) 2012 the patient underwent several tests, examinations and treatments due to her condition; however, her condition did not improve. Reportedly, on (b)(6) 2012 the patient underwent an exploratory laparotomy, lysis of adhesions and enterorrhaphy to repair the damage to her bowel. The bowel was found to be deseasonalized and had to be repaired with sutures. There was no other damage noted to the patient's bowel and hemostasis was excellent. The procedure was completed and the patient tolerated the procedure well. The patient left the or in satisfactory condition. The patient was discharged from the hospital on (b)(6) 2012. No medical information is provided since the patient's discharge on (b)(6) 2012. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not have any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained a bowel injury during a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334683 | 3305530 | 3202125 | 2955842201303140 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2012-09-12 | Injury | HospitalizationLife ThreateningRequired Intervention | As part of a legal mediation effort, on (b)(6) 2013, intuitive surgical, inc. (isi) received information, including medical records, of a patient who underwent a da vinci s hysterectomy procedure on (b)(6) 2012 for uterine fibroids, menorrhagia and anemia. One week prior to the da vinci surgical procedure ((b)(6) 2012), the patient went to the emergency room (er) with complaints of generalized weakness, dizziness, and nausea. For one week prior to the er visit, she had heavy period and passed some clots. The patient's potassium and hemoglobin levels were low. She received a blood transfusion and her repeat hemoglobin was 12 and her potassium was normal. She was diagnosed with acute blood loss anemia, symptomatic secondary to menorrhagia and hypokalemia. The patient's condition was overall stable and was sent home the same day. On (b)(6) 2012, the patient underwent a da vinci s hysterectomy procedure. During the procedure, it was noted that the patient had a markedly large uterus (16-week uterus with multiple uterine fibroids) along with dense pelvic adhesions from the omentum to the anterior abdominal wall. The adhesions were lysed with monopolar scissors and pk instrument without complication, taking about 10 minutes to complete. Following the hysterectomy, no lacerations were found in the vagina and the vaginal cuff was sutured with 2 layers of running suture. Good hemostasis and vaginal cuff closure were both noted at the end of the procedure. According to the post-operative report, the patient tolerated the procedure well, had no complications, and she was taken to the recovery room in stable condition. Three days later after the da vinci procedure ((b)(6) 2012), the patient went to an emergency room (er) at another hospital complaining of abdominal pain and noting that she had not had a bowel movement for 7-8 days. She reported she had been discharged home following the hysterectomy with a hemoglobin of 7. 7, and in the er, the patient was also noted to have leukocytosis (wbc 23000). She was admitted to the hospital, given soapsuds enemas for constipation, transfused with 3 units of blood for her low hemoglobin and started on antibiotics for a urinary tract infection. Her abdominal pain resolved and she was discharged home on (b)(6) 2012. Ten days after the da vinci procedure ((b)(6) 2012), the patient was transported via emergency medical services (ems) to the same er with complaints of heavy vaginal bleeding and a progressively worsening foul-smelling vaginal discharge. A ct scan of the abdomen and pelvis revealed evidence of a pelvic abscess and she was admitted and started on antibiotics. On (b)(6) 2012, a gynecologist performed a diagnostic laparoscopy with evacuation of a hematoma and repair of vaginal cuff dehiscence. According to the patient's operative report, the doctor identified a loosened blue vicryl stitch in the vaginal cuff as well as copious amounts of blood clots (hematoma) posterior to the cuff and inflammation anterior to the cuff. The hematoma was evacuated and the cuff was surgically closed and intact at the end of surgery. The surgeon also noted that the patient had dense bowel adhesions. The vaginal cuff was noted to be intact and there was no bleeding. The patient tolerated the procedure well. Postoperatively, the patient's condition quickly improved and she was discharged home on (b)(6) 2012. Isi reviewed the post-operative report and did not find any allegation of a malfunction of a da vinci system, instrument or accessory. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2012 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334689 | 3305535 | 3202130 | 2955842201303141 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-04-30 | Injury | Required Intervention | As part of a legal mediation effort, on july 25th, 2013 intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si hysterectomy, left salpingo-oophorectomy, adhesiolysis, and cytoscopy procedure on (b)(6) 2012. The legal complaint alleged that the patient sustained vaginal cuff dehiscence as a result of a robotic-assisted hysterectomy. According to the legal complaint, multiple adhesions were found on the left abdominal wall, the cul-de-sac and left pelvic gutter, and around the left adnexal mass during the surgery. In addition, the legal complaint indicated that the procedure was tolerated well and no complications were noted. Based on the operative report, a cystoscopy was performed at the conclusion of the surgical procedure which confirmed good efflux of indigo carmine from both ureters. Also, it was noted in the operative report that there were no injuries to the bladder wall and the patient was taken to the recovery room in stable condition. On (b)(6) 2012, the patient was evaluated at a clinic for complaints of heavy bleeding and clots, noted to be the second episode of bleeding. The information provided indicated that the patient was seen in an emergency room (er) the preceding weekend for bleeding, however, the medical records related to that er visit were not provided. Upon examination, the physician found a 1cm cuff defect in the midline with visible clot and increased bleeding with manipulation/exam. The physician repaired the cuff with sutures. During the surgery, the physician noted that a vaginal cuff opening of approximately 2-3 cm with bleeding from the posterior cuff was observed. At an unspecified time shortly after the (b)(6) 2012 surgery, an inferior vena cava filter was placed due to contraindication of use of anticoagulant therapy in the presence of active bleeding. On (b)(6) 2012, the patient underwent another surgical procedure by another surgeon for repair of a complete mucosal separation of the vaginal cuff. The surgeon indicated that the vaginal cuff had complete separation of the mucosa with intact sutures in the posterior vaginal mucosa and lacerations directly across from the sutures suggesting a possible traumatic cause of the cuff separation. The legal complaint indicated that the surgeon did not elaborate further on how this trauma could have occurred. The post-operative records were not provided. On (b)(6) 2012, the inferior vena cava filter was removed. The legal complaint suggested that the vaginal bleeding was controlled and had not recurred since the final surgery. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. Isi has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. The operative report did not indicate that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the surgical procedure. This complaint is being reported due to the following conclusion: the legal complaint alleged that the patient sustained a vaginal cuff dehiscence as a result of undergoing a da vinci hysterectomy procedure. However, at this time, it is unknown how the patient's bowel injury occurred and if the da vinci surgical system caused or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334710 | 3305540 | 3202135 | 2955842201303143 | 1 | NAY | Manufacturer | Other | Other | InitialFollowup | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | No | No | Reuse | 2013
IS3000 A60P8 | 2011-12-06 | Injury | Required Intervention | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical, inc. (isi) received information regarding a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2011. It was reported in the urologist consult operative report that an injury to the patient's bladder occurred during the hysterectomy procedure. The planned surgical procedure was converted to open laparotomy and a urologist was consulted to evaluate the patient's bladder. The patient's operative report for the da vinci hysterectomy was not provided; however, the urologist's consultation report was provided. The surgeon stated in his consultation report that the assessment of the patient's bladder found a 1cm defect located posterior inferior portion of the bladder. No other damage to the patient's bladder was noted. The reported laceration was repaired with suture. It was noted that there was clear efflux from both ureteral orifices. A foley catheter was placed. On (b)(6) 2011 the catheter was removed and no complications were noted at that time. No other information was provided. | Based in the information provided, intuitive surgical has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The patient's medical records do not contain any allegation of a malfunction of a da vinci system, instruments or accessories. Attempts have been made to gather additional information from the risk management group at the site, however, as of the date of this report no response has been received. No previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained a bladder injury during a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334715 | 3305541 | 3202298 | 2955842201303144 | 1 | NAY | Manufacturer | OtherConsumer | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS2000 | 2012-04-19 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013 intuitive surgical received information including medical records of a patient who underwent a da vinci s hysterectomy procedure with bilateral salpingo-oophorectomy, extensive lysis of adhesions and right ureterolysis on (b)(6) 2012. The patient was discharged from the hospital on (b)(6) 2012. The patient's operative report (op) for the surgical procedure on (b)(6) 2012 indicated that the patient underwent the da vinci hysterectomy procedure with bilateral salpingo-oophorectomy due to uterine fibroids, menometrorrhagia and mild cervical dysplasia. Based on the information in the patient's op, there was no report that a malfunction of the da vinci surgical system, instruments or accessories occurred during the procedure. The surgeon describes in the op that there were extensive adhesions and he decided to perform extensive lysis of adhesions between the bladder and the anterior part of the lower uterine segment. The surgeon states in the op that the right ureter was very close to the level of dissection. The surgical procedure was successfully completed and the patient was taken to the recovery room with stable vital signs and in satisfactory condition. The patient was discharged the following day. At a routine follow up appointment on (b)(6) 2012, the patient indicated that she was experiencing pelvic pressure in the midline and that she felt that she had to urinate all the time. The patient was referred back to the hospital as an outpatient. Per medical records, the patient went to the emergency room (b)(6) 2012 and was admitted to the hospital with fever, and right hydronephrosis. Ct scan showed pelvic fluid collection, possible hematoma. Triple iv antibiotics were administered to the patient. The patient underwent cystoscopy, right retrograde pyelography, right ureteral stent insertion, and evacuation of vaginal cuff hematoma on (b)(6) 2012. The contrast exited either posteromedially or medially from the ureter very close to the patient's bladder and into the vagina. The surgeon was unable to get contrast into the more proximal ureter and the surgeon was unable to successfully intubate the proximal ureter. The patient's bladder drained at the conclusion of the case. The patient tolerated the procedure without any immediate complications. The operative report stated a post op diagnosis of ureterovaginal fistula. On (b)(6) 2012, the patient underwent an open right ureteroneocystostomy to repair the right ureteral stricture. The op from this surgery notes that the right paravesical space was found to be very stuck and exhibited inflammation. The patient's right ureter was anastomosed to her bladder and a ureteral stent was placed into the collecting system. After completion of the procedure, good hemostasis was noted. Per the information in the op, the patient tolerated the open procedure without any immediate complications. | Based in the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date found that no system errors were generated that would have caused or contributed to post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci hysterectomy procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334691 | 3305547 | 3202142 | 2955842201303142 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS2000 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-02-09 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on july 25th, 2013 intuitive surgical received information regarding a patient who underwent a da vinci si cystic right salpingo oophorectomy and cystectomy procedure on (b)(6) 2012. In late (b)(6) 2012, a ct scan was performed showed that the patient had a complex cystic lesion on her right ovary. In late (b)(6) 2012 the patient experienced nausea and back pain. A ct scan was ordered, which showed a very marked hydronephrosis. The surgeon noted that the patient's ureter was very dilated down to the pelvic brim in the region where the prior ovarian cyst was. The surgeon made the decision to proceed with a cystoscopy right retrograde pyelogram and possible stent insertion. On (b)(6) 2012 the patient underwent a right retrograde pyelogram ureteroscopy and insertion of a double j stent. The ureteroscopy demonstrated a probable scarred area with a small caliber. The stent was placed with some difficulty. During the procedure, the ureteroscopy also showed a high grade obstruction with a pale area of the ureter, which was felt to represent an area of ischemia. A follow up ct scan was performed on (b)(6) 2012 and showed hydronephrosis with a malpositioned ureteral stent. Secondary to the recurrent hydronephrosis, the decision was made to proceed with the placement of a nephrostomy tube. On (b)(6) 2012 a nephrostomy tube was placed successfully. The patient underwent ureteral reimplantation surgery for an obstructed ureter on (b)(6) 2012. The patient was discharged on (b)(6) 2012 on antibiotics and pain medication. On (b)(6) 2012 the patient had her nephrostomy tube removed. On (b)(6) 2012 a "lystoscopy" was performed and the patient's ureteral stent was removed. On (b)(6) 2012, the patient began having problems with flank pain. A renal and bladder ultrasound performed on (b)(6) 2012 that showed hydronephrosis in the patient's right kidney, suggesting a recurrent obstruction of the right ureter. A ct-scan was performed, which showed a hydroureteronephrosis down to her anastomosis of the right side, thus indicating that the ureter was not draining well. On (b)(6) 2012, the surgeon proceeded with cystoscopy and ureteral stent placement that remained in the patient until (b)(6) 2012, at which time she had another cystoscopy performed with removal of the stent. The patient continued to have some symptoms and made the decision to be evaluated at another hospital. During the patient's evaluation at the new facility it was discovered that her renal function was severely compromised from the ureteral injury. It was recommended to the patient, that she have her right kidney removed. On (b)(6) 2012 the patient underwent removal of her right kidney. The surgery was performed without complication. | Based in the information provided, intuitive surgical, inc. Has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2012 found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: the patient had post operative complications after a da vinci si cystic right salpingo oophorectomy and cystectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334753 | 3305558 | 3202153 | 2955842201303147 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2011-03-16 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013 intuitive surgical received information including medical records of a patient who underwent a da vinci si hysterectomy procedure with bilateral-oophorectomy and cystoscopy procedure on (b)(6) 2011. On (b)(6) 2011 the patient was admitted into the emergency room of another hospital due to what the patient reported as experiencing abdominal pain. The attending surgeon suspected possible acute abdominal infection and detected hernia at the trocar site. During pelvic exploration the surgeon noted a large amount of fluid and suspected a ureteral injury. The patient required an exploratory laparotomy, a cystostomy, ureteral and bladder exploration, placement of bilateral ureteral stents, and an appendectomy. The patient underwent this procedure without any complications and she was noted to have remarkable post-operative recovery with normal return of all bladder function. The patient's operative report (op) regarding the surgery date of (b)(6) 2011 indicates that the patient's preoperative diagnosis noted tp have uterine fibroids and pelvic pain. During the surgical procedure it was discovered that the patient uterus was attached to the anterior abdominal wall with dense bands of adhesions covering the fundus to the anterior abdominal wall, encasing the bladder to the point where the bladder flap could not be identified. The ovaries exhibited cysts bilaterally and appeared to be benign. Based on the information indicated in the patient's op report, there was no indication that a malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure and there were no surgical anomalies encountered during the surgical procedure. After completion of the procedure, all of the patient's pedicles were inspected and the trocars were removed under direct visualization to ensure hemostasis. The patient wounds were then closed. A cystoscopy was performed and found that there was no trauma to the patient's bladder. Both of the patient's ureters were inspected and documented to be ejecting indigo carmine stained urine. The patient final cystoscopy was normal and the patient's urethra was inspected. A foley catheter was placed, the patient was taken down from the lithomy position, extubated and transported to the recovery room in satisfactory condition. Patient was discharged from the hospital to go home on (b)(6) 2011. The patient's discharge report alleged that the patient had an unremarkable postoperative course with normal return of bowel and bladder function. She voided successfully without any problems. The patient's progress notes alleged that the patient was admitted into the er on (b)(6) 2011 due to severe pain. A ct scan performed on the patient showed that the patient had free flowing fluid in her pelvis, a hernia at a trocar site and probable gallstones. During the laparotomy procedure, a large to moderate amount of fluid was encountered. There was a large amount of inflammation involving the entire pelvis. Several loops of large and small bowel were matted down on the pelvis and the appendix was at the center of this. The appendix was inflamed, but it appeared more likely to be secondary to periappendicitis. Careful inspection of the patient's sigmoid colon found no injury and there was no evidence of fecal material or any inflammatory fluid. The small bowel was also inspected and there was no evidence of injury. Inspection of the patient's left infundibulopelvic ligament region found that there was a large inflammatory reaction and appeared to be clear fluid emanating from the region of the left ureter. There was also an incarcerated left lower quadrant trocar site hernia with segmental omentum. The hernia was reduced and a small area of necrosis in the omentum was excised and ligated and the trocar site hernia was repaired. An appendectomy was also performed. A intravenous pyelography (ivp) was performed and no obvious extravasation was identified. Upon opening the patient's bladder, blue dye was observed. After exploration was completed, there still appeared to be fluid rolling up, particularly from the left ureteral region where there appeared to be inflammation. Stents were placed. The patient tolerated the procedure well, was extubated and transferred to the recovery room in stable condition. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The operative reports do not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2011 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical hysterectomy procedure with a bilateral-oophorectomy. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334723 | 3305566 | 3202161 | 2955842201303146 | 1 | NAY | Manufacturer | OtherConsumer | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS 1200 | 2009-01-06 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(6) 2013 intuitive surgical, inc. (isi) received information, including medical records, of a patient who underwent a da vinci hysterectomy procedure on (b)(6) 2009 for symptomatic fibroid uterus. According to the operative reports, the surgeon evaluated the patient's pelvis during the da vinci surgical procedure and indicated there was a markedly enlarged multi-fibroid uterus. After the vaginal cuff was closed, the surgeon noted a small amount of welling of methylene blue from the left pelvic brim. The surgeon suspected a bladder or urethral injury and urology was consulted. When urology arrived, a cystoscopy with retrograde placement of a guidewire was performed and revealed there was a left ureteral transection close to the bladder neck. The foley catheter was replaced in the urinary bladder and the patient was then undocked from the da vinci system. The urologist examined the patient's pelvis through a small amount of dissection in the camper's fascia. The small amount of dissection revealed the transected left ureter and charring was observed at the end of the ureter. The urologist repaired the ureteral injury by directly implanting the ureter into the bladder. He also performed ureteral neocystostomy, ureteral stent placement, and bladder closure. The patient reportedly tolerated the procedure very well and was transferred to the recovery room. The patient was discharged 3 days after the surgical procedure. The patient was advised to have a foley bag for 2-4 weeks. The operative reports do not have any allegation of a malfunction of a da vinci system, instrument or accessory. According to the patient's medical records, the patient had her ureteral stent removed in (b)(6) 2009. During a doctor's visit on (b)(6) 2011, the patient reported she had been experiencing chronic pain on the left flank, left abdomen, left leg, and back. The patient stated that this pain started since her hysterectomy procedure on (b)(6) 2009. During the examination, the surgeon noted that the patient did not have any gross edema or any evidence of acute injury. The patient was advised to continue the medication pain management and physical therapy. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. The patient's current condition is not known at this time. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci s surgical hysterectomy procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334740 | 3305573 | 3202168 | 2955842201303145 | 1 | NAY | Manufacturer | OtherConsumer | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS 1200 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-05-02 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(6) 2013, intuitive surgical, inc. (isi) received information regarding a patient that underwent a da vinci si hysterectomy procedure on (b)(6) 2012. The information provided states that a few weeks post op, the patient developed significant urinary incontinence. A ct scan showed that the patient had right hydronephrosis and hydroureter down to the pelvis without discreet fluid collection to the vaginal wall. The fluid was found to be consistent with urine. On (b)(6) 2012, the patient underwent a 2nd da vinci surgical procedure to repair her ureter and a cystoscopy with bilateral retrograde pyelograms. Ureteral stents were placed to repair the patient's ureter. However, the stent placement was unsuccessful due to ureteral obstruction and a possible ureterovaginal fistula. A right nephrostomy tube was then placed. The patient tolerated the procedure well and after completion of the procedure the patient was extubated and taken to the post-anesthesia care unit for routine post-surgical care. The patient was transferred to the floor after being stabilized. The patient did well postoperatively. The patient voided spontaneously, had adequate pain control with pain medication, and ambulated without difficulty. The patient was discharged from the hospital on (b)(6) 2012. | Based in the information provided, intuitive surgical has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however, as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained a ureter injury during a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334763 | 3305584 | 3202135 | 2955842201303149 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | No | No | Reuse | 2013
IS3000 A60P8 |  | Injury | Required InterventionHospitalizationDisability | As part of a legal mediation effort on (b)(6) 2013 intuitive surgical received a legal complaint regarding a patient who underwent a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure on an unknown date. The legal complaint alleges that during the surgical procedure, the patient sustained bowel perforation. As a result of her bowel injury the patient developed sepsis, experienced septic shock, renal failure, hypoxic brain injury, thyroid injury, gallbladder damage, cognitive deficits, and urinary and bowel incontinence. Per the information in the legal complaint, the patient has suffered severe and permanent injuries including severe internal injuries, pain, disability, physical impairment, mental impairment, disfigurement, mental anguish, inconvenience and loss of capacity for the enjoyment of life, all of which are continuing and are permanent in nature. No other information was provided. | Based on the limited information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the hospital's risk management department. The director of risk management indicated that due to legal constraints she was unable to provide information regarding the patient or the reported event. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained a bowel perforation during a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334773 | 3305585 | 3202180 | 2955842201303150 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-08-21 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical received information including medical records, of a patient who developed post- surgical complications after having a da vinci si hysterectomy procedure on (b)(6) 2012 to treat endometriosis. The operative report does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. The patient's procedure went as planned without any complications. The patient was discharged and was readmitted to the hospital on (b)(6 )2012 with lower abdominal pain, chills, fever and bloody discharge. Test results revealed that the patient had a very high white blood cell count indicative of infection. The patent's pelvic exam demonstrated a slight bloody discharge from the vagina, no palpable pelvic mass was detected. No clear signs of abcess and infection were reported. The patient was started on iv antibiotics and remained in the hospital for five days. She was discharged home on oral augmentin and flagyl. On (b)(6) 2012 the patient developed a fever of 102. The patient was admitted in the hospital for further management. At the hospital she stayed afebrile, and had leukocytosis of 19. 5 on admission. The patient had a ct scan of her abdomen and pelvis done which showed marked thickening of the right sided colon as well as transverse colon. It was noted that there was also irregular fluid collection above the vaginal cuff. The patient was started on iv ceftazidime, flagyl and vancomycin, an infectious disease consult was called to speak to the patient with management of the infection and antibiotics. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical hysterectomy procedure. | DA VINCI SI SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334774 | 3305587 | 3202182 | 2955842201303148 | 1 | NAY | Manufacturer | OtherConsumer | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2010-10-07 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(6) 2013 intuitive surgical, inc. (isi) received information, including medical records, of a patient who underwent a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure on (b)(6) 2010 for treatment of chronic pelvic pain and abnormal uterine bleeding. The patient reportedly had a small perforation of the bladder during the procedure, which the surgeon repaired. The surgeon repaired the bladder injury but also requested an intraoperative consult from another surgeon to perform a cystoscopy to ensure there are no other abnormalities or injuries and to assess drainage from the ureteral orifices. The bladder was found to be fully distended and there was no evidence of significant extravasation of urine or leakage. The patient was taken to recovery in stable condition. The operative reports did not contain any report of a malfunction of the da vinci si system, instruments, or accessories that caused or contributed to the bladder injury. Ten days after the surgical procedure ((b)(6) 2010), the patient had her followup appointment to remove the foley catheter. During this visit, she had complaints of a malodorous vaginal discharge. The patient was admitted to the hospital because there was a concern that urine was leaking from the foley catheter. The surgeon suspected a communication between the vagina and the bladder. The following day ((b)(6) 2010), a pelvic exam was performed. There was some evidence of wound dehiscence at the top of the vaginal cuff and there was probably urine within the vagina. A ct urogram and ct pelvis were performed on (b)(6) 2010, which confirmed a vesicovaginal fistula with inflammatory changes around the urinary bladder and vagina, and sigmoid diverticulosis. No discharge summary is available for any additional information. On (b)(6) 2010, the patient was seen by the urologist who had performed the cystoscopy on (b)(6) 2010. The patient was discharged with a foley catheter and returned back on (b)(6) 2010. At this time a repeated cystoscopy demonstrated a persistent defect at the vaginal cuff and fistula. The cystoscopy demonstrated foreign bodies in the bladder consistent with suture material. Two pieces were removed from the fistula track; however, the urologist recommended another cystoscopy under anesthesia. Based on the patient's medical records, the patient continued to be treated from (b)(6) 2010 through (b)(6) 2011 for cystograms and vesicovaginal fistula repairs. The patient's current condition was not provided. A retrograde exam performed on (b)(6) 2011 showed no evidence of fistula, however, on (b)(6) 2011 the patient reported incontinence. The vaginal exam did not show any evidence of fistula. Kegel exercises was recommended and the patient was required to report back in 6 months. No additional medical records related to the patients current medical condition has been provided. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. Isi made an attempt to contact the site to obtain more information about the reported event but has not received any additional information. If additional information is received a follow up medwatch report will be submitted to the fda. No previous complaint was reported relating to this event. Based on the provided information, there was no evidence that the da vinci si system, instruments and accessories malfunctioned. However, this complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334768 | 3305592 | 3202187 | 2955842201303151 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS2000 A51P7 | 2012-08-07 | Injury | Required Intervention | As part of a legal mediation effort, on july 25th, 2013, intuitive surgical, inc. (isi) received information including medical records of a patient that underwent a da vinci s radical prostatectomy with bilateral pelvic lymph node dissection procedure on (b)(6) 2012. The patient had a history of psa with subsequent diagnosis of adenocarcinoma of the prostate. During the patient's post-operative recovery, it was noted that the patient had rectal bleeding. A rectal examination and proctoscopy showed that the patient had sustained a small rectotomy. The patient was immediately taken back into the or to repair the injury. Isi's review of the patient's operative (op) report related to the patient's da vinci s prostatectomy procedure found no anomalies occurred during the surgical procedure. There was no report that a malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure. A general surgeon was consulted to repair the damage to the patient's rectum. The damage to the patient's rectum was repaired using the da vinci surgical system. A two layered closure was performed to repair the affected area. The patient's rectum was inspected after the repair was completed. No other damage was found. After the surgery was completed, the patient was awakened from anesthesia and transferred to the recovery room in stable condition. Based on the information indicated in the op report, the surgical procedure was successfully completed. The patient's discharge summary indicated that the patient was discharged from the hospital on (b)(6) 2012. Prior to the patient's release, the patient was ambulating, tolerating an oral diet, had positive bowel movements and flatus and his pain was well controlled. The patient was hemodynamically stable, afebrile and his labs were within normal range. The patient was discharged home without complication. | Based on the information provided, operative, discharge and follow up notes, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda the operative reports do not contain any report of a malfunction of a da vinci system, instrument or accessory. No previous complaint was reported relating to this event. Baed on the information provided it was determined that there was no direct correlation to the da vinci s prostatectomy procedure and the post surgical complications experienced by this patient. This complaint is being reported due to the following conclusion: the patient sustained a rectal injury during a da vinci s prostatectomy procedure. Review of the site's system logs for the reported procedure date of (b)(6) 2012 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334783 | 3305597 | 3202192 | 2955842201303153 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS2000 A51P7 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-10-03 | Injury | Hospitalization | As part of a legal mediation effort on (b)(4) 2013 intuitive surgical, inc. Received information including medical records of a patient who experienced post-surgical complications after undergoing a da vinci si ovarian cystectomy procedure (b)(6) 2011. On (b)(6) 2011 the patient returned to the hospital emergency room complaining of severe abdominal pain in the periumbilical area. During a ct scan it was discovered that the patient had an incarcerated hernia. On (b)(6) 2011, the patient underwent a surgical procedure to repair the hernia. The patient's operative (op) report for the surgical procedure of (b)(6) 2011 alleged that during the da vinci si ovarian cystectomy procedure, it was noted that the patient side manipulator arms on the patient side cart collided a lot while docked to a single port site. To resolve the issue, the surgeon made the decision to create additional port sites to perform the procedure using multiple ports. There were no other system issues documented in the op report. The surgical procedure was successfully completed. The patient tolerated the procedure well and was transferred to the recovery room in stable condition. A ct scan of the patient's abdomen and pelvis on (b)(6) 2011 showed that the patient had a midline ventral hernia, just superior to the umbilicus. The hernia contained omental fat that was mildly edematous. Subcutaneous edema was also noted to be tracking along the patient lower left anterior abdominal wall. No focal fluid collection or abscess was discovered. A 3 cm cystic structure was also noted in the right adnexa, possibly representing an ovarian cyst. During the surgical procedure to repair the patient's hernia on (b)(6) 2011, the surgeon identified a large defect to the left of the umbilicus as well as above the umbilicus which contained incarcerated tissue. The tissue was stuck and adherent and was coming out through the defect. The incarcerated tissue consisted of omentus and there was some semblance of a sac forming in this area. Examination of the omentum showed strangulated omentum in several areas. The strangulated tissue was transected until viable omentu remained. A separate umbilical hernia defect measuring 1 cm in size was discovered. There were other small defects also found in the anterior rectus sheath and the surgeon stated that perhaps these occurred during replacement of a needle or trocar. The peripheral defects and hernia were repaired with sutures. Due to the patient's obesity, her incisional hernia was reinforced with mesh to reduce the risk of recurrence. The patient's umbilicus was excised and mesh was placed to cover the entire affected area. An umbilectomy was also performed, as the blood supply to the patient's umbilicus was suspected. The surgical procedure was completed. Based on the information in the op report, the patient tolerated the procedure well and did not experience any complications. The patient was discharged from the hospital on (b)(6) 2011. The patient's medical records indicated that during a post-surgical follow-up examination on (b)(6) 2011 the patient was found to be doing well and has not experienced any complications or infections. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The operative reports do not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2011 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si ovarian cyst procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334794 | 3305606 | 2245206 | 2955842201303152 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-09-14 | Injury | Required Intervention | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical, inc. (isi) received information including medical records of a patient that underwent a da vinci si right salpingo-oophorectomy procedure on (b)(6) 2011. Per the records received the surgery was unremarkable. Immediately following the initial surgery, the patient was brought back to the operating room for hypotension and was noted to have a large amount of blood in the upper abdomen and pelvis. The patient underwent a diagnostic laparoscopy which was converted to an exploratory laparotomy with no clear finding in regard to the absolute source of the bleeding. The patient went to the recovery room in a stable condition. The patient recovered well and was discharged on day 5 ((b)(6) 2011) the patient was then seen and readmitted on (b)(6) 2011, with a peri-hepatic abscess. The patient required antibiotics and a four day hospitalization, and underwent a ct guided interventional radiology drainage of the abscess. Per the initial operative report for the (b)(6) 2011 procedure, the patient tolerated the oophorectomy procedure well. The patient was taken to the recovery room reportedly in stable condition. Per the discharge summary completed at the hospital, the patient underwent an uneventful robotic-assisted right salpingo oophorectomy. However, in the postop period, she was noted to have hypotension and severe abdominal pain. There was concern for an intra-abdominal bleed. A decision was made to proceed with a ct scan to evaluate the abdomen and pelvis. Ct scan of the abdomen revealed intra-abdominal blood. Patient was taken back to the operating room, where a diagnostic laparoscopy was done initially; however, the bleeder was unable to be isolated, so the patient was converted to an intra-abdominal procedure, exploratory laparotomy. The suspected area of bleeding was ligated; however, no definite area was found to be bleeding. The patient post-operative period was unremarkable. She did receive a transfusion of packed red blood cells, which she tolerated well. Her diet was advanced slowly due to multiple procedures and intra-abdominal blood. However, her diet was progressed and she ambulated and she had bowel movements, was passing flatus and her pain was well controlled. She was discharged home on post-operative day number 5, staples were removed. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2011 found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334787 | 3305610 | 3202205 | 2955842201303154 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P2 | 2011-12-16 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(6) 2013, intuitive surgical, inc. (isi) received information regarding a patient that underwent a da vinci si hysterectomy procedure on (b)(6) 2011 and was found to have a ureter injury 14 days later. Note that all of the information received is the patient's statement and no medical records were provided. The patient was diagnosed with uterine fibroids, heavy menstrual bleeding, and pelvic pain. The da vinci si hysterectomy procedure was completed as planned and the patient described her recovery from the procedure as better than expected. On (b)(6) 2011 the patient experienced a constant flow of urine that leaked from her vagina uncontrollably and at all times. This was stated by the patient to be due to a thermal burn sustained on the lower portion of the patient's right ureter. The burn was alleged to have created a uterovaginal fistula as diagnosed by the operating surgeon. The patient states that she required three additional hospitalizations with three additional unspecified surgeries during a six month course of treatment. Her treatment included the placement and replacement of long ureteral stents which caused her physical discomfort and visible blood in her urine. The patient states that the long term stenting resulted in a low tolerance for standing or walking as it caused pain and discomfort. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained a thermal burn on the lower portion of her right ureter during a da vinci si hysterectomy procedure. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334789 | 3305611 | 3202206 | 2955842201303156 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P2 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-01-09 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical, inc. (isi) received information regarding a patient that underwent a da vinci si hysterectomy procedure on (b)(6) 2013. The information alleges that during the surgical procedure, the patient's sigmoid colon was damaged per the on (b)(6) 2013, per the information received. The patient visited the hospital emergency room complaining of abdominal pain. The patient underwent a diagnostic laparoscopy and exploratory laparotomy on (b)(6) 2013 to repair damage to the sigmoid colon. The patient was discharged from the hospital on (b)(6) 2013. On (b)(6) 2013 the patient returned to the hospital with excessive swelling to her legs. Her physicians reportedly suspected deep vein thrombus; however, that was ruled out and the patient was discharged from the hospital. On (b)(6) 2013 the patient was admitted into the hospital with a c-diff infection. No other information was provided. | Based on the information provided, intuitive surgical has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2013 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient allegedly sustained an injury during a da vinci surgical hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334805 | 3305625 | 3202220 | 2955842201303155 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
420021 | 2011-12-21 | Injury | Required Intervention | As part of a legal mediation effort, on (b)(4) 2013 intuitive surgical, inc. Received information including the medical records of a patient who was about to undergo a da vinci si hysterectomy and bilateral salpingo-oophorectomy, when the patient's aorta was damaged during placement of a 12 mm trocar for the camera port on (b)(6) 2011. The patient's operative report (op) indicated that the surgeon immediately made the decision to start an exploratory laparotomy, which revealed that the patient's aorta was damaged. A vascular surgeon was consulted for evaluation and repair of the patient vessel. After the repair was completed, the procedure was completed using open surgical technique. There was no statement in the op report that a malfunction of the da vinci surgical system, instruments or accessories occurred that caused or contributed to the injury sustained by the patient. The open procedure was completed and there were no other complications encountered. The patient was transferred to the intensive care unit in stable condition. The op report concerning the repair of the patient's iliac indicated that the surgeon repaired the damage to the patient's vessel using suture and after the closure there was no bleeding. Reportedly, a complete inspection of all other areas was performed and there was no evidence of significant bleeding. Per the information indicated in the patient's consultation report, during the patient's post-operative care the patient's hemoglobin was found to be 3. 5. Reportedly the family declined to undergo a blood transfusion due to religious beliefs. It was reported in the patient's discharge summary that during the patient's postpartum hospital course, 5 days post op the patient continued to improve. Her workup for possible infections were all negative, which included continued monitoring of her temperature, physical exam and ct scan of the abdomen and pelvis. The patient was discharged from the hospital on (b)(6) 2011 in stable condition. | Based on the information as indicated in the patient's op report, it has been concluded that the injury sustained by the patient was not due to a malfunction of the da vinci surgical system, instrument or accessories, but was rather caused by the surgeon's technique while attempting to create the camera port site incision. If additional information is received regarding this report, a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: during creation of a port site while placing a 12 mm trocar, the patient's aorta was damaged. | 12 MM BLUNT OBTURATOR | ENDOSCOPIC ACCESSORY | 3334792 | 3305630 | 3202225 | 2955842201303157 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | 420021 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
| 2010-02-26 | Injury | HospitalizationLife ThreateningRequired Intervention | As part of a legal mediation effort, on (b)(4) 2013 intuitive surgical inc. (isi) received information concerning a patient who underwent a da vinci hysterectomy procedure on (b)(6) 2010. The information received states that the patient sustained a thermal monopolar injury to her ureter during the surgical procedure which caused documented extravasation of urine into the peritoneal cavity and that this was directly causative to the fluid collection and increased pelvic pain that was then misdiagnosed as an ovarian cyst. The information states that the patient underwent surgical castration as a direct result of a complication from her da vinci hysterectomy. The documentation states that at the time of the surgical procedure, the patient was found to have a small uterus, normal ovaries bilaterally, and no evidence of extra-cervical disease. During post operative follow up the patient experienced shoulder pain and received physical therapy for a diagnosis of compression from surgical positioning. The patient received non steroid anti-inflammatory treatment. The patient was discharged from the hospital on (b)(6) 2010. On (b)(6) 2010, the documentation states that the patient returned to the hospital complaining of severe abdominal pain and significant vaginal drainage. The patient was evaluated by urology service and underwent cystoscopy, cystogram under anesthesia, bilateral retrograde pyelogram, and right ureteral stent placement. The surgery was stated to show right ureteral extravasation. In (b)(6) 2011 the patient was evaluated for pelvic pain and found to have a large left sided pelvic cystic mass, which was felt to be consistent with an ovarian cyst. After receiving antibiotics, the patient underwent a da vinci bilateral salpingoophorectomy procedure and the opening a fluid collection in the pelvis. It was indicated that the patient had increasing fever and leukocytosis and developed an ileus and partial small obstruction, which required an extended hospitalization. The patient was discharged with on oral antibiotics from the hospital on (b)(6) 2011. The information provided states that on (b)(6) 2011, the patient underwent a ct scan. The ct scan showed a marked inflammation and thickening of the sigmoid colon and an inflammatory process involving the bowel. A radiologist reviewed the films from 2010 and noted that the patient had multiple fluid collections in the pelvis which may have been related to urinary injury. Infection of these collections was suspected and inflammation of the small bowel and colon may be secondary. The patient admitted and treated with antibiotics for 7 additional days. She had a pic line placed and received home infusions in (b)(6) 2011. | Based on the limited information and lack of any additional clinical information from the hospital or surgeon, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. A follow-up mdr will be submitted if additional information is received. The documentation provided does not have any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained a ureteral injury during a da vinci hysterectomy procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334833 | 3305632 | 3202227 | 2955842201303158 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional |  | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2011-04-21 | Injury | HospitalizationLife ThreateningRequired Intervention | As part of a legal mediation effort, on (b)(6) 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si prostatectomy procedure on (b)(6) 2011. The provided document stated that the patient sustained perforations of the small bowel during the surgical procedure. The patient's operative report, dated (b)(6) 2011, noted that the patient had significant comorbidities and prior surgeries of exploratory laparotomy, colectomy, colostomy and closure of colostomy. The surgeon included in the operative report that the patient had severe adhesions particularly on the left side where the colostomy was present. The robotic ports were placed in a specific way to take down the adhesions. The provided document stated that with minimal cautery the adhesions were taken down in order for the robotic ports to all be placed, except for the one on the left hand side. The surgeon was able to perform lysis of the adhesions to gain adequate exposure of the prostate gland. The left sided adhesions were not fully taken down. The remainder of the procedure was completed successfully without any difficulties. The operative note indicated that the patient's bowel was carefully inspected for injury and there was one area of perforation noted in the colon area, which was closed with sutures. The anastomosis was tested for water tightness and there were no complications noted. The patient medical records indicate that on (b)(6) 2011, the patient was evaluated in an emergency room (er) for reported complaints of worsening nausea and vomiting. A chest x-ray reportedly revealed a significant amount of air under his diaphragm. The patient underwent an exploratory laparotomy and numerous adhesions in both the right and left lower quadrants of his abdomen were observed which required about 30-45 minutes for lysis. During the procedure, the patient's operative report indicated that the surgeon discovered 2 enterotomies in the small bowel that were found in the left lower quadrant with abscess activity. The small bowel was resected and the anastomosis was stapled. During the procedure, the patient's bladder was noted to be adhered to the bowel and this was taken down. At the conclusion of the procedure, urology service came into the operating room to assist with the infusion of 120 ml of saline into the patient's foley catheter and bladder. No evidence of extravasation was found and the patient was transferred to the icu in stable condition. No additional discharge note information is available. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. Isi has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. The operative notes provided with the legal complaint did not indicate that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the da vinci prostatectomy procedure. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2011. No system errors were found to have occurred during the reported da vinci si surgical procedure. This complaint is being reported due to the following conclusion: the legal complaint alleged that the patient sustained injuries to his small bowel as a result of undergoing a da vinci surgical procedure. However, at this time, it is unknown how the patient's bowel injuries occurred and if the da vinci surgical system caused or contributed to the patient's injury. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334835 | 3305634 | 3202229 | 2955842201303159 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS2000 | 2009-02-17 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on july 25th, 2013 intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci s hysterectomy procedure on (b)(6) 2009. The patient's medical history summary was provided. The information received stated that during the surgical procedure, a failure of the three-dimensional optical system occurred while the surgeon was suturing the vaginal cuff and the patient sustained a vaginal cuff dehiscence. The patient's medical records indicate that the patient underwent the da vinci si hysterectomy on (b)(6) 2009 for chronic abdominal pain, chronic menometrorrhagia, chronic pelvic pain, dysmenorrhea, and symptomatic leiomyoma. The medical summary indicates that equipment failure occurred presumed digitizer on the left 3d eye system on the console at the point of suturing. Both 2d and 3d visual systems were then utilized to complete the case in a very safe manner without complications to the patient. The patient experienced post-operative nausea for 2 days before she was discharged from the hospital on (b)(6) 2009. The patient indicated that she was sensitive to narcotic medications and that they produce nausea. By the second day post op the patient reached almost complete resolution of nausea, the patient was ambulating, tolerating diet, and was discharged home. The patient's medical records indicate that the patient missed her six-week post-operative appointment. However, on (b)(6) 2009, the patient reportedly developed sudden onset right lower quadrant pain while walking. The patient was seen in an emergency room (er) and admitted to a hospital on (b)(6) 2009. On (b)(6) 2009, the patient underwent a proctoscopy and exploratory laparotomy for closure of vaginal cuff and drainage. During the surgical procedure it was noted that there was a wide opened hole in the vaginal cuff that evidently had been plugged by omentum initially. The operative report states that once the omentum had been removed and the defect exposed then irrigated the pelvis out thoroughly, made sure that the bladder flap were pushed anteriorly and the vaginal cuff with five 1 vicryl pop off sutures. The medical summary noted that the surgeon made a nice repair with no obvious visceral injury. The medical history states that blood loss was minimal and the patient was returned to her room in stable condition at the conclusion of the surgical procedure. The patient did well post operatively and was discharged on (b)(6) 2009. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Based on the operative report, it is unknown at what point the equipment failure occurred during the surgical procedure. It is unknown if the equipment failure occurred during suturing of the vaginal cuff or at the conclusion of the surgical procedure when sutures were used to close the surgiports. This complaint is being reported due to the following conclusion: the patient sustained vaginal cuff dehiscence after undergoing a da vinci si hysterectomy procedure. However, at this time, it is unknown if the da vinci si surgical system caused or contributed to the patient's injury. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334826 | 3305640 | 3202235 | 2955842201303160 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | Yes | Health Professional | IS2000 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2012-12-03 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013 intuitive surgical received information including medical records of a patient who underwent a da vinci s lysis of adhesions, bilateral salpingo-oophorectomy, appendectomy, cystoscopy and cervical strumectomy which were performed on (b)(6) 2012. Per the operative reports, dense abdominal pelvic adhesions, consistent with the patient's history of prior surgeries, laparoscopic banding and endometriosis. There was a dense adhesion surrounding the lap band from the omentum and the small bowel to the anterior abdominal wall requiring approximately an hour and a half of adhesiolysis to normalize the anatomy. The same was found in the pelvis with the ovaries. After the procedure was completed, the patient was taken to the recovery room in stable condition. The op report did not allege that a malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure and there were no anomalies noted. Per the information in the op report, the surgical procedure was completed and the patient tolerated the procedures well. The surgeon's post-operative diagnosis were pelvic adhesive disease, abdominal adhesive disease, and endometriosis. On (b)(6) 2012, a ct abdomen and pelvis revealed extensive pneumoperitoneum and a large pelvic collection suspicious for an abscess. On (b)(6) 2013, she was taken back to the operating room and underwent repair of a perforated ileum and lysis of adhesions. The patient had a complicated post-operative course, including icu stay for one month, remaining intubated (followed by tracheotomy placement), anemia, sepsis, and pneumonia. She was discharged stable. Per the medical record for the post-operative follow up visit on (b)(6) 2013 the patient reported that she had improved and was doing better. | Based on the limited information provided, intuitive surgical has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. On (b)(4) 2013 intuitive surgical contacted the risk management department at the hospital to gather additional information about this case, but they were unable to provide any information regarding this case. Risk management indicated that she would have to follow up with the surgeon and they would contact intuitive surgical if any information becomes available. Review of the site's system logs for the reported procedure date of (b)(6) 2012 found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained an injury during a da vinci s bilateral salpingo-oophorectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334866 | 3305656 | 3202251 | 2955842201303161 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-05-12 | Injury | Required InterventionLife Threatening | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical inc. (isi) received information concerning a patient who underwent a da vinci si total laparoscopic hysterectomy with left salpingo-oophorectomy and right salpingectomy and diagnostic cystoscopy procedure on (b)(6) 2011. According to the information provided, the patient was found to have bowel adhesions to the anterior uterine surface and omental adhesions to the anterior abdominal wall during the surgical procedure performed on (b)(6) 2011. The operative report states that great use of cautery was used during the procedure. A diagnostic cystoscopy was performed and found to be unremarkable. The diagnostic cystoscopy also revealed good flow from both ureteral orifices. The patient was reportedly discharged on an unspecified date after an unremarkable stay. On an unspecified date in late (b)(6) 2011, the information indicated that the patient began experiencing brown vaginal discharge as well as flatulence from her vagina. Based on visits with a physician in late (b)(6) 2011, the physician noted that the patient was having intercourse with her spouse. However, the patient adamantly denied that this was the case and she recalled having sexual intercourse with her husband near his birthday in late (b)(6) 2011. The physician referred the patient to a colo-rectal surgeon. On (b)(6) 2011, the patient was admitted to a hospital and underwent a diagnostic laparoscopy which was converted to an open exploratory laparotomy with small bowel resection and incidental appendectomy. During the surgery, it was noted that the patient had marked inflammation in the pelvic cavity as well as a hole in the colon, with a fistula communicating from the distal ileum. It was stated that these conditions can be attributed to the use of monopolar energy from the prior surgery. On (b)(6) 2011, the patient had a revision procedure to the vaginal cuff. It was noted that there was no identifiable cuff dehiscence during the surgical procedure. The patient's vaginal cuff was reported reinforced. She was hospitalized for an unspecified duration. On (b)(6) 2012, the patient underwent a colonoscopy procedure and a digital exam of the vagina revealed discharge with some stool at fingertip. It was indicated that the procedure found some connection between the bowel and vagina. On (b)(6) 2012, the patient underwent an unspecified second procedure for further evaluation. No clinical information was provided regarding the second procedure. The information provided states that her symptoms had improved but she was still not back to her normal state. | Based on the limited information available, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not have any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however, as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2011. No system errors were found to have occurred during the surgical procedure. This complaint is being reported due to the following conclusion: the patient sustained an injury to her colon as a result of undergoing a da vinci surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334868 | 3305659 | 3202254 | 2955842201303162 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS1200 A43P9 | 2011-08-02 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci partial nephrectomy procedure on (b)(6) 2011. The documentation provided states that the patient developed a hematoma after undergoing the surgical procedure. Approximately one month prior, this patient had undergone a cholecystectomy procedure with no complication. No intra-operative complications were noted in the operative report of the da vinci partial nephrectomy procedure performed on (b)(6) 2011. At the conclusion of the surgical procedure, the patient's peritoneal contents were inspected and no injuries or defects were found. The patient was reportedly sent to the recovery room in good condition. During the course of the patient's hospitalization, the patient had a reported complaint of retrosternal discomfort on (b)(6) 2011. The patient was evaluated and found to have a non-q myocardial infarction. On (b)(6) 2011, the patient underwent a left heart cardiac catheterization procedure with selective right and left coronary angiography and left ventriculography. According to the patient's medical records, the patient tolerated the procedure well and was discharged from the hospital the same day. On (b)(6) 2011, the patient returned to the hospital with reported complaints of extreme anxiety. The patient's medical records indicate that she was managed medically and she improved. The patient was discharged home that same day without any complications. On (b)(6) 2011, the patient was hospitalized for lower abdominal pain and an elevated white blood cell count. The patient reported a one-week history of chills, generalized malaise, nausea without vomiting, abdominal pain that was worse on the left side, and frequent stools. A ct scan performed on the patient showed what appeared to be a large left perinephric abscess. On (b)(6) 2011, the patient underwent ct guided drainage of the suspected abdominal abscess. During the procedure, the patient was found to have an infected hematoma rather than a diverticular abscess. A percutaneous catheter was placed for drainage and she was discharged home that same day on oral antibiotics. On (b)(6) 2012, the patient was evaluated for microscopic hematuria. The patient underwent a cystoscopy procedure. The patient was found to have negative hematuria. The physician stated in the patient's medical records that she was doing reasonably well. No additional information was provided. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient developed a hematoma as a result of undergoing a da vinci partial nephrectomy procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334869 | 3305660 | 3202255 | 2955842201303163 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2009-02-11 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(6) 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci s prostatectomy procedure on (b)(6) 2009. The documentation provided states that the patient developed an abscess as a result of undergoing the surgical procedure and was hospitalized. The patient's operative report for (b)(6) 2009, indicated that he underwent a radical prostatectomy, followed by a bilateral pelvic lymph node dissection. The patient had been diagnosed with adenocarcinoma of the prostate. No intra-operative complications were reported on the operative report. The operative report also noted that the patient tolerated the procedure well. The patient was discharged home the following day on (b)(6) 2009, and was reported to be doing fine. The discharge summary indicated that the patient was afebrile, his urine was clear, with minimal drainage. He was discharged with a foley catheter. On (b)(6) 2009, the patient had a post-operative follow-up appointment with the surgeon. The patient's catheter was removed. The patient's port sites looked good. During a follow up visit on (b)(6) 2009 the patient reported dysurial pain in his right testicle. The patient was placed on oral antibiotics and vicodin for pain control. During a follow up visit on (b)(6) 2009, a ct scan was performed and indicated a left inguinal mass causing compression of localized venous vasculature with edema in the lower extremities. The patient was admitted to the hospital the following day. On (b)(6) 2009 a drain was placed to remove pelvic fluid collection that was causing lymphatic and renal obstruction. The swelling improved overnight and the patient was discharged the next day. A follow-up appointment on (b)(6) 2009, revealed that the patient was healing well from surgery and he did not have urinary incontinence. On (b)(6) 2009, the patient was evaluated in an emergency room (er) and found to have a recurrent lymphocele. The patient was treated for nausea and pain. The patient reportedly declined hospitalization. On (b)(6) 2009, the patient had another follow-up appointment with the surgeon who performed the da vinci prostatectomy procedure. According to the patient's medical records, he was down to putting out about 200 cc every 2 hours out of his jp drain. The medical record also noted that when the patient pushes betadine in, he can only get a few cc before it starts to drain out around, so the space is probably almost completely obliterated at this point. The patient was scheduled to see a radiologist in 3 days. No further clinical information was provided. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2011. No system errors were found to have occurred during the reported da vinci si surgical procedure. This complaint is being reported due to the following conclusion: the patient sustained post-surgical complications as a result of undergoing a da vinci prostatectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334881 | 3305666 | 3202261 | 2955842201303164 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2010-12-15 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(6) 2013 intuitive surgical received information including medical records of a patient whose bladder was reportedly injured during a da vinci si hysterectomy and bilateral oophorectomy on (b)(6) 2010. The inadvertent damage to the patient's bladder was repaired and the surgical procedure was completed. The operative report for the patient's for the surgery date of (b)(6) 2010 indicated that an inadvertent cystotomy approximately 2 cm was created and was instantly recognized and repaired with sutures. Indigo carmine was administered to the patient and there was no passage of dye into the peritoneal cavity noted after the completing the repair. Reportedly, the vaginal cuff was completed and was oversewn with a good seal. Hemostasis was checked and assured. The patient's ureters were inspected and noted to be in their normal anatomical positions bilaterally. The procedure was completed a leaf of interceed was placed over the vaginal cuff. The patient was closed, extubated and taken to the recovery room in stable condition. There were no other complications noted. The patient's history and physical report stated that during a post-op follow up on (b)(6) 2011, the patient complained that 3 days post op she began to experiencing increased fluid from the vagina when she stands, walks, sneezes or coughs. The patient was concerned that these may be from her bladder. A ct cystogram documented vesicovaginal fistula. The patient's discharge summary stated that this possibly was due to her intervention procedure in (b)(6) 2010. Reportedly, the patient was instructed to maintain her foley catheter for another 4 to 6 weeks and then follow up with the urologist. The patient was discharged from the hospital on (b)(6) 2011. The patient's medical records indicated that during a follow-up examination, the patient reported that she felt that she had an infection. A cystoscopy was performed on the patient and a urinalysis revealed that the patient did not have an infection. Reportedly, the patient underwent a cystoscopy, bilateral retrograde pyelograms, abdominal exploration with lysis of adhesion, closure of vesico-vaginal fistula with omental graft interposition, and sp tube replacement. During the procedure there was no ureteral extravasation, hydronephrosis or filling defect noted. After completion of the procedure good hemostasis was assured. The patient was closed and the patient was transferred to the pacu in stable condition after extubation. The patient was discharged from the hospital on (b)(6) 2011. No additional information was provided since the patient's last discharge. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained an injury during a da vinci si hysterectomy and bilateral oophorectomy and experienced post surgical complications post-operatively. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334915 | 3305671 | 3202266 | 2955842201303167 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2011-12-06 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2011. The documentation received states that the patient developed a ureteral vaginal fistula as a result of the surgical procedure. The legal claims also alleged that the patient sustained kidney damage that was caused by a series of infections she contracted during her post-surgical struggles. Additionally it was stated that the patient had an unremarkable medical history but had been dealing with menorrhagia. Prior to undergoing the hysterectomy, the patient did not have any problems with her kidneys, ureters, or bladder. The operative report indicated that during the da vinci si hysterectomy procedure the patient had extensive lysis of adhesions and bilateral ovarian cystectomy. After the patient's vaginal cuff was closed, the patient was given 5 cc indigo carmine dye intravenously and cystoscopy was then performed. The operative report states that the patient's right ureteral orifice did not have any flow after over 40 minutes of observation. A urologist was then consulted. The urologist observed narrowing of the patient's right ureter and some mild hydroureteronephrosis. The urologist implanted a ureteral stent. The patient tolerated the procedure well. The patient was discharged home from the hospital the following day on (b)(6) 2011. The patient's medical records state that the patient was seen by the surgeon on (b)(6) 2011. The surgeon informed the patient that the right flank pain she had experienced over the previous several months might have be related to possible congenital ureteral stricture. On (b)(6) 2011, the patient was evaluated due to complaints of increasing right flank pain. The patient complained of fatigue, low grade fever, and hematuria starting the day before. Ivp shows a high grade obstruction on the right hand side with stenosis. According to the urologist's notes, this could be a stricture that is blocking this ureter. The patient underwent another cystoscopy in addition to a right ureteroscopy and ureteral dilation on (b)(6) 2011. Another ureteral stent was placed that same day. An x-ray performed that day revealed that the patient had a right ureteral stricture and a questionable ureterovaginal fistula on (b)(6) 2012, the patient came to see the doctor because the stent fell out. The patient underwent another cystoscopy, a right ureteroscopy, a right retrograde and laser of ureterovaginal fistula, and placement of another stent. On (b)(6) 2013 the patient visited the doctor and the stent was removed the same day. On (b)(6) 2013 the patient underwent a right distal ureteral re-implantation using a pso as hitch and mobilizing / pexing of the right ovary. The patient was discharged on (b)(6) 2013. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2011. No system errors were found to have occurred during the reported da vinci si surgical procedure. This complaint is being reported due to the following conclusion: the documentation received alleged that the patient sustained a ureteral vaginal fistula as a result of undergoing a da vinci hysterectomy procedure. However, at this time, it is unknown how the patient's injury occurred and if the da vinci surgical system caused or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334896 | 3305703 | 3202298 | 2955842201303166 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS 3000 A60P8 | 2012-03-12 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(6) 2013 intuitive surgical received information, including medical records, of a patient who developed post- surgical complications after having a da vinci si hysterectomy with left salpingo-oophorectomy procedure on (b)(6) 2012. It was noted that a 360 degree circumferential cutting of the cervix from the vagina was performed using monopolar hot shears. A small laceration was detected at the vaginal cuff; however, the laceration was closed with sutures. The pelvis was irrigated and found to be completely hemostatic before the procedure was completed. The patient was discharged home on post- operative day one. The patient went in to the office for a follow-up visit on (b)(6) 2012 and the abdominal incisions were described as healing well. The patient went to the emergency room on (b)(6) 2012 with complaints of abdominal pain. Patient was examined and a ct scan of the abdomen and pelvis revealed a small fluid collection in the cul-de-sac near the cuff with a fleck of air measuring 2 cm. Differential diagnoses include a small abscess or seroma. The patient subsequently saw another surgeon on (b)(6) 2012 and his examination revealed an extremely tender vaginal cuff, and surgeon admitted her to the hospital with suspected vaginal cuff abscess. Laboratory work during the admission was significant for anemia and leukocytosis (elevated white blood count). The patient was treated with intravenous antibiotics and was discharged on (b)(6) 2012 on oral antibiotics. The patient's discharge diagnosis also included vaginal cuff abscess. As of (b)(6) 2012 the abscess/cellulitis was resolved. The last noted visit was on (b)(6) 2013 at which time estrogen hormone treatment was initiated for peri/post menopausal symptoms. It was also noted that the patient had vulvovaginitis and was treated with an antibiotic. | Based on the information provided, intuitive surgical has not determined the root cause of the post- surgical complications experienced by the patient. If additional information is received a follow -up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however, as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. The patient's condition resolved with antibiotic therapy without the need for any additional procedure. This complaint is being reported due to the following conclusion: the patient suffered an abscess after a da vinci si hysterectomy with left salpingo-oophorectomy procedure. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334904 | 3305718 | 3202313 | 2955842201303168 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS 3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2010-12-22 | Injury | Required InterventionHospitalization | On july 25th, 2013 intuitive surgical received a legal complaint including medical records of a patient who underwent a da vinci si hysterectomy and bilateral salpingo-oophorectomy with pelvic and paraaortic lymph node dissection on (b)(6) 2010. The patient had a diagnosis of grade ii endometrial cancer and fibroid uterus. The surgical procedure was completed as planned with the removal of uterus, tubes and ovaries through vagina with periaortic lymphnode dissection. The vaginal cuff was closed using interrupted stitches without any difficulty. The surgical procedure was successfully completed and the patient was extubated and transferred to the recovery room in stable condition. The patient was discharged from the hospital on (b)(6) 2010 in good condition with a normal hematocrit level without complications as per the medical records. After arriving home the patient began to develop nausea and abdominal pain. This persisted to dehydration and she presented to hospital's er remarkably dehydrated on (b)(6) 2010. A ct-scan performed showed findings consistent with a possible pelvic abscess. She was admitted and underwent a ct guided catheter placement in an attempt to drain the abscesses. The patient was discharged hypotensive with acute abdomen on (b)(6) 2011 and underwent an exploratory laparotomy, with the findings of a 2mm left ileac artery injury. She was also noted to have a cautery effect and injury on the aorta. Both the left iliac artery and aorta were repaired. Secondary to the acute blood loss, the patient was sent back to the icu, where the she remained mechanically ventilated. On (b)(6) 2011, the patient experienced an episode of fairly sudden shock and vascular compromise. She was rushed back to the operating room again and underwent a re-exploration of her abdomen, with findings of an aortic rupture in the same area as the previously noted cautery injury, and was repaired again. She received extensive volumes of fluid, blood, platelets and fresh frozen plasma throughout the course of her hospitalization. She was again maintained on a ventilator following surgery and was followed by the renal service, pulmonary service, and infectious disease service until the point she was finally extubated. She was discharged home on (b)(6) 2011. The patient represented to the hospital on (b)(6) 2011 complaining of lower abdominal pain. A ct-scan revealed an extravasation of arterial contrast at the level of the aortic bifurcation, with possible pseudoaneurysm or arterial perforation. The patient underwent another exploratory laparotomy. She was noted during that surgery to have an issue with the left iliac, close to the area of prior disruption, but not related to a prior repair. Because of the extent of the injury and dense reactive tissue to the aortic and iliac vessels and the amount of blood and inflammatory response, the surgeon decided to do an extra anatomic "axillobifeonoral" bypass. The anatomic bypass involves the ligation of the aorta and the bilateral ileac vessels. A graft had to be taken from her subclavian axillary veins through an incision into her left clavicle area. A graft anastomosis was then performed. The patient was sent back to icu and slowly improved. The patient was discharged from the hospital on (b)(6) 2011. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. There was no allegation in the patient's operative report that a malfunction of the da vinci surgical system, instruments or accessories occurred. No previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2010 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient sustained injuries during a da vinci si surgical procedure resulting in subsequent post-surgical complications. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334940 | 3305757 | 3202352 | 2955842201303169 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | Yes | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
|  | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on july 25, 2013, intuitive surgical inc. (isi) received information concerning a patient who underwent a da vinci hysterectomy procedure on an unspecified date in (b)(6) 2008. The documentation provided alleged that the patient's vaginal wall was burned during her hysterectomy and she suffered a rectal-vaginal fistula resulting in a colostomy, multiple ct scans, several outpatient surgeries and procedures and 2 additional major surgeries with hospital stays. The documentation provided states that the patient had a fibroid growth on her uterus for many years and had for some time been experiencing increasing abdominal pain. The patient claimed that as a result of the burn to her vaginal wall, she developed a vaginal-rectal fistula and stool was coming out of her vagina. The documentation states that the patient had a colostomy performed on an unspecified date followed by an attempt to repair the fistula. In addition, the patient reportedly developed abdominal pain and bloating from the procedures and medications, which caused her to spend a day in the er. This surgery was performed in (b)(6) 2008. According to the allegations, during her stay in the er, ct scans and an outpatient colonoscopy procedure were performed to determine if the fistula was healed enough to be reversed. The documentation stated that when the reversal was done, the surgeon discovered part of the patient's bowel needed to be resected. The date of the bowel resection procedure was not provided. On an unspecified date a few months later, the patient developed an incisional hernia at the site of the colostomy and went in for surgery again on (b)(6) 2008. The documentation also states that the patient has scars that are irritating and she has abdominal pain on the left side that requires medication to make it bearable. The patient's physician informed her that the pain is from adhesions from the multiple surgeries. No additional patient recovery information was provided. | Based on the limited information provided, and in absence of medical / clinical operative, discharge and post operative documentation, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained a burn to her vaginal wall during a da vinci hysterectomy procedure. However, at this time, it is unknown how the patient's injury occurred and if the da vinci surgical system caused or contributed to the patient's injury. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3334952 | 3305772 | 3202367 | 2955842201303170 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional |  | No | No |  | Device Not Retu | No | Unkown | 2013
IS3000 A60P8 | 2011-04-07 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(6) 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2011. The legal complaint alleged that the patient developed a pelvic abscess as a result of undergoing the surgical procedure and was hospitalized. The patient's procedure notes state that the pre-operation diagnosis was symptomatic uterine fibroids, pelvic pain, and menometrorrhagia. The patient was noted to have a relatively large uterus, approximately 16 weeks with one large "funduai" fibroid. Ovaries and tubes reported to be normal. After removal of the uterus, a small vaginal laceration was noted to be repaired with 2. 0 vicryl without incident. The vaginal cuff was closed using the v-loc running suture with reported excellent closure. There were no signs of leakage of urine and no reported evidence of injury. The patient was brought to the recovery room in stable condition. The patient was discharged on (b)(6) 2011 (one day after surgery) with no reported complications. According to the patient's medical records, the patient was presented to the emergency room with symptoms of fever, chills, and increased abdominal discomfort on (b)(6) 2011. A ct scan revealed that the patient had a bowel injury. On (b)(6) 2011, the patient underwent an exploratory laparotomy, abdominal washout and drainage of the vaginal cuff abscess, however, there was no evidence of an injury to her bowel. The patient was hospitalized for 8 days for treatment of the abscess. The patient was discharged home with oral antibiotics on (b)(6) 2011. The patient's discharge summary states that she had a very light amount of wound drainage but no signs of infection. No further clinical information was provided after the date of discharge from the hospital on (b)(6) 2011. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. The operative report does not contain any allegation of a malfunction of a da vinci system, instrument, or accessory. No previous complaint was reported relating to this event. Isi has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. This complaint is being reported due to the following conclusion: the patient had a bowel injury which was discovered after a da vinci hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335026 | 3305823 | 3202418 | 2955842201303173 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS1200 A51P8 | 2011-10-24 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on 07/25/2013, intuitive surgical received information of a patient who underwent a da vinci s hysterectomy with bilateral salpingo-oophorectomy procedure on (b)(6) 2011 for severe cervical dysplasia. This patient had a family history of ovarian cancer. Prior to the procedure, the patient was informed of the surgery risks, including cuff dehiscence, port cuff healing. The patient agreed and signed the consent to procedure and risks thereof. The monopolar scissors instrument was used to transect the vaginal cuff from the cervix and was done in a circumferential manner. The patient's entire cervix, uterus, tubes, and ovaries were removed. The surgeon noted achieving excellent vaginal cuff closure after suturing as well as applying an interceed adhesion barrier to the cuff. The da vinci system was undocked and the procedure switched to traditional laparoscope for the remainder of the procedure. The patient was taken to the recovery room in stable condition. The patient had 3 post-operative check-ups with her doctor from (b)(6) 2011. During these visits, her vaginal cuff was noted to be intact and healing well. Approximately 2 months after the da vinci surgical procedure ((b)(6) 2012), the patient reported a sensation followed by vaginal bleeding during sexual intercourse. She went to the emergency room (er) where she diagnosed with vaginal cuff dehiscence with evisceration. On the same day, the patient underwent an emergency diagnostic laparoscopy and repair of vaginal cuff procedure. The surgeon found a small umbilical hernia, which was repaired. There was no evidence of bowel strangulation or other damage. There were no apparent complications from the procedure. The patient was discharged 2 days later ((b)(6) 2012). As of (b)(6) 2012, the patient's vaginal cuff was healing well and the patient has not reported to have any bleeding. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2011 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci s surgical hysterectomy procedure with bilateral salpingo-oophorectomy. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335045 | 3305838 | 3202433 | 2955842201303174 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS1200 A51P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P2 | 2010-07-16 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on july 25th, 2013, intuitive surgical, inc. (isi) received information including medical records of a patient that underwent a da vinci si hysterectomy procedure with right salpingo-oophorectomy procedure on (b)(6) 2010. The patient was diagnosed with a large 16 week sized fibroid uterus and right ovarian adhesions to the uterus. The da vinci hysterectomy procedure with right salpingo-oophorectomy was completed as planned and the patient was placed in the recovery room in satisfactory condition. The patient's postoperative course was significant for development of anemia, which required a transfusion of two units of packed cells. The patient was discharged as planned. On (b)(6) 2010 the patient was re-admitted to the hospital experiencing a low grade fever, nausea, vomiting, abdominal pain and a feeling of malaise. Upon evaluation the patient was found to have developed a fluid collection in her pelvis and underwent a ct guided percutaneous drainage of the collection on (b)(6) 2010. Approximately 16 ml of dark bloody serosanguinous fluid was removed. The ct scan also revealed that the patient's sigmoid colon has pericolonic inflammation. An x-ray of the patient's abdomen was performed on (b)(6) 2010 to confirm that the inflammation and fluid had ceased. The patient was discharged on (b)(6) 2010 and prescribed antibiotics. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The operative reports do not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained a large pelvic abscess which required her to have additional hospitalized after a da vinci si hysterectomy procedure with right salpingo-oophorectomy. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335092 | 3305860 | 3202455 | 2955842201303176 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P2 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P2 | 2011-01-28 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013 intuitive surgical received information including medical records, of a patient who had a bladder injury during da vinci si hysterectomy and bilateral salpingo-oophorectomy procedure that was performed on (b)(6) 2011. During the procedure, the surgeon noticed a 1. 5-2cm rent (cystotomy) in the posterior aspect of the bladder. The surgeon made the decision to convert the procedure to open surgical procedure in order to get full exposure. The surgeon ordered a urological consultation to address the bladder injury. The consulting surgeon then performed a repair of the bladder rent. The surgeon irrigated the abdomen and hemostasis was assured. A drain was placed in the posterior cul-de-sac. No further complications were noted. The patient was discharged on (b)(6) 2011 with a jp drain for wound drainage and a foley catheter for urinary drainage, both which were removed by (b)(6) 2011. A cystogram on (b)(6) 2011 showed no visualized bladder leak. At the patient's recorded last post-op visit on (b)(6) 2011, the surgeon cleared the patient and stated that no further follow up was necessary. The patient has not had any additional medical treatment. The operative report does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. No further clinical information was provided about the patient's current status. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: during the da vinci abdominal hysterectomy bilateral salpingo-oophorectomy procedure, the patient's bladder was injured. | DA VINCI SI SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335078 | 3305869 | 3202464 | 2955842201303175 | 1 | NAY | Manufacturer | OtherConsumer | Health Professional | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P2 | No | Yes | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-02-01 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical, inc. (isi) received a medical records summary regarding a patient who underwent a da vinci si right salpingo-oophorectomy procedure on (b)(6) 2013. The information received alleges that the patient's operative report described that the patient had a mass that adhered to the iliac vein and during dissection the iliac vein was torn. The procedure was converted to an open laparotomy. A vascular surgeon was consulted to repair the patient's vessel. The medical summary as provided indicates that during the surgical procedure the damage to the patient's vessel was noted during dissection using a harmonic shears instrument. The open laparotomy procedure was successfully completed and the patient tolerated the procedure well. During the patient's hospitalization, the patient was diagnosed with right common femoral deep vein thrombus and was treated with medications to treat her condition. On the 7th day of the patient's hospitalization the patient was anemic and was given two units of red blood cell. On the 11th day the patient's pain was well controlled, tolerating a regular diet, passing flatus and bowel movements, ambulating without difficulty and had improvement in her leg pain. The patient's exertional dyspnea improved and she was deemed stable for discharge. During the patient post-operative follow-up on (b)(6) 2013, the patient indicated that she was doing much better and that her main concern was that she was bleeding from her midline abdominal incision. On (b)(6) 2013, during a routine follow up the patient complained of abdominal discomfort, loss of appetite and nausea. On (b)(6) 2013, the patient underwent an abdominal ct scan and it was negative. During a routine follow-up on (b)(6) 2013, the patient was found to be recovering well. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however, as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(4) 2013, found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. No previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient's sustained an injury during a da vinci si right salpingo-oophorectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335127 | 3305909 | 3202504 | 2955842201303177 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS 2000 A51P8 | 2012-10-31 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on 07/25/2013 intuitive surgical, inc. (isi) received information, including medical records, of a patient who underwent a da vinci s hysterectomy and cystoscopy procedure on (b)(6) 2012 due to a history of dysfunctional uterine bleeding, unresponsive to medical management. One day prior to the procedure ((b)(6) 2012), she was diagnosed with bacterial vaginitis and placed on antibiotics. According to the operative report, there were no reports of complications during the surgical procedure. A cystoscopy was performed and indigo carmine was confirmed to come out both ureteral orifices. The surgeon noted that the patient had a normal pelvis, normal upper abdomen and had normal flow into the bladder of urine after the case from the ureters. The bladder reportedly had good integrity with no leaks. The patient was taken into the recovery room in stable condition. The patient was discharged on (b)(6) 2012, she was afebrile, voiding and tolerating a regular diet. On (b)(6) 2013, the patient went to her doctor's office with complaints of recurrent vaginal discharge and post coital bleeding. On (b)(6) 2013, the vaginal cuff was noted to be not completely healed but it was intact. Later the same day, she went to the emergency room (er) for vaginal bleeding and abdominal cramping. A ct pelvic scan was performed, which revealed that the walls of the vaginal cuff were diffusely mildly thickened and there was a small amount of free fluid within the posterior cul de sac. The patient was treated with medications and was advised to followup with her gynecologist. The patient was sent home the same day. On (b)(6) 2013, the patient visited her doctor for acute onset of pelvic pain and was prescribed medications for a vaginal cuff cellulitis. The patient reportedly did not have any recurrent problems until approximately 4. 5 months after the da vinci procedure. On (b)(6) 2013, the patient went to the er with complaints of abdominal pain, vaginal bleeding, and vomiting. A ct scan of the abdomen and pelvis showed a large amount of free air in the belly, and a suspected vaginal and/or rectal peritoneal fistula. She was taken to the operating room for a suspected vaginal cuff dehiscence. A diagnostic laparoscopy was performed and the vaginal cuff was resected and closed along with lysis of adhesions of the descending colon. She was discharged home on (b)(6) 2013 in stable condition. Follow up in the doctor's office on (b)(6) 2013 internal for hemorrhoids only and her vaginal cuff was noted to be intact. The patient was prescribed a topical anti-inflammatory and antipruritic medication and was asked to return to the clinic as needed. No additional information since the (b)(6) 2013 visit is available. | Based on the information provided, intuitive surgical, inc. (isi) has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The operative reports do not have any allegation of a malfunction of a da vinci system, instrument or accessory. No previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2012 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci s surgical hysterectomy procedure. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335172 | 3305944 | 3202539 | 2955842201303178 | 1 | NAY | Manufacturer | OtherConsumer | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS 2000 A51P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-12-08 | Injury | Required InterventionHospitalization | As part of a legal complaint effort, on (b)(6) 2013 intuitive surgical received information including medical records, of a patient who developed post- surgical complications after having a da vinci si bilateral salpingo-oophorectomy procedure on (b)(6) 2011 for a left ovarian complex cyst. The patient history showed multiple abdominal surgeries in the past and the patient was expected to have a large amount of adhesions in the small and large intestine. The adhesions were removed before beginning the bso procedure. These adhesions were removed with scissors. During the da vinci si procedure bilateral salpingo-oophorectomy procedure it was noted bt the surgeon that extensive omental and colon adhesions were removed with scissors and that a very small rent in the small intestine was likely due to a thermal burn which was then repaired during this procedure and excellent hemostasis was noted. Post- surgery the patient had some abdominal pain and she was discharged on (b)(6) 2011. The following day the patient complained about abdominal pain and was readmitted on (b)(6) 2011 with abdominal pain and fever. She was treated with pain medications and keflex for a probable postoperative ileus and discharged on (b)(6) 2011. On (b)(6) 2011, the patient returned to the emergency room with persistent pain. The patient was evaluated and a ct scan of the abdomen revealed a defect in the right lower quadrant of the abdomen with leaking air and oral contrast, and the patient was taken to the operating room for a distal ileal perforation repair with drainage of anterior abdominal wall abscess. The procedure was completed successfully and the patient was transferred to the recovery room in stable condition. It was noted that the patient's bowel perforation was presumed to be secondary to the patient's surgery 6 days prior. Patient was treated for sepsis and acute renal failure postoperatively and was discharged on (b)(6) 2011. In (b)(6) 2012, the patient had plastic surgery to excise the abdominal scar from the above mentioned surgery, along with lysis of adhesions and repair of an incisional hernia. The operative report does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. No further clinical information has been provided about the patient's current health status. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not have any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Intuitive surgical inc. (isi) has reviewed the system logs for this site with a procedure date of (b)(6) 2011 and confirmed that there were no system errors generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical bilateral salpingo-oophorectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335265 | 3305995 | 3202298 | 2955842201303165 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | Yes | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2010-07-27 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(6) 2013, intuitive surgical inc. (isi) received information including medical records concerning a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2010. The documentation received states that the patient developed a pelvic abscess following a da vinci si hysterectomy procedure and was readmitted to the hospital for further treatment. The operative report for (b)(6) 2010, indicated that the patient underwent a da vinci assisted pelvic hysterectomy in addition to plication of the uterosacral ligaments and drainage of bilateral ovarian cysts. The patient was found to have an enlarged uterus of approximately 18 weeks with abdominal cysts. There was a cyst on both ovaries containing clear liquid. According to the operative report, the patient's large uterus was pulled out through the vagina with some difficulty. The vaginal cuff was closed with vicryl suture. The patient was transferred to the recovery room in stable condition. As per the medical record on (b)(6) 2010, the patient came in an emergency room (er) for reported complaints of nausea, vomiting, and not feeling well. The patient's medical records indicated that she had normal bowel movements since discharge, had minimal vaginal bleeding, and she was able to void without difficulty. A ct scan performed on the patient suggested evidence of some hydronephrosis, more on the right side and fluid collection with in the lower area of the pelvis. The patient was also reported to have an elevated white blood cell count of 24,000. The patient was then admitted to the hospital and treated with iv antibiotic therapy. A stent was also placed to drain the fluid which was negative for any evidence of ureteral transection as her bun and creatinine were equal to the blood levels. The patient was discharged home with oral antibiotics on (b)(6) 2010. On (b)(6) 2010 the patient was discharged, and was reported to be feeling good and she was afebrile and her abdomen was soft and nontender. No further clinical information was provided since the date of discharge on (b)(6) 2010. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient developed a pelvic abscess as a result of undergoing a da vinci hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335394 | 3306088 | 3202683 | 2955842201303181 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 |  | Injury | Required Intervention | As part of a legal mediation effort, on (b)(6) 2013, intuitive surgical, inc. (isi) received information including medical records of a patient that underwent a da vinci si hysterectomy procedure on an unknown date. Per the records, at the time of surgery, she was noted to have adhesions of her small and large intestines, which were taken down with meticulous sharp dissection. Post operatively the patient developed a fever. On (b)(6) 2010, the patient underwent diagnostic laparoscopy which reportedly showed a perforation of the distal ileum and adhesions. The patient underwent laparotomy through a midline abdominal incision and the perforation in the small bowel was repaired. Secondary to adhesions of the ileum to the right ovary and tube, the decision was made to remove the right ovary and tube and a portion of the ileum was removed, and an anastomosis was created. She discharged home on (b)(6) 2010. Per the documentation provided the patient had urinary retention requiring intermittent self catheterization. On (b)(6) 2010, she underwent cystoscopy which shoed the right ureteral orifice and tunnel were somewhat elevated and swollen. No indigo carmine dye could be seen coming from the right ureteral orifice. The patient was admitted on (b)(6) 2010 with a diagnosis of ureterovesical vaginal fistula. She underwent repair of the fistula with right ureteral re-implantation. The patient was discharged on (b)(6) 2010. At the time of discharge, a right ureteral stent had been placed in the right ureter. According to the documentation provided, the patient continued to experience problems relating to a "vesiculovaginal" fistula, which was repaired on (b)(6) 2010. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The operative reports do not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained post operative complications after a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335407 | 3306095 | 3202690 | 2955842201303182 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS 3000 A60P8 | 2011-11-21 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(6) 2013 intuitive surgical received information, including medical records, of a patient had a da vinci si hysterectomy and right salpingo oophorectomy procedure on (b)(6) 2011. No complications were noted during the da vinci si surgical procedure. The operative report does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. The patient came back to the hospital on (b)(6) 2011 with symptoms of abdominal pain and urinary leakage. The patient was evaluated and a ct scan of the abdomen and pelvis showed mild hydronephrosis, with spillage of contrast into the abdomen. The patient had percutaneous stenting of both ureters. There were no complications noted, and the patient was discharged on (b)(6) 2011. Per the medical records provided, the patient went back to the hospital on (b)(6) 2011 with a complaint of leaking from the vagina. On (b)(6) 2011 the patient went in surgery for pelvic exploration. The surgeon performed a bilateral ureteral reimplantation and closure of the left ureterovaginal fistula. A suprapubic tube was placed and the patient tolerated the surgical procedure well. The patient was discharged on (b)(6) 2011 with the suprapubic catheter in place. Per this legal complaint the patient continues to suffer from bladder symptoms and some pelvic pain. No further clinical information was provided. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not have any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Isi has reviewed the system logs for this site with a procedure date of august (b)(6) 2011 and confirmed that there were no system error messages. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical hysterectomy and right salpingo oophorectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335441 | 3306098 | 3202693 | 2955842201303171 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS 3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-04-14 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(6) 2013 intuitive surgical received information including a medical summary regarding a patient who underwent a da vinci si hysterectomy on (b)(6) 2011. The medical history summary states that this patient was planned to undergo a da vinci hysterectomy with planned ureteral stent placement. The patient had evidence of normal urethra and bladder. The ureteral stents were placed successfully by the urologist and foley catheter was placed at the end of the procedure. The patient was re-opened for the da vinci hysterectomy procedure. The medical history summary states that the surgeon who was performing the hysterectomy procedure noted an approximately 1-1. 5cm of bladder cystotomy while performing the anterior colpotomy. The bladder cystotomy was repaired with no further complication. The surgeon made the decision to perform a supracervical hysterectomy to avoid any additional trauma or injury to the patient's bladder. The patient tolerated the procedure well. According to the medical history information provided, prior to the da vinci surgical system being docked to the patient, the patient underwent a planned bladder cystoscopy with placement of bilateral stents, a foley catheter was placed and a uterine manipulator. Traction was encounter while installing the uterine manipulator causing a perforation to the patient's uterus. During the da vinci surgical procedure, the surgeon experienced difficulty dissecting the bladder away from the cervix, and the decision was made to perform a supracervical hysterectomy to avoid any additional trauma or injury to the bladder. The surgical procedure was completed. The patient tolerated the procedure well. Post-operatively the surgeon discussed with the patient and her family that he made the decision to perform a supracervical hysterectomy instead of a total hysterectomy. After the patient was cleared by cardiology following an episode of supraventricular tachycardia, she was discharged from the hospital on about (b)(6) 2011. On (b)(6) 2011 the patient was admitted into the hospital's emergency room with a complaint of abdominal pain and during her hospitalization she experienced an episode of tachycardia. An abdominal ct scan showed that the patient had left-sided hydronephrosis and hydroureter with a marked delayed contrast of the left kidney. The ct scan also showed thickening of the posterior bladder wall and free fluid in the pelvis subcutaneous along the right ventral abdominal wall. The patient's urine was positive for protein, ketones, hematuria to 50 and red blood cells. The patient underwent a cystoscopy, left retrograde pyelogram and stent insertion. A surgeon who had been consulted during the surgical procedure indicated that the repair to the patient's bladder was performed correctly without any leakage and there was no transection or injury to the patient's ureter. It was stated this most likely is secondary to the edema after having initial ureteral stent placement for the da vinci surgery. The patient was scheduled for release from the hospital the next day. No additional records were made available. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however, as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient sustained an injury during a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335495 | 3306187 | 3202782 | 2955842201303184 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-04-13 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013 intuitive surgical received information including medical records of a patient who underwent a da vinci si hysterectomy procedure, a cystoscopy and lysis of adhesion procedure on (b)(6) 2012. The patient was discharged from the hospital on (b)(6) 2012 in stable condition. On (b)(6) 2012, the patient returned to the hospital's emergency room complaining of abdominal pain. The patient's operative (op) report for the surgical procedure dated (b)(6) 2012 indicated that the patient's preoperative diagnosis was chronic pain and abnormal vaginal bleeding. The op report did not allege that a malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure and there were no anomalies noted. Per the information in the op report, the surgical procedure was completed and the patient tolerated the procedure well. The patient was transferred to the recovery room in stable condition. The patient called the doctor's office on (b)(6) 2012 complaining of nausea and vomiting. An e-script for zofran was sent in for her. On (b)(6) 2012, the patient underwent a cystoscopy, bilateral retrograde ureterogram, bilateral double ureteral stent placement and then a robotic-assisted laparoscopic bladder repair and closure. During the cystoscopy procedure it was discovered that the patient had a 1 cm bladder rupture or perforation. Bilateral retrograde ureterograms performed found that the bilateral collecting system appeared to be normal. A robotic-assisted bladder repair was performed and the procedure was completed and the bladder was carefully evaluated which showed no other injury. The procedure was completed and the patient was extubated and transferred to the post anesthesia care unit in stable condition. There were no surgical complications. The op alleges that the patient was discharged in stable condition on (b)(6) 2012. On (b)(6) 2012, during follow up examination, the patient complained of abdominal pain and nausea approximately 30 minutes after eating. An examination of the patient's abdomen found that the patient did not have cva tenderness and that her laparoscopic incisions were healing well. The patient's abdomen was soft and her bowels were active. There were no signs of acute abdomen or rebound. Zofran was prescribed to prevent nausea. On (b)(6) 2012, the patient underwent a follow examination. The patient complained of back pain due to the stents. The patient's abdomen was found to be soft, nontender and nondistended. No organomegaly palpable, no hernias and the incisions were healing well. The patient had no erythema or discharge. A pelvic examination found that the patient's urethra palpates were normal and the vagina was without lesions. Prolapse was noted. The sutures were in place and the vaginal cuff was intact. On (b)(6) 2012 a follow up examination found that the patient's abdomen was soft, nontender and nondistended. The patient had no masses palpable. No organomegaly palpable. No hernias. All of the patients' incisions were found to be healing well and she had no erythema or discharge. The patient's pelvic examination the patient's external urethral meatus appeared normal and the urethra palpates were normal. The patient's vagina did not exhibit lesions and no prolapse was noted. The patient's cervix was surgical absent and the sutures were in place. The patient's vaginal cuff was completely healed and there were no pelvic masses noted. There were no masses in the cul-de-sac. The patient's bladder was nontender. The patient had no inguinal lymphadenopathy. On (b)(6) 2012, the patient underwent a follow-up examination. The patient's pelvic examination found that the patient was healing well. On (b)(6) 2012 during a follow-up examination, she complained of discharge, painful intercourse and dyspareunia, it was discovered that the patient had a little 2 mm granulation tissue. This was treated with silver nitrate. No pelvic vinasses and the bladder was not tender. | Based on the information provided, intuitive surgical has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2012 found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: the patient sustained an injury during a da vinci surgical hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335508 | 3306200 | 3202795 | 2955842201303185 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-01-11 | Injury | Required InterventionHospitalization | As part of a legal mediation effort, on (b)(4) 2013, intuitive surgical, inc. Received information including medical records of a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2013. The operative (op) report for the patient's hysterectomy procedure on (b)(6) 2013 was not provided. The patient's medical summary alleges that the patient underwent the da vinci hysterectomy procedure secondary to failed medical management of dysmenorrhea and menorrhagia. The surgical procedure was apparently complicated due to dense pelvic adhesions. The patient was discharged the following day with her postoperative course complicated by panic attack. Post operatively the patient reported that she had some right shoulder pain as well as bloating and was managed conservatively with rest, hydration and stool softeners, however, her pain worsened. The patient's medical reports provided state that when the patient returned to the hospital on (b)(6) 2013, the patient described her pain as being so intense that she could not eat or drink. The patient also reported fever and vomiting. Abdomen and pelvic ct tests were ordered. No significant intra-abdominal findings noted. In the right adnexa, multi-cystic collection was seen. The patient also reported constipation and stool softer was prescribed. On (b)(6) 2013, the patient was readmitted due to fever and abdominal pain. She consented for a diagnostic laparoscopy. During the surgical procedure the patient's vaginal cuff was found to be intact; however, there was some greenish discharge coming from the cuff. A manual examination found that the patient had a large immobile mass in the right lower quadrant approximately 6cm. Under direct visualization multiple dense adhesions of the mentum and bowel to the anterior abdominal wall were discovered. During take down of the patient's adhesions, the surgeon made the decision to continue dissection using traditional laparotomy techniques to safely take down the patient's adhesions, since the patient's bowel had adhered to the anterior abdominal wall. During the surgical procedure it was discovered that the patient had an inflamed bowel with some serosal denudation and a pelvic cuff abscess. The abscess was evacuated. Examination of the patient's bowel found several areas of serosal inflammation. A general surgeon was consulted and the decision was made to manage the condition with bowel rest. No repair was required. The patient's abdomen was then copiously irrigated and a jackson pratt drain was placed. The patient tolerated the procedure well and was taken to the post-anesthesia care unit in stable condition. On (b)(6) 2013, the patient discharged herself from the hospital against the advice of the nursing staff. The patient's medical summary alleges that on (b)(4) 2013, the patient returned to urgent care complaining of abdominal pain with nausea and vomiting. When the patient was informed that she had signed herself out of the hospital against the hospital advice, the patient stated that she did not know what she was doing and did not realize that she was at home until awakened at home. During a post-operative follow-up and examination the patient's abdominal incision was found to be well healed and the staples were removed. A follow-up examination on (b)(4) 2013, found that the patient was doing well, had nonspecific gi issues and pain that was consistent with post-operative recovery. The patient had voiced concerns regarding the post-surgical complications she experienced. It was explained to the patient that the post-surgical complications she experienced were likely secondary to the extensive dissections due to her scar tissue and was unrelated to the da vinci surgical technology. On (b)(6) 2013, during a follow-up examination it was discovered that the patient was well healed. On (b)(6) 2013, during a routine follow-up pelvic exam, the patient was healing well and her pain was controlled. The patient had no active issues. | Based in the information provided, intuitive surgical has not determined the root cause of the injury sustained by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however, as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT | 3335555 | 3306222 | 3202220 | 2955842201303186 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-10-25 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013, intuitive surgical received information, including medical records, of a patient who underwent a da vinci si vaginal hysterectomy with bilateral salpingectomy and cystoscopy in addition to a right inguinal hernia repair with mesh placement on (b)(6) 2011. The operative report does not contain any report of a malfunction of a da vinci system, instrument or accessory. Post-operatively, the patient complained of pelvic pain. On (b)(6) 2011 the patient had a ct scan of the abdomen and pelvis which showed ovarian vein thrombosis. The patient was started on lovenox and antibiotics and transitioned to coumadin (anti-coagulant). On (b)(6) 2011 the patient underwent a diagnostic laparoscopy; the post- operative diagnosis was vaginal cuff bleeding. A small adhesion was taken down that was attached to the inguinal hernia repair site. The patient also received 3 blood transfusions during that admission, according to the medical records. At the patient's 6 week postoperative hysterectomy checkup on (b)(6) 2011, the patient complained of occasional random pain from her hernia repair but was doing well otherwise. In (b)(6) 2012, she reported having less anxiety but continued having some persistent problems with pain from her hernia. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complications experienced by the patient. If additional information is received a follow-up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. Review of the site's system logs for the reported procedure date of (b)(6) 2011 found that no system errors were generated that would have caused or contributed to the post-surgical complications experienced by the patient. This complaint is being reported due to the following conclusion: the patient experienced post-surgical complications after undergoing a da vinci si surgical vaginal hysterectomy with bilateral salpingectomy and cystoscopy procedure. | DA VINCI SI SURGICAL SYTSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335565 | 3306252 | 3202847 | 2955842201303183 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-02-01 | Injury | Required InterventionHospitalization | As part of a legal mediation effort on (b)(4) 2013 intuitive surgical, inc. Received information regarding a patient who underwent a da vinci si hysterectomy on (b)(6) 2011 due to persistent abnormal uterine bleeding. The patient's operative report does not contain information regarding any malfunction of the da vinci surgical system, instruments or accessories occurred and there was no indication that the patient experienced any intra-operative complications. The patient's operative report stated that the patient's vaginal cuff was closed with no difficulty and the cuff was found to be hemostatic. The procedure was completed and the patient tolerated the procedure well. The patient was awakened and transferred to the recovery room in stable condition. On (b)(6) 2011, the patient underwent a vaginal cuff closure procedure due to vaginal cuff dehiscence. Per patient's operative report, during examination it was discovered that the patient's vaginal cuff was open in the midportion about 2 cm in length. Bowel content was visible through the cuff opening. The edges of the cuff were found to be intact. The cuff was closed using suture and was closed without difficulty. A second layer of suture was placed along the entire length of the vaginal cuff to provide additional support. The patient remained hemostatic. There were no complications during the procedure. The patient was awakened from anesthesia and was transferred to the recovery room in stable condition. The patient was discharged from the hospital on (b)(6) 2011. | Based on the information provided, intuitive surgical has not determined the root cause of the post-surgical complication experienced by the patient. If additional information is received a follow up medwatch report will be submitted to the fda. The documentation provided does not contain any allegation of a malfunction of a da vinci system, instrument or accessory. Attempts have been made to gather additional information from the risk management group at the site, however as of the date of this report no response has been received. In addition, no previous complaint was reported relating to this event. This complaint is being reported due to the following conclusion: the patient experienced vaginal cuff dehiscence post op a da vinci si hysterectomy procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3335525 | 3306257 | 3202114 | 2955842201303187 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-25 | 2013-08-23 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-25 | Yes | No | Reuse | 2013
IS3000 A60P8 | 2013-07-22 | Other | Other | It was reported that prior to starting a da vinci si hiatus hernia procedure, the light emitting diode (led) indicator on the patient side manipulator (psm) arm 2 was noted to be green and the error message arm touching was observed during startup of the system. Unable to resolve the issue, the surgeon made the decision to abort the planned surgical procedure and rescheduled it for a later date after the patient had been placed under anesthesia. No patient harm, adverse outcome or injury was reported. | The investigation conducted by the field service engineer concluded that the issue experienced by the site was associated with a patient side manipulator (psm) set-up joint (suj). The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The suj supports the instrument and camera arms and are also used to set up the initial positions of the arms. The system was repaired by replacing the suj potentiometer. This complaint is being reported due to the following conclusion: the patient had been placed under anesthesia after surgeon made the decision to abort the surgical procedure due to a malfunction of the system. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3329985 | 3301113 | 3197708 | 2955842201303098 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-07-22 | 2013-08-21 | Yes | Yes | Health Professional | IS3000 A60P8 | Yes | Yes | 2013-07-22 | Yes | No | Reuse | 2013
400272 | 2013-08-19 | Malfunction |  | During robotic case, the harmonic scalpel machine beeped and a message displayed that stated "replace instrument, blade error detected, unplug hand piece and replace instrument. Instrument removed, and replaced after the intuitive rep completed troubleshooting device. Unable to correct issue. No harm to pt. |  | HARMONIC ACE CURVED SHEARS INSERT | ROBOTIC HARMONIC | 3339993 | 3308632 | 3205227 | MW5031587 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-08-20 | 2013-08-21 | No | Yes | Health Professional | 400272 |  | Yes |  |  |  |  | 2013
IS3000 A60P7 | 2013-06-05 | Injury | Required InterventionHospitalization | It was reported that a patient underwent a da vinci si hysterectomy surgical procedure and sustained a bowel injury. Intuitive surgical inc. (isi) spoke to the patient on (b)(4) 2013 and obtained the following information: on (b)(6) 2013, the patient underwent a hysterectomy procedure using a da vinci si system. On (b)(6) 2013, the patient developed sharp abdominal pains and also experienced gas like symptoms. She contacted her surgeon and was advised to take gas medication for her symptoms. On (b)(6) 2013 around 5:30-6:00pm, her symptoms progressively got worse, had a fever of 102 and was taken to the hospital by paramedics. On (b)(6) 2013, the patient went in for an x- ray and it was discovered that the patient had a cut bowel. The patient went in for surgery to repair her cut bowel and it was also discovered that patient had an abscess. The patient was admitted in the hospital for 14 days for an infection and recovery of the cut bowel. No further information was provided by the patient. | Isi contacted the manager of the o. R. On (b)(4) 2013 to obtain additional clinical information about the incident and she refused to provide information to isi. Isi left a message for the surgeon to contact isi and has not received any new clinical information about the incident. Isi has reviewed the system logs for this site with a procedure date of (b)(6) 2013 and no system errors were found to have occurred during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the patient sustained a cut bowel after having a da vinci si procedure. She had to have another procedure to repair her cut bowel and was in the hospital for 14 days for an infection and recovery of her bowel injury. However, at this time it is unknown how the patient obtained the bowel injury and if the da vinci si surgical system caused or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3320354 | 3291536 | 3188131 | 2955842201303064 | 1 | NAY | Manufacturer | OtherConsumer | Other | Initial | 2013-07-26 | 2013-08-16 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2013-07-26 | Device Not Retu | No | Reuse | 2013
420194 | 2010-08-05 | Injury | Required Intervention | On (b)(4) 2013, intuitive surgical inc. (isi) received uf/importer report (b)(4). The report indicated that the patient underwent a robotic myomectomy procedure on (b)(6) 2010 and a robotic hysterectomy and bilateral salpingectomy without complications on (b)(6) 2012. The report also indicated that on 04/18/13, a ct scan was performed on the patient's abdomen due to complaints of abdominal pain. The ct scan found a single metal structure in the deep right pelvis likely hemostatic surgical clip. The same surgeon reportedly performed both robotic surgical procedures on (b)(6) 2010 and (b)(6) 2012. According to the surgeon, the patient had no other abdominal surgeries except for the 2 reported robotic surgical procedures. The surgeon also denied that surgical clips were used during the surgeries. The reporter indicated that on (b)(6) 2013, the patient was readmitted for a laparoscopy to remove the foreign object and for lysis of adhesions. The pathology report described the foreign object that was removed as a 0. 7 x 0. 4 x. 0. 1 cm dark gray piece of metal. | On (b)(4) 2013, intuitive surgical inc. (isi) contacted the initial reporter of this complaint, a nurse, and obtained additional information regarding the reported event. The initial reporter indicated that the site concluded that the foreign object that was retrieved from the patient is the cross-hatched pad that is supposed to be adhered to one of the tips of a mega needle driver instrument. She did not have the lot of the suspect mega needle driver instrument involved with the reported event. The initial reporter did not know if the suspect instrument was used in subsequent surgical procedures. She also did not know current status or location of the suspect instrument involved with the reported event. To her knowledge, the patient was currently doing fine. The site was unable to confirm when the fragment fell into the patient. The system logs did not indicate that a mega needle driver instrument was used in the (b)(6) 2012 surgery. The system logs did indicate that a mega needle driver was used in (b)(6) 2010 surgery, and that the same instrument was used in three subsequent surgeries. That same instrument was subsequently returned for evaluation, and was found to be missing its pad. We are unable to confirm whether the missing pad is the fragment that was retrieved from the patient. Based on the information provided, this complaint is being reported due to the following conclusion: the patient underwent a laparoscopy procedure to remove a foreign metal object. This. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3320376 | 3291549 | 3188144 | 2955842201303065 | 1 | NAY | Manufacturer | Other | Nurse | Initial | 2013-07-21 | 2013-08-16 | Yes | Yes | Health Professional | 420194 | No | Yes | 2013-07-23 | Device Not Retu | No | Unkown | 2013
IS3000 A60P8 | 2013-07-18 | Other | Other | It was reported that during a da vinci si myThe investigation conducted by the intuitive surgical, inc. (isi) field service engineer (fse) was unable to replicate the cautery issue experienced by the site. The fse tested the site's da vinci surgical system and it was found to function within specifications. Functional testing of the system in conjunction with the site's gyrus generator, pk forceps instruments and pk instrument cord found that all of the components used functioned without any issues noted. The pk forceps instruments fired when cautery energy was activated. Review of the site's system logs by the fse found that no system errors were generated that would have caused or contributed to the issue experienced by the site. This complaint is being reported due to the following conclusion: the da vinci surgical procedure was converted to open as a result of not being able to use the pk dissecting forceps instrument to control the patient's bleeding. As of (b)(6) 2013, there have been no reported recurrences of the issue at this hospital.omectomy procedure, the surgeon was unable to activate cautery energy using the pk dissecting forceps instrument while attached to the gyrus generator. The surgical staff reseated the instrument and the pk instrument cord; however, the issue persisted. The site contacted intuitive surgical, inc. (isi) for technical support engineer assistance while the surgeon was in the process of converting the case to a traditional open surgical procedure. The site confirmed with the tse that the cautery cable to the gyrus generator was seated correctly. The tse suggested that the site change the pk instrument cord and the cautery cable; however, the site did not have a spare pk instrument cord or cautery cable. The tse was unable to confirm with the site if the gyrus generator was operational while disconnected from the da vinci surgical system. On (b)(4) 2013, isi contacted the initial reporter of this complaint. The initial reporter indicated that she was circulating during the surgical procedure and that during the surgical procedure, the surgeon experienced difficulty getting the pk dissecting forceps instrument to function. During this time the patient was bleeding and the surgeon attempted to use a bipolar instrument to control the bleeding however, his attempts were unsuccessful. As a result, the surgeon made the decision convert the planned surgical procedure to open surgical techniques to quickly control the patient's bleeding and to complete the surgical procedure. It was stated that the bleeding experienced by the patient was not due to an injury to the patient; but rather normal bleeding as part of the surgical procedure. The patient required a blood transfusion and was given 4 bags of red blood cells and 2 units of plasma. The patient was stabilized. No other information was provided by the initial reporter. On (b)(4) 2013, isi contacted the surgeon who performed the surgical procedure. The surgeon indicated that a few minutes into the surgical procedure, he was unable to get the pk instrument to cauterize when he pressed the foot pedal on the surgeon side cart. The surgeon indicated that he observed that the cautery icons in the high resolution stereo (hrsv) showed that cautery energy was activated, but the instrument was not cauterizing. The bleeding experienced by the patient was due to the patient's fibroids. Unable to get the pk instrument to work, the surgeon attempted to use a bipolar instrument, however, he was not getting the desired cautery and therefore made the decision to convert the procedure to open. The patient tolerated the open procedure well and was discharged from the hospital two days post op. The patient was noted as recovering well. | The investigation conducted by the intuitive surgical, inc. (isi) field service engineer (fse) was unable to replicate the cautery issue experienced by the site. The fse tested the site's da vinci surgical system and it was found to function within specifications. Functional testing of the system in conjunction with the site's gyrus generator, pk forceps instruments and pk instrument cord found that all of the components used functioned without any issues noted. The pk forceps instruments fired when cautery energy was activated. Review of the site's system logs by the fse found that no system errors were generated that would have caused or contributed to the issue experienced by the site. This complaint is being reported due to the following conclusion: the da vinci surgical procedure was converted to open as a result of not being able to use the pk dissecting forceps instrument to control the patient's bleeding. As of (b)(6) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3321750 | 3293203 | 3189798 | 2955842201303070 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-07-18 | 2013-08-16 | Yes | No | Health Professional | IS3000 A60P8 | Yes | Yes | 2013-07-18 | Yes | No | Reuse | 2013
IS3000 A60P8 | 2013-07-18 | Other | Other | It was reported that at the beginning of a da vinci si hysterectomy procedure, the monopolar and bipolar cautery energy were not working. The site power cycled the system and reseated the system cables; however, the issue persisted. The site contacted intuitive surgical, inc. (isi) for technical support engineering (tse) for troubleshoot assistance. The site indicated to the tse that the light emitting diode (led) indicator at the port of the personality module energy device (pmed) was blue and that they did not have a backup energy control cable. The site stated that they switched to an external footswitch and were continuing with the case. Approximately 17 minutes later, the isi clinical sales representative on site for the surgical procedure called isi technical support and reported that the surgeon made the decision to convert the planned procedure to open surgical techniques. No patient harm was reported. On (b)(4) 2013, isi contacted the initial reporter of this complaint. The initial reporter indicated that the surgeon made the decision to complete the procedure using open surgical techniques due to the cautery issue. There was no harm to the patient and the patient tolerated the open procedure. The initial reporter indicated that it was not reported to him that the patient returned to the hospital due to experiencing any post-surgical complications. | The investigation conducted by field service engineering concluded that the cautery issue experienced by the site was associated with the personality module energy device (pmed). Functional testing of the system found that there was no link established between the pmed and energy devices, thus causing the issue experienced by the site. The pmed is designed to interface and to control functions within various cautery generators or esu's (electro-surgical unit), which are activated by a foot pedal switch. The pmed provides enhanced control, enables multiple pedal inputs, multiple or expanded generator control and system visibility of foot pedal stimulus. The system was repaired by replacing the affected pmed. The pmed was returned and evaluated. Engineering evaluation was able to replicate the customer reported failure mode. A cautery test performed with the pmed installed on a test in-house system found that it was not recognized when the system was powered on. Engineering evaluation also found that the light emitting diodes on the pmed were not lit and there were no communication signals, thus indicating that the pmed had a power issue. This complaint is being reported due to the following conclusion: the planned surgical procedure was converted to open due to a malfunction of the system. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3321743 | 3293332 | 3189927 | 2955842201303069 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-07-18 | 2013-08-16 | Yes | Yes | Health Professional | IS3000 A60P8 | Device Returned To Manufacturer | No | 2013-07-18 | Yes | No | Reuse | 2013
IS3000 A60P8 | 2012-06-30 | Injury | Required Intervention | On (b)(6) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2012. According to the patient's medical records, she sustained a small rent/tear in the posterior wall of the bladder. Isi has reviewed the medical records for the patient. Based on the operative report of the hysterectomy procedure performed on (b)(6) 2012, there is no indication that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the surgical procedure. The surgeon indicated that after closing off the vaginal cuff with the suture, we then placed indigo carmine in the patient's vessels. The surgeon also indicated, we then performed a cystoscopy after we removed all the trocars under direct visualization. At the cystoscopy, both ureters were noted to be ejecting indigo carmine. We then saw a small little hole, pin size hole. The operative report indicated that an intraoperative consult was performed by a surgeon from urology and it was confirmed that there was an incidental cystotomy. The patient's incision sites were then reopened and 2 robotic trocars were used in the lateral ports and then a midline camera port. The urology surgeon repaired the bladder injury by performing a standard laparoscopic interrupted suture just above this area on the bladder to use as a holding stitch. The surgeon then performed identical stitching, once we identified the actual hole, by lifting up on the bladder on the holding stitch. Two stitches were placed through this and were tied down using surgeon's knots. A repeat cystoscopy was then performed and it was revealed that part of the bladder wall was tearing with tying down of the knots, indicating very fragile tissue. The surgeon noted in the operative report that there had been overall a decent closure of the bladder, and with filling the bladder with irrigation, the bladder was able to hold the majority of it. He further noted that they felt like this closure be sufficient given the location, the size, and the overall bulking up of the tissue behind it with the stitches. A foley catheter was then placed in the patient and the bladder was drained. There were no reported complications during the repair of the patient's bladder. The urology surgeon noted in the patient's medical records that the bladder injury was most likely due to adhesions and difficulty with dissection due to previous cesarean sections and pelvic surgery. The patient ws discharged from the hospital on (b)(6) 2012. A post-operative visit on (b)(6) 2012, was found to be unremarkable. On (b)(6) 2012, the patient had complaints of urethral and right lower quadrant pain and diarrhea. The patient's medical records indicate that a urine culture revealed enterococcus and she was started on macrobid, an antibiotic. No additional clinical records were available or provided for review after the date of (b)(6) 2012. | Intuitive surgical inc. (isi) has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Based on the review of patient's medical records and the operative reports, there is no indication that a malfunction of the da vinci si system, and instrument, or an accessory occurred during the hysterectomy procedure. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2012. No system errors were found to have occurred during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the patient sustained a bladder injury during a da vinci si surgical procedure. However, at this time, it is unknown what the cause is of the patient's bladder injury. It is also unknown if the da vinci si surgical system caused or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3321991 | 3293493 | 3190088 | 2955842201303071 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-06-13 | 2013-08-16 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-06-13 | Device Not Retu | No | Reuse | 2013
40027202 | 2013-07-17 | Injury | Required Intervention | It was reported that during a da vinci si myomectomy procedure, a piece of blade from the harmonic ace curved shears insert installed on a harmonic ace curved shears instrument broke off and fell into the patient. The fragment was retrieved and the surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | Intuitive surgical inc. (isi) has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. The harmonic ace curved shears insert was returned to isi and evaluated. The curved blade on the insert was found to be broken at the distal end. The broken fragment was not returned to isi. The broken fragment measured roughly. 550 x. 085. The blade was observed to be fractured distal to the clamp arm pin, near the transition from straight to curved. Engineering concluded that the insert damage was likely due to mishandling/misuse. Engineering evaluation also found the shaft of the insert to be slightly bent roughly 4 above the distal end. The distal end of the insert was observed to wobble when the shaft was rotated. Engineering also concluded that the damage to the instrument's shaft was likely due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings of handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Based on the information provided, this complaint is being reported due to the following conclusion: the initial reporter indicated that a fragment from the harmonic ace curved shears insert broke off, fell into the patient, and was retrieved. However, at this time, it is unknown how the fragment was retrieved from the patient. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 3322021 | 3293499 | 3190094 | 2955842201303072 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-07-19 | 2013-08-16 | Yes | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2013-07-19 | Yes | No | Reuse | 2013
IS3000 A60P8 | 2012-12-13 | Injury | Required InterventionLife Threatening | On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci si surgical procedure on (b)(6) 2012. The legal complaint alleged that the patient suffered a perforated colon and experienced intense pain and stool drainage from the surgical site, developed a severe infection, and underwent an emergency repair, debridement, and abdominal washout procedure two days later. The legal complaint also indicated that the patient had considerable pain, disability, convalescence and disfigurement thereafter. | Intuitive surgical inc. (isi) has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2012. No system errors were found to have occurred during the surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the legal complaint claimed that the patient underwent a da vinci si surgical procedure, sustained a perforated colon, and required emergency repair. However, at this time, it is unknown how the patient's colon was perforated and if the da vinci si surgical system caused or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3322044 | 3293502 | 3190097 | 2955842201303074 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-19 | 2013-08-16 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-19 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-07-10 | Death | Death | It was reported that during a da vinci si nissen fundoplication procedure, after the surgeon completed suturing, the patient expired. Reportedly, the patient's demise was related to a pulmonary embolism and that the patient's demise was unrelated to a malfunction of the da vinci surgical system, instruments or accessories. It is believed that the patient experienced complications to bleeding. No other information was provided. | Based on the limited information provided, intuitive surgical, inc. Has not determined the root cause of the patient's demise. Several attempts have been made to verify additional information regarding the reported event; however, no additional information has been received. A follow-up medwatch report will be submitted to the fda if additional information is received. Review of the site's system logs for the reported procedure date found that no system errors were generated that would have caused or contributed to the patient's death. This complaint is being reported based on the following conclusion: the patient expired during a da vinci surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3319757 | 3290825 | 3187420 | 2955842201303049 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-07-16 | 2013-08-15 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-16 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-07-16 | Injury | Required InterventionHospitalization | It was reported that during a da vinci si gynecological procedure, the patient sustained an injury to her ureter. Reportedly, the patient remained hospitalized and it is unknown when the patient would be discharged. On (b)(6) 2013, intuitive surgical contacted the physician's assistant (pa) who assisted during the surgical procedure. The pa indicated that the damage to the patient's ureter was discovered several days post op while the patient was still hospitalized. The pa also indicated that the patient also developed a fistula. The pa did not provide any other information. | Based on the limited information provided, intuitive surgical, inc. (isi) has not determined the root cause of the injury sustained by the patient. Isi has made several follow-up attempts to obtain additional information regarding the reported event; however, no additional information has been provided. A follow-up medwatch report will be submitted to the fda if additional information is received. Review of the site's system logs for the reported procedure date found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: the patient sustained an injury during a da vinci surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3319746 | 3290846 | 3187441 | 2955842201303050 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-07-16 | 2013-08-15 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-16 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-07-12 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, while performing suture, the surgeon sewed into the patient's bladder, which created a hole in the patient's bladder. The patient's bladder was repaired and the patient was discharged from the hospital the same day. On (b)(6) 2013, intuitive surgical, inc. Contacted the physician's assistant (pa) that was assisting during the da vinci surgical procedure. The pa indicated that after the surgeon had completed closure of the patient's vaginal cuff and after the da vinci surgical system was undocked from the patient, a cystoscopy performed on the patient found that the surgeon had inadvertently sutured the patient's bladder during closure of the patient's vaginal cuff. The da vinci surgical system was re-docked to the patient and the surgeon removed the suture from the patient's bladder and re-sutured the patient's vaginal cuff to correct the misplaced suture. No repair to the patient's bladder was required. The pa indicated that no malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure and that inadvertent suture of the patient's bladder by the surgeon was due to the surgeon's technique. The pa indicated that to the best of his knowledge, the patient has not returned to the hospital due to experiencing any post-surgical complications. | Based on the additional information provided by the site, the da vinci surgical system did not malfunction in a way that caused or contributed to the injury sustained by the patient; but was due to the surgeon's technique, while suturing the patient's vaginal cuff. Review of the site's system logs for the reported procedure date found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: a patient sustained a bladder injury during use of the da vinci surgical system. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3319775 | 3290852 | 3187420 | 2955842201303052 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-07-16 | 2013-08-15 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-07-16 | Device Not Retu | No | Reuse | 2013
400272 | 2013-08-12 | Malfunction |  | The white tip portion of the shears became loose. Device was replaced with a new shear insert. The defective item was removed off the surgical field. What was the original intended procedure?hiatal hernia. Device #1is this a laboratory device or laboratory test?no. |  | HARMONIC ACE CURVED SHEARS INSERT | SYSTEM SURGICAL COMPUTER | 3323009 | 3294441 | 3191036 | 3294441 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-08-14 | 2013-08-14 | No | Yes | Physician | 400272 | No | No Answer Provided |  |  |  |  | 2013
42020505 | 2013-07-30 | Injury | Required Intervention | It was reported that during a da vinci si prostatectomy procedure on (b)(6) 2013, a piece of the cable on the bipolar forceps instrument snapped off and fell into the patient. The broken piece was retrieved; however, it is unknown how the fragment was retrieved. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Intuitive surgical inc. (isi) has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 3311927 | 3283198 | 3179793 | 2955842201303020 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-07-30 | 2013-08-12 | Yes | Yes | Health Professional | 42020505 | No | No | 2013-07-30 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2013-07-12 | Death | Death | It was reported that a patient who underwent a da vinci si mitral valve repair procedure on (b)(6) 2013, expired post-operatively. A surgeon who was part of the cardiac team at the site informed an intuitive surgical inc. (isi) clinical sales representative (csr) that the surgical procedure was abandoned since the patient developed pulmonary edema. No further clinical information was provided about the procedure or about the patient. | Intuitive surgical inc. (isi) has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2013. No system errors were found to have occurred during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: it was reported that the patient who underwent a da vinci si surgical procedure passed away post-operatively. However, at this time, it is unknown if a malfunction of the da vinci si system, an instrument, or an accessory occurred during the surgical procedure. It is also unknown if the da vinci si surgical system caused or contributed to the patient's demise. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3311948 | 3283208 | 3179803 | 2955842201303025 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2013-08-01 | 2013-08-12 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2013-08-01 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-08-01 | Other | Other | It was reported that during a da vinci si surgical procedure on (b)(6) 2013, the site noticed that the left master tool manipulator (mtml) was hanging down in the surgeon side console(ssc). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The surgeon was moving the mtml; however, the mtml did not control anything. The mtml arm was reportedly sluggish. The site attempted to restart the system; however, the alleged issue continued. The surgeon decided to abort the surgical procedure post anesthesia with port placement prior to contacting intuitive surgical inc. (isi) technical support engineer (tse). It is unknown at this time whether the surgeon aborted the surgical procedure due to the alleged issue with the mtml. | Isi has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of date of this report. A follow-up mdr will be submitted if additional information is received. An isi field service engineer (fse) went to the site on (b)(4) 2013 and noticed that the only way the site could pass homing was to disable the mtml arm remote arm controller boards (rac1). Rac1 was replaced in order to resolve the error 25580 which occurred during surgery. This system error code 25580 is reported when the da vinci safety system determines a hardware voltage level is out of range. During test drive, fse noticed error 23005 which occurred on the psm1 axis 6 during test drive. This system error code 23005 is reported when there is an issue with the output power limit. Fse had ran an electrical safety test and the safety test passed testing. Fse also ran a system verification which also passed testing. Isi has reviewed the system logs for this site with a procedure date of (b)(4) 2013 and confirmed error 25580 was reported. Error 23005 was in the system log during test drive by the fse on (b)(4) 2013. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3311961 | 3283230 | 3179825 | 2955842201303026 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-08-01 | 2013-08-12 | Yes | Yes | Health Professional | IS3000 A60P8 | Device Returned To Manufacturer | No | 2013-08-01 | Yes | No | Reuse | 2013
IS3000 A60P3 | 2013-06-27 | Injury | Required Intervention | It was reported that 2 days post a successful da vinci si cholecystectomy procedure, the patient was taken back into the or presenting symptoms of pain. It was discovered that the patient had sustained a laceration of the small bowel. Reportedly, the patient was released from the hospital and is recovering fine. | On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the surgeon who performed the da vinci surgical procedure. The surgeon indicated that during the planned surgical procedure, no apparent malfunction of the da vinci surgical system, instruments or accessories occurred. The surgeon indicated that he did not know what caused the damage to the patient's bowel; however, it was his belief that the patient's injury was not caused by the da vinci surgical system, instruments or accessories. The surgeon declined to provide any other details regarding the reported event. Review of the site's system logs for the reported procedure date found that no system errors were generated that would have caused or contributed to the injury sustained by the patient. This complaint is being reported due to the following conclusion: the patient's bowel was damaged during a da vinci surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3307021 | 3277482 | 3147499 | 2955842201303010 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-07-10 | 2013-08-10 | Yes | No | Health Professional | IS3000 A60P3 | No | No | 2013-07-10 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2012-12-19 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the customer experienced non-recoverable system error code 212. With the assistance of an intuitive surgical, inc. (isi) technical support engineer (tse), the site removed the instrument installed on the patient side manipulator (psm) arms and powered off the system. When the site restarted the system, the issue recurred. The tse instructed the site to exercise the left master tool manipulator (mtml) on the surgeon side cart and perform a fault override; however, system error code 212 recurred. Unable to resolve the issue the surgeon made the decision to abort the planned surgical procedure. No patient harm was reported on (b)(6) 2013, isi contacted the site's or unit assistant. The or unit assistant indicated that the planned surgical procedure was completed using open surgical techniques. She indicated that she was unable to provide any other details regarding the reported event. | The investigation conducted by the field service engineer concluded that system error code 212 experienced by the customer was associated with the left master tool manipulator (mtml) and the multiple servo driver (msd) pca board. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system contains five multiple servo driver (msd) pca boards, which provide drive current to the servo motors associated with each degree of freedom. The system was repaired by replacing the affected mtml and msd pca board. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 212 occurs when the actual voltage to drive current through the motors deviates from the expected voltage by a specified amount. Upon determining this condition, the safety systems put the da vinci in a non-recoverable safe state. The mtml and the msd board were returned to intuitive surgical, inc. (isi) for failure analysis investigation. Engineering confirmed the customer reported system error code 212 and found that the main cable harness and axis 3 motor were defective. The msd board output was found to be out of the acceptable range. As of (b)(6) 2013, there have been no reported recurrences of the issue at this hospital. This complaint is being reported due to the following conclusion: the surgeon made the decision to convert to open surgical techniques after encountering an recurring system error code 212 with the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3304991 | 3275500 | 3172095 | 2955842201302991 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-08-09 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2012-12-19 | Yes | No | Reuse | 2013
IS3000 A60P8 | 2013-07-11 | Death |  | It was reported that during a da vinci si thoracic procedure, the patient expired. Reportedly, prior to the patient's demise, while the surgeon was manipulating the master tool manipulators (mtms) on the surgeon side console, the surgeon was unable to control or move the instrument installed on a patient side manipulator (psm) arm 1. Reportedly, the site observed that the led on psm 1 was solid blue in color when the issue occurred and they were eventually able to resolve the issue. On (b)(6) 2013, intuitive surgical contacted the surgeon who performed the surgical procedure. The surgeon indicated that the patient's demise was unrelated to his inability to manipulate the instrument installed on psm 1, but rather the patient's condition. The patient was very sick and was not doing very well under anesthesia. The surgeon indicated that he had intended to perform lung decort using the da vinci surgical system; however, after the system was docked to the patient and when he sat down at the ssc, he was unable to move the instrument that was installed on the psm. The patient had chest adhesions and as a result he had to create larger port site incisions. He did not observe any restrictions around the port sites that would have caused or contributed to the issue that he experienced. The surgeon indicated that the patient had severe vena cava obstruction and was a dialysis patient. The surgeon indicated due to the patient's venous blockage the patient was bleeding from his port site incisions and during the time that the issue occurred, the patient began to code. The system was immediately undocked from the patient and resuscitation was performed on the patient for 45 minutes to 1 hour without success. On (b)(6) 2013, the site's da vinci coordinator provided additional information regarding the reported event. The da vinci coordinator indicated that the cannulae used during the surgical procedure were inspected prior to use and there was no damage found. A post-op inspection of the cannulae was also performed and there was no damage to the cannulae noted. | The investigation conducted by the field service engineer (fse) was unable to duplicate the issue experienced by the site. Functional testing of the system found that the system functioned within specification. This complaint is being reported due to the following conclusion: a patient death occurred during a da vinci surgical procedure. Based on the information provided, it has been determined that the da vinci surgical system, instruments and/or accessories did not cause or contribute to the patient's demise; but rather the patient's condition. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3305565 | 3276053 | 3172648 | 2955842201303004 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-07-11 | 2013-08-09 | Yes | Yes | Health Professional | IS3000 A60P8 | Yes | No | 2013-08-14 | No | No | Reuse | 2013
42000606 | 2013-05-28 | Injury | Required Intervention | On (b)(4) 2013, intuitive surgical, inc. Received (b)(4) from the fda. The event details are provided below: robotic needle holder noted to fray apart at tip of instrument. No pieces noted to come off, instrument removed from patient. No pieces noted in operative site. X-ray taken at end of procedure to verify. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. This complaint is being reported due to the following conclusion: it was reported that during a da vinci surgical procedure, the cable on the large needle driver instrument frayed, requiring the patient to undergo an x-ray to determine if any piece(s) from the instrument fell into the patient. Per the information provided, there were no fragments from the instrument observed falling into the patient. In addition, the x-ray did not reveal that the patient retained any fragments from the instrument. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3306039 | 3276513 | 3173108 | 2955842201303000 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-12 | 2013-08-09 | Yes | Yes | Health Professional | 42000606 | No | No | 2013-07-12 | Device Not Retu | No | Unkown | 2013
IS3000 A60P6 | 2013-07-10 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the site experienced issues with the endoscopic camera manipulator (ecm) arm. The site contacted their distributor for trouble shooting assistance. The distributor representative arrived on site and found that the cables connecting the camera roll motor with the camera mount on the ecm were broken. The site attempted to proceed with the procedure with the ecm in its broken state; however, everytime the surgeon manipulated the ecm from the master tool manipulators (mtms) on the surgeon side cart, the ecm arm rolled a bit, but the movement was non-intuitive. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The investigation conducted by an intuitive surgical, inc. (isi) representative confirmed the customer reported complaint. The fse repaired the system by replacing the endoscopic camera manipulator (ecm) arm. The ecm provides the sterile interface for the 3d endoscope. The ecm has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the ecm is returned (post engineering evaluation) or if additional information is received. This complaint is being reported due to the following conclusion: the planned surgical procedure was converted to open surgical techniques due to a malfunction of the da vinci surgical system. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3306223 | 3276701 | 3173296 | 2955842201302997 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-07-10 | 2013-08-09 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2013-07-10 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2010-09-28 | Death | Death | On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2010. The legal complaint alleged that in the immediate days following the procedure, the patient experienced symptoms of severe abdominal pain and extreme distension of abdomen, trouble standing or walking, difficulty breathing, lack of bowel movements and flatus, and black foul smelling emesis. The legal complaint also indicated that on (b)(6) 2010, the patient was transported to a hospital via ambulance and admitted for severe abdominal sepsis and inflammation of the abdomen and abdominal tissue. In addition, the patient was operated on immediately after admittance to the hospital. According to the legal complaint, during the operation, doctors observed the patient was septic and had a punctured colon. The patient was kept in the intensive care unit (icu) in severe critical condition maintained by a ventilator, so that the doctors could continue to treat her. Also, the legal complaint indicated that despite invasive medical procedures every two days over the duration of two weeks in attempt to treat the patient, the sepsis was too advanced and irreversible. The legal complaint claimed that the patient passed away on (b)(6) 2010, due to sepsis and hypotension caused by a punctured colon from the da vinci surgical system surgery. | Intuitive surgical inc. (isi) has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2010. No system faults/errors were found to have occurred during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the legal complaint claims that the patient underwent a da vinci si surgical procedure and subsequently passed away due to sepsis and a punctured bowel. At this time, it is unknown if a malfunction of the da vinci si system, and instrument, or an accessory occurred during the surgical procedure. It is also unknown if the da vinci si surgical system caused or contributed to the patient's injury and subsequent demise. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3306356 | 3276828 | 3173423 | 2955842201302996 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-12 | 2013-08-09 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-07-12 | Device Not Retu | No | Reuse | 2013
VS3000 | 2013-06-25 | Death | Death | (b)(6) female undergoing a robotic myomectomy and bilateral neosalpingostomy had an intraoperative death. Near the end of the procedure, the pt's blood pressure was not detected in the presence of a pulse and normal oxygenation levels. The rhythm was noted to be pulseless electrical activity (pea). Despite acls procedures and an attempt to insert an intra-aortic balloon pump to sustain the pt, all resuscitative efforts were unsuccessful. Resuscitative efforts were sustained for over 2 hrs at no time was a problem with the robotic equipment reported. This was an unexpected death. Provisional autopsy results: av septum revealed an area of bleeding and fatty infiltration. There was fatty infiltration of the liver. However, in the absence of other findings, definitive characterizations for the cause of the arrhythmia and left ventricular failure is not possible. The final results are still pending. |  | DA VINCI ROBOT | VISIONCART | 3320387 | 3289234 | 3147499 | MW5031411 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-08-08 | 2013-08-09 | Yes | No | Health Professional | VS3000 | Yes | Yes |  |  |  |  | 2013
VS3000 | 2013-06-25 | Death | Death | (b)(6) female undergoing a robotic myomectomy and bilateral neosalpingostomy had an intraoperative death. Near the end of the procedure, the pt's blood pressure was not detected in the presence of a pulse and normal oxygenation levels. The rhythm was noted to be pulseless electrical activity (pea). Despite acls procedures and an attempt to insert an intra-aortic balloon pump to sustain the pt, all resuscitative efforts were unsuccessful. Resuscitative efforts were sustained for over 2 hrs at no time was a problem with the robotic equipment reported. This was an unexpected death. Provisional autopsy results: av septum revealed an area of bleeding and fatty infiltration. There was fatty infiltration of the liver. However, in the absence of other findings, definitive characterizations for the cause of the arrhythmia and left ventricular failure is not possible. The final results are still pending. |  | DA VINCI ROBOT | VISIONCART | 3320387 | 3289234 | 3147499 | MW5031411 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-08-08 | 2013-08-09 | Yes | No | Health Professional | VS3000 | Yes | Yes |  |  |  |  | 2013
VS3000 | 2013-06-25 | Death | Death | (b)(6) female undergoing a robotic myomectomy and bilateral neosalpingostomy had an intraoperative death. Near the end of the procedure, the pt's blood pressure was not detected in the presence of a pulse and normal oxygenation levels. The rhythm was noted to be pulseless electrical activity (pea). Despite acls procedures and an attempt to insert an intra-aortic balloon pump to sustain the pt, all resuscitative efforts were unsuccessful. Resuscitative efforts were sustained for over 2 hrs at no time was a problem with the robotic equipment reported. This was an unexpected death. Provisional autopsy results: av septum revealed an area of bleeding and fatty infiltration. There was fatty infiltration of the liver. However, in the absence of other findings, definitive characterizations for the cause of the arrhythmia and left ventricular failure is not possible. The final results are still pending. |  | DA VINCI ROBOT | VISIONCART | 3320387 | 3289234 | 3147499 | MW5031411 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-08-08 | 2013-08-09 | Yes | No | Health Professional | VS3000 | Yes | Yes |  |  |  |  | 2013
IS3000 A60P6 | 2013-07-11 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the surgeon experienced non-intuitive motion while manipulating the instruments installed on the pt side manipulator (psm) arms from the master tool manipulators. It was also reported that the surgeon experienced difficulty using cautery. The site contacted intuitive surgical, inc (isi) for troubleshooting assistance, after the surgeon had made the decision to complete the surgical procedure using open surgical techniques. No pt harm was reported. Review of the site's system logs by the tse found that there were no system errors codes generated related to the reported issue. On (b)(4) 2013 isi contacted the nurse manager who initially reported this complaint. The nurse manager indicated that she was present during the surgical procedure. The pt was placed in the trendelenburg position and that she observed that the port site incisions were slightly close and the psms were "too straight". She indicated that the surgical staff attempted to troubleshoot the issue, but were unable to resolve the issue. The nurse manager was unable to provide the details concerning the specific troubleshooting techniques that were performed to resolve the issue. On (b)(4) 2013 isi contacted the surgeon who performed the surgical procedure. The surgeon indicated that approx 20 minutes into the da vinci surgical procedure, he observed that the image through the high resolution stereo viewer (hrsv) was inverted. After activating cautery energy while using the right master tool manipulator, tissue on the left side was cauterized. There was no injury to the pt, as the affected tissue was being cauterized as part of the surgical procedure. The surgeon indicated that there were no issues with placement of the port site incision or the portion of the psm during the surgical procedure. He indicated that the surgeon side console settings were found to be correct and there were no camera issues that occurred during the surgical procedure. The surgeon indicated that the pt tolerated the open procedure well and is recovering well. | The investigation performed by intuitive surgical, inc (isi) field service engineering was unable to replicate the issue experienced by the surgeon. This complaint is being reported due to the following conclusion: the surgeon made the decision to convert to open surgical techniques after experiencing non-intuitive movement. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3337801 | 3289502 | 3186097 | 2955842201303015 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-07-11 | 2013-08-09 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | Yes | 2013-07-11 | Yes | No | Reuse | 2013
| 2013-05-22 | Malfunction |  | During the procedure the vessel sealer blade would not retract into the instrument, causing a revocable fault. The instrument was removed and a new one was opened. Since the blade would not retract, it left the blade exposed. No harm to patient. What was the original intended procedure?operative procedure:1. Da vinci robotic laparoscopic hysterectomy. 2. Bilateral salpingectomy. 3. Morcellation of uterus with fibroids. Device #1is this a laboratory device or laboratory test?no. |  | DA VINCI SI VESSEL SEALER | SYSTEM SURGICAL COMPUTER | 3320170 | 3291298 | 3187893 | 3291298 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-08-09 | 2013-08-09 | No | Yes | UNKNOWN |  | Yes | No Answer Provided |  |  |  |  | 2013
| 2013-05-20 | Malfunction |  | The robotic instrument, the endowrist one da vinci si vessel sealer was not recognizd as a medical device and therefore unable to use during procedure. What was the original intended procedure?1. Da vinci robotic laparoscopic hysterectomy with bilateralsalpingo-oophorectomy. 2. Abdominopelvic washings. 3. Bilateral pelvic lymphadenectomy. 4. Paraaortic lymphadenectomy. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST | SYSTEM SURGICAL COMPUTER | 3328428 | 3299572 | 3196167 | 3299572 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-08-09 | 2013-08-09 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2013
IS3000 A60P6 | 2013-07-09 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the cannula mount on patient side manipulator (psm) arm 1 snapped and pieces of the cannula mount broke off. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. On (b)(4) 2013 isi contacted the surgical technician who initially reported this complaint. The surgical technician indicated that the cannula mount broke while the patient side assistant was docking the psm to the patient. Due to the broken cannula mount, the surgeon made the decision to complete the surgical procedure using open surgical techniques. It is unknown what caused the cannula mount to break. The surgical technician indicated that there was no harm to the patient and the patient tolerated the open procedure fine. There were no intra-operative complications experienced by the patient as a result of the reported event. | The investigation conducted by intuitive surgical, inc. (isi) field service engineering (fse) confirmed the customer reported failure mode. The fse found that the cannula mount on patient side manipulator (psm) arm 1 was broken on the right wing at the lower hinge that connects the wing to the clamp. The fse was able to work the hinge closed and it opened and closed freely. A test cannula was installed onto the cannula mount; however, the holding force was found to be inadequate. The psm cannula mount accessory is designed to hold the cannula securely in place. The cannula mount on psm 1 was replaced to repair the system. The cannula mount on psm 2 was replaced as a precaution. The fse provided retraining to the nurse manager on the proper closing techniques of the cannula mount. The da vinci surgeon si surgical system user manual instructions demonstrate how to properly close the cannula mounts and specifically states: seat the cannula mount molding. Make sure the molding snaps over the cannula mount this complaint is being reported based on the following conclusion: the surgeon made the decision to convert the planned surgical procedure to open due to a broken cannula mount. As of (b)(6), there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3301970 | 3272972 | 3169565 | 2955842201302985 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-07-09 | 2013-08-08 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2013-07-09 | Yes | No | Reuse | 2013
420205 | 2013-07-31 | Malfunction |  | Broken bipolar/spring. All parts were retrieved. What was the original intended procedure?robotic assisted myomectomy/morcellation. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST PK DISSECTING FORCEPS | SYSTEM SURGICAL COMPUTER | 3312758 | 3283765 | 3180360 | 3283765 | 1 | NAY | User Facility |  |  | Initial | 2013-08-08 | 2013-08-08 | No | Yes | Invalid Data | 420205 |  | No Answer Provided |  |  |  |  | 2013
| 2013-07-31 | Malfunction |  | Robot vessel sealer was used by surgeon but it was not cutting well. Item was removed from use and a new one was opened. No problems with the new one. The device and its original packaging were given to the hospital's materials management department to return to the manufacturer. |  | ENDOWRIST ONE | SYSTEMSURGICALCOMPUTER CONTROL | 3337071 | 3308203 | 3204798 | 3308203 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-08-06 | 2013-08-06 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
IS2000 A51P8 | 2013-07-03 | Other | Other | It was reported that during a da vinci s nephrectomy procedure, the surgical staff encountered an unrecoverable fault 23 error code with a message indicating that the fiber cable required cleaning. Due to the reported issue, the surgeon made the decision to convert the da vinci s surgical procedure to open surgical techniques. No patient harm, adverse outcome, or injury was reported. After the da vinci s surgical procedure was converted to open surgical techniques, the surgical staff contacted an intuitive surgical inc. 's (isi) technical support engineer (tse). The tse advised the surgical staff to power down the system, and to clean both ends of the red fiber cable and the connections on the patient side cart (psc) and surgeon side console (ssc). | On (b)(4) 2013, an isi field service engineer (fse) contacted the site to follow up regarding the reported issue. The fse indicated that after the surgical staff spoke to the tse on (b)(4) 2013, they repaired the system by reseating the red fiber cable from the ssc to the psc. It is unknown if the surgical staff cleaned the ends of the red fiber cable prior to reseating the cable. According to the da vinci s system user manual, the red system cable and it's receptacles contain fiber optics. Clean both cable and receptacles by blowing the fiber ends with a generic canned aero duster as needed. The system was able to power on and homed normally. The system was monitored for several cases and there were no recurrences of the reported issue. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2013. The system log revealed that the site began to receive multiple non-recoverable 23 error codes more than 2 hours after the da vinci s surgical procedure had started. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon converted the da vinci s surgical procedure to open surgical techniques after encountering multiple non-recoverable 23 error codes. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3292066 | 3262824 | 3159399 | 2955842201302896 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-07-03 | 2013-08-05 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2013-07-03 | Device Not Retu | No | Reuse | 2013
| 2010-07-29 | Injury | Required InterventionHospitalization | On (b)(6) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci prostatectomy procedure. The legal complaint alleged that the patient underwent surgery on (b)(6) 2010, which resulted in damage, including a large rectal laceration with recto-vesical fistula that required intraoperative repair with sutures, a loop ileostomy and further inpatient hospitalization. The legal complaint also indicated that the patient continues to suffer from chronic abdominal pain, incontinence, multiple hernias and intestinal scarring which has required him to undergo multiple additional surgical procedures. In addition, the legal complaint indicated that due to the injuries sustained during the da vinci surgical procedure, the patient had and will continue to have multiple painful additional medical tests and procedures, physician consultations and additional surgeries and has suffered and will continue to suffer pain, loss of function, emotional distress, and permanent injury. | Intuitive surgical inc. (isi) has made several attempts to contact the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. At this time, it is unknown what facility/site the reported event occurred. In addition, the details of how the patient was injured are unknown. Based on the information provided, this complaint is being reported due to the following conclusion: the legal complaint alleged that as a result of undergoing a da vinci prostatectomy procedure, the patient was injured and required additional surgical procedures. However, at this time, it is unknown if a malfunction of the da vinci surgical system caused or contributed to the patient's injury. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3292070 | 3262846 | 3159421 | 2955842201302897 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-07-03 | 2013-08-05 | Yes | No | Health Professional |  | No | No | 2013-07-03 | Device Not Retu | No | Unkown | 2013
IS2000 | 2013-07-15 | Other | Other | On (b)(6) 2013, the reporter contacted intuitive surgical inc. Technical support engineer (tse) alleging flickering of the right eye of the da vinci s system during a cystectomy procedure. The surgeon was concerned about the constant flickering of the right eye and decided to convert the surgical procedure to open surgery. The patient was under anesthesia for 30 minutes prior to the surgeon deciding to convert the surgical procedure to open surgery. Reporter mentioned that the flickering of the right eye was intermittent and could be seen on both the surgeon console and the touchscreen. The patient was fine after the open surgical procedure and did not have any complications due to open surgery. Reporter contacted tse after the surgeon decided to convert the surgical procedure to open surgery. | The device has not yet been returned to isi for investigation. If any addition information becomes available and if the product is returned, intuitive surgical will evaluate the product and inform fda of the product investigation in a supplemental report. Intuitive surgical was able to speak to the reporter on (b)(4) 2013 to obtain additional information: the site had the same issue last week with the flickering of the right eye; however, when they experienced the issue last week, the flickering was intermittent and happened only once in a while. A replacement of a blue cable was delivered to the site, which they had not received at the time of the initial call to the tse. The camera head was also replaced last week. Reporter mentioned that they received the blue cable after initially speaking to the tse on (b)(4) 2013; however the flickering still occurred on the surgeon console and the touchscreen. Field service engineer (fse) came on site on (b)(4) 2013 and replaced the camera cable which resolved the issue. Fse was present at the next da vinci procedure with no problems. The complaint is being reported because the surgeon converted the da vinci procedure to open surgical technique due to the constant flickering of the right eye of the surgeon console. | DA VINCI S SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3292863 | 3263561 | 3160136 | 2955842201302915 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-07-15 | 2013-08-05 | Yes | Yes | Health Professional | IS2000 | No | Yes | 2013-07-15 | No | No | Reuse | 2013
IS3000 A60P6 | 2013-07-25 | Other | Other | On (b)(6) 2013, the surgeon contacted intuitive surgical inc. (isi) technical support alleging that an error 702 on psm1 occurred during a da vinci si lobectomy procedure. Error 702 is triggered when communication to instrument interface board fails. The patient side manipulator (psm) is an instrument arm located on the patient side cart (psc) that provides sterile interface for the endowrist instrument. The surgeon contacted isi technical support and indicated that the system started ok with no problems; however, after sometime the psm arm 1 failed. The surgeon had not started the surgical procedure yet; however, the patient was under anesthesia. The site received a message to disable psm1 or restart the system. The site restarted the system several times; however, psm1 continued to fail. The isi technical support reviewed the system logs and noticed the following error messages 25521, 702 and 23034 were all on the psm1. Per surgeon the psm1 was replaced on (b)(4) 2013. Isi technical support confirmed when looking at the logs that psm1 was replaced and that the logs also showed that the alleged error 702 had already appeared during that time. The surgeon felt uncomfortable and was unwilling to proceed with the surgery and decided to abort the surgical procedure post anesthesia. | The device has not yet been returned to isi for investigation. If any addition information becomes available and if the product is returned, intuitive surgical will evaluate the product and inform fda of the product investigation in a supplemental report. An isi field service engineer (fse) went to the site on (b)(4) 2013 and replaced psm1. Fse ran a system test and verified that the system was ready for use with no problems found after the replacement of psm1. The subject psm1 is being requested back for investigation. Isi has made several attempts to contact the surgeon to obtain additional information concerning the reported event and on (b)(4) 2013 isi received an email from the surgeon stating that the da vinci system was inspected prior to the surgical procedure and passed testing. The patient was under anesthesia for approximately one hour prior to the surgeon aborting the procedure. Surgeon confirmed that there were no ports placed on the patient at the time of the alleged issue. This particular surgery was not rescheduled for a da vinci procedure and the surgeon decided to do a vats lobectomy on (b)(6) 2013. The patient was discharged on (b)(6) 2013 with no complications. Surgeon mentioned that the psm1 was replaced and was working as intended. The complaint is being reported since the surgeon decided to abort the surgical da vinci procedure post anesthesia due an error 702 on the psm1. | DA VINCI SI SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3292872 | 3263585 | 3160160 | 2955842201302916 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-07-25 | 2013-08-05 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2013-07-25 | Yes | No | Reuse | 2013
IS3000 A60P8 | 2013-07-03 | Other | Other | It was reported that during a da vinci si beating heart revascularization procedure, the surgical staff indicated that patient side manipulator 1 (psm1) had jerky movements and was not responding to the surgeon's movements. The psm is an instrument arm located on the patient side cart (psc) that provides sterile interface for the endowrist instrument. In addition, the surgical staff indicated that clicking sounds could be heard from psm1. Due to the psm1 issue, the surgeon made the decision to convert the surgical procedure to open surgical techniques. | On (b)(4) 2013, an intuitive surgical inc. (isi) field service engineer (fse) performed a field evaluation at the site. Through troubleshooting, the fse was unable to replicate the reported issues with psm1. However, per the request of the surgical staff, the fse replaced psm1. The fse then performed a system test that passed according to manufacturer specifications. On (b)(4) 2013, isi contacted the initial reporter of this complaint, the robotics coordinator, and obtained additional information regarding the reported event. According to the robotics coordinator, the surgical staff reseated the sterile adapter on psm1 several times when the reported issue began. Since the reported issue with psm1 could not be resolved, the surgeon decided to convert the surgical procedure to open surgical techniques. The robotics coordinator indicated that the psc was only docked to the patient for a short period of time before the surgical procedure was converted to open surgery. She denied that the patient experienced any intra-operative or post-operative complications. The robotics coordinator also indicated that the reported issue has since been resolved after psm1 was replaced. As of the date of this report, there have been no reported recurrences of the alleged psm issue. The psm1 has been returned to isi for evaluation. However, at this time, the evaluation of the psm has not been completed. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted post-engineering evaluation or if additional information is received. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2013. No system errors were found to have occurred during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon converted the da vinci si surgical procedure to open surgical techniques after encountering jerky movements with psm1. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3291901 | 3262747 | 3159302 | 2955842201302895 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-07-03 | 2013-08-02 | Yes | Yes | Health Professional | IS3000 A60P8 | Device Returned To Manufacturer | No | 2013-07-03 | No | No | Reuse | 2013
| 2013-07-31 | Malfunction |  | Da vinci prograsp instrument stopped working while md was using instrument. Jaws would not open. Or tech started to remove the instrument from the sterile field and the silver metal tab button broke off on the outside of the pt. Button retrieved and instrument removed from the field and replaced with another da vinci instrument. Original instrument had 4 lives left. What was the original intended procedure?total laparoscopic hysterectomy. |  | ENDOWRIST PROGRASP | SYSTEM SURGICAL COMPUTER CONTR | 3297383 | 3268250 | 3164825 | 3268250 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-08-02 | 2013-08-02 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
IS3000 A60P8 | 2013-06-26 | Injury | Required Intervention | Event Description: --/-- It was reported that during a da vinci si surgical procedure, the surgeon encountered a recoverable fault while using the vessel sealer instrument. The initial reporter (robotics coordinator) of this complaint contacted an intuitive surgical inc. (isi) technical support engineer (tse) for assistance. However, during the call, the robotics coordinator did not have time to troubleshoot. The robotics coordinator indicated that the surgeon chose to convert the da vinci surgical procedure to traditional laparoscopic techniques. Initially, no patient harm, adverse outcome, or injury was reported. The tse reviewed the site's system logs and found an error 32001 related to the vessel sealer instrument. An error 32001 indicates that the vessel seal instrument cutting blade cannot be retracted and may be exposed. | Manufacturer Narrative: --/-- The vessel sealer instrument was returned to isi and evaluated. Engineering was unable to install the instrument through an 8mm cannula based on the condition the device was received. Engineering evaluated the instrument by installing the device in a larger diameter stapler cannula. The instrument blade was not found to be dislodged. Engineering indicated that the instrument was able to cut and seal properly. However, according to engineering, bio-debris was found in both blade tracks. Engineering indicated that excessive bio-debris may have contributed to blade jams. Engineering also found the instrument's pitch/yaw wrist disc dislodged from the middle wrist disc. The range of motion of the instrument was limited due to the dislodgement. Engineering concluded that the dislodgement was likely due to mishandling/misuse. --/-- Additional Manufacturer Narrative: --/-- On (b)(6) 2013, an isi field service engineer (fse) performed a field evaluation at the site. The fse repaired the da vinci si surgical system by replacing the instrument control box (icb). The fse tested and verified the system to manufacturer specifications while using a vessel sealer training instrument. On (b)(6) 2013, isi contacted the robotics coordinator for follow-up information. The robotics coordinator indicated that the surgical procedure was successfully completed via traditional laparoscopic techniques with no report of injury. However, contrary to the initial report the robotics coordinator also indicated that there was a post operative complication which required the patient to return to surgery. The robotics coordinator indicated that the conversion from the da vinci procedure to laparoscopic during the (b)(6) 2013 case was directly due to the issue of the fault experienced with the vessel sealer instrument. On (b)(6) 2013, isi contacted the robotics coordinator to request further information, at which time the robotics coordinator indicated that she was not present in the operating room (or) when the surgical procedure was starting. When robotics coordinator arrived in the or, she indicated that the surgeon was attempting to release tissue that was being held by the vessel sealer instrument and for an unknown reason the surgeon was reportedly unable to release the grips on the instrument. The surgical assistant at the patient side cart (psc) was unable to troubleshoot the reported problem with the instrument, however, the robotics coordinator stated that the reported issue was resolved after a fault override was performed. At an unspecified time after the surgical procedure was completed via traditional laparoscopic techniques, the robotics coordinator indicated that the patient was brought back to the or due to experiencing unspecified post-operative complications. The robotics coordinator was unable to provide details of what post-operative complications the patient sustained. In addition, she did not know if the surgical complications were related to the da vinci si surgical system or the traditional laparoscopic surgery. The icb was returned to isi and evaluated. Engineering evaluation was unable to replicate the reported issue by installing the icb on an in-house printed circuit assembly (pca) system and exercising a vessel sealing instrument multiple times. The vessel sealing instrument cauterized, sealed, and cut without any errors. The vessel sealer instrument involved with this complaint has been requested to be returned to isi for evaluation. However, at this time the instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Based on the information provided, this complaint is being reported due to the following conclusion: the initial reporter indicated that the patient sustained an unspecified post-operative complication and was brought back to the or. However, at this time, it is unknown if the da vinci si system, an instrument, or an accessory caused or contributed to the patient's reported injury. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | DA VINCI SI SURGICAL SYSTEM | Reuse | 3289316 | 3260270 | 3156825 | 2955842201302866 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-07-02 | 2013-08-01 | Yes | No | Health Professional | IS3000 A60P8 | Device Returned To Manufacturer | No | 2013-07-25 | Yes | No |  | 2013
| 2013-07-23 | Malfunction |  | During surgery, the physician was trying to coagulate a tissue and noticed that the intuitive surgical pk dissecting forceps(electrosurgical instrument) weren't working. He noticed the pk wire from the device was loose/broken. Broken device was removed and replaced with new one. No harm was noted to patient and no parts of instrument was left inside. No injury to patient or staff. What was the original intended procedure?gyn surgery. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST PK | INSTRUMENT COMPUTER CONTROLLED | 3325940 | 3297768 | 3194363 | 3297768 | 1 | NAY | User Facility |  |  | Initial | 2013-08-01 | 2013-08-01 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
|  | Injury | HospitalizationLife ThreateningDisability | Underwent robotic -davinci- proctectomy. Rectum was perforated and not detected until sepsis infection had developed requiring extended hospital stay and numerous surgeries to clean out the infection and perform a colostomy. |  | DAVINCI | SURGICAL ROBOT | 3297382 | 3265967 | 3162542 | MW5031278 | 1 | NAY | Voluntary |  | Patient | Initial | 2013-07-31 | 2013-07-31 | Yes | Yes | Service Personnel |  | No | No |  |  |  |  | 2013
| 2013-07-25 | Injury | Other | An (b)(6) year old female underwent fiberoptic bronchoscopy with lavage, cervical mediastinoscopy, right thoracoscopy, robotic middle lobectomy and mediastinal lymph node dissection on (b)(6) 2013. Intraoperatively, the surgeon identified that one of the two strut arms of the clip applier was fractured, a piece of metal approx one to two centimeters in size. The surgeon's attempts to locate this metal tip fragment was unsuccessful. |  | DAVINCI S ROBOTIC SMALL CLIP APPLIER 8MM | DAVINCI S ROBOTIC SMALL CLIP A | 3339233 | 3308756 | 3222227 | 3308756 | 1 | DWF | User Facility |  | RISK MANAGER | Initial | 2013-07-29 | 2013-07-31 | Yes | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2013
IS3000 A60P8 | 2013-06-26 | Death | Death | It was reported that during a da vinci si nissen fundoplication procedure, one of the patient's vessels was compromised while the surgeon was performing his last stitch. The bleeding was secured and the surgical procedure was converted to open surgical techniques. The patient subsequently expired in the operating room due to blood loss and pre-existing conditions. | On (b)(4) 2013, intuitive surgical inc. (isi) contacted the surgeon's nurse who was scrubbed in during the surgical procedure and obtained additional information regarding the reported event. According to the nurse, the patient was a high risk patient who had huge cardiovascular issues. The surgical procedure took about 2. 5-3. 0 hours. Towards the end of the surgical procedure, the surgeon was in the process of placing his last stitch when the needle held by a megasuture cut needle driver instrument controlled by his right hand punctured a vein behind the patient's diaphragm. The nurse indicated that the patient had a malformation of the vasculature behind her diaphragm and this was one factor as to why the vein was pierced. Due to the vascular malformation, the surgeon was not expecting the vein to be present behind the patient's diaphragm. Immediately after the vein was pierced, the nurse was able to control the bleeding by clamping down on the damaged vessel with a laparoscopic grasper instrument. Within a minute or two after the vessel was damaged, the surgeon converted the surgical procedure to open surgical techniques as an attempt to repair the vein. The surgeon and a cardiovascular surgeon who was present in the operating room (or) attempted to repair the vessel. However, the nurse indicated that the patient subsequently expired since anesthesia could not get ahead and keep up with the demands of the patient's condition which had rapidly deteriorated. The nurse indicated that the patient's pre-existing medical conditions were another factor as to why the patient's condition rapidly deteriorated. The nurse indicated that no malfunction of the da vinci si system, an instrument, or an accessory caused or contributed to the patient's demise. On (b)(4) 2013, isi contacted the risk management dept. At the site and obtained additional information regarding the reported event. The risk manager indicated that the surgical procedure was not recorded. She also reviewed the operative report of the case and verified that there was no notation that a malfunction of the da vinci si surgical system occurred during the surgical procedure. The risk manager also confirmed that the information provided to isi by the surgeon's nurse on (b)(6) 2013, was correct. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2013. No system errors were found to have occurred during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the surgical staff indicated that the patient expired during a da vinci si surgical procedure. However, at this time there is no evidence that user-error or a malfunction of the da vinci system, an instrument, or an accessory caused or contributed to the patient's demise. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3279665 | 3250790 | 3147345 | 2955842201302782 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-06-27 | 2013-07-29 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-06-27 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-06-26 | Injury | Other | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that a piece about 2-3 mm in length was missing from the tip of the pk dissecting forceps instrument and a piece of cable was also sticking out at the end of the instrument. The surgical staff indicated that a fragment from the instrument fell into the patient but was retrieved. The planned surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) obtained additional information from the site regarding the reported event. The scrub technician indicated that she performed a quick inspection of the pk dissecting forceps instrument prior to the surgical procedure and no issues were observed. The surgical procedure was not recorded. The surgeon reportedly used the pk dissecting forceps instrument for a short period of time before noticing problems with the device. At the time the event occurred, the surgeon was also using a monopolar curved scissors (mcs) instrument. The scrub technician did not recall observing any instrument collisions during the surgical procedure. When the pk dissecting forceps instrument was removed, the scrub technician noticed that a cable was sticking out and torn up. The pk dissecting forceps instrument was replaced with another instrument. A small piece of the cable was found on an unspecified specimen that was removed from the patient. An x-ray was taken per hospital policy and found to be negative for foreign bodies. No post-operative complications were reported. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned to intuitive surgical inc. (isi) and evaluated. Engineering evaluation did not find any fragments missing from the instrument. The pitch up cable was found to be broken at the distal clevis hub. The broken cable segment that contained the crimp was still installed in the instrument's clevis. The clevis did not exhibit any damage or wear marks. The other cables of the instrument's wrist did not exhibit any damage. No other instrument damage was found. Intuitive surgical inc. (isi) has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Based on the information provided, this complaint is being reported due to the following conclusion: the surgical staff indicated that a fragment from the instrument broke off, fell into the patient, and was retrieved. However, at this time, it is unknown how the fragment was retrieved. | PK DISSECTING FORCEPS INSTRUMENT | Reuse | 3279669 | 3250792 | 3147347 | 2955842201302780 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2013-06-26 | 2013-07-29 | Yes | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-06-26 | No | No |  | 2013
400180 | 2013-06-25 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, the surgical staff observed the monopolar curved scissors (mcs) tip cover accessory installed on an mcs instrument fall off and into the patient's pelvic cavity. The mcs tip cover accessory was retrieved and the surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The tip cover has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, intuitive surgical inc. (isi) contacted the initial reporter (a nurse) of this complaint and obtained information regarding the reported event. The nurse indicated that the surgical staff regularly uses electro lube. The nurse was instructed to verify proper use of the electro lube. During the surgical procedure, two patient side manipulators (psm) were in use. The psm is an instrument arm located on the patient side cart (psc) that provides sterile interface for the endowrist instrument. Isi has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon indicated that the mcs tip cover accessory fell off the mcs instrument. The tip cover fell into the patient but was retrieved. However, at this time it is unknown how the tip cover was retrieved. | MONOPOLAR CURVED SCISSORS TIPCOVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3279714 | 3250839 | 3147394 | 2955842201302784 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-06-28 | 2013-07-29 | Yes | Yes | Health Professional | 400180 | No | Yes | 2013-06-28 | Device Not Retu | Yes | Unkown | 2013
IS3000 A60P8 | 2011-04-14 | Injury | Required Intervention | On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci si total hysterectomy and bilateral salpingo-oophorectomy procedure on (b)(6) 2011. The legal complaint alleged that the usage of the da vinci robot resulted in an injury to the bladder, which caused the client to endure over a year of pain, medication, doctors' visits, and finally another surgery to further correct the damage done by the da vinci. | Intuitive surgical inc. (isi) has made several attempts to contact the site and surgeon to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Isi has reviewed the medical records of the patient. The operative report of the patient for (b)(6) 2011, did not reveal that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the surgical procedure. According to the operative report, the bladder flap was then developed without complication, and during the development of the left bladder flap, a defect was noted. It was felt that we had entered the vagina. At this time, we decided to continue the procedure and repair the vaginal defect at the completion of the procedure. After completion of the surgical procedure, the da vinci si system was undocked from the patient. At that point, the surgeon then proceeded to perform a cystoscopy. Based on the cystoscopy, the surgeon noted that it was clear that what we had felt was a vaginal defect was actually an incidental cystotomy. Upon inspection of the patient's bladder, a spillage of methylene blue into the abdominal cavity was observed. A urology surgeon was contacted and the da vinci si system was re-docked to the patient. The urology surgeon repaired the cystotomy utilizing the da vinci si system. After completion of the repair, the urology surgeon performed a cystoscopy to inspect the cystotomy repair. The surgeon noted that the repair appeared to be tight without any evidence of persistent leakage and there was bilateral urine flow from both ureteral orifices. She also indicated that the da vinci system was again undocked and a foley was replaced in the bladder. The operative report indicated that the patient was returned to the dorsal supine position , and the general anesthetic was reversed without complication. The patient was discharged from the hospital the following day, (b)(6) 2011. According to the patient's medical records, the patient had a follow-up appointment in a clinic on (b)(6) 2011, for the cystotomy procedure. A cystogram performed that same day showed that the bladder is healed with no extravasation. The medical record also indicated that the patient was able to urinate part of infused volume into her bladder. On (b)(6) 2012 the patient underwent a cystoscopy procedure for a preoperative diagnosis of recurrent pyelonephritis post cystotomy and history of left flank pain with voiding contraction. According to the operative report, the bladder mucosa was normal except for a second ureteral orifice located posterolaterally with adjacent scar tissue. This gives it the appearance of a reimplanted ureter. The patient was found to have a left ectopic ureter. The patient's medical record from (b)(6) 2012, indicated that the patient had sustained a left ureteral injury during a hysterectomy procedure that was managed with a direct reimplant at the time. The medical record also indicated that since the surgical procedure, the patient complained of chronic pain on the left and she was noted to have vesicoureteral reflux through this ureter. In addition, the medical record indicated that the patient underwent two deflux injections via cystoscopy that were unsuccessful. On (b)(6) 2012, the patient underwent an open intravesical ureteral reimplant. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2011. No system errors were found to have occurred during the reported da vinci surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the patient sustained an injury to her bladder and ureter during a da vinci si surgical procedure. However, at this time, it is unknown how the patient's injuries occurred. It is also unknown if a malfunction of the da vinci si system, an instrument, or an accessory caused or contributed to the patient's reported injuries. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3279783 | 3250892 | 3147447 | 2955842201302783 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-06-28 | 2013-07-29 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-06-28 | Device Not Retu | No | Unkown | 2013
42022703 | 2013-06-21 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, a piece from the tip of the pk dissecting forceps instrument broke off and fell into the patient. The broken instrument piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the yaw pulley was broken. The side of one yaw pulley was missing a. 225 x. 060 piece. The broken piece allowed the grip to move within the pulley and have a larger range of motion in yaw. Engineering evaluation also found that one conductor wire as broken at the yaw pulley exit, on the same side as the yaw pulley breakage. The wire was detached from its connection at the grip. The instrument failed electrical continuity testing. The yaw pulley showed no signs of arcing. The distal end of the instrument's main tube exhibited deep scratches and gouges and has various has various scratch marks with light material removal. The scratches are. 090 -. 290 in length and not aligned with the tube axis. No other damage was found. Engineering concluded that the damage to the main tube was likely due to misuse/mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Intuitive surgical, inc. Has made several attempts to contact the site to obtain additional information regarding the reported event. No additional information has been provided. A follow-up medwatch report will be submitted to the fda if additional information is received. Based on the information provided, this complaint is being reported due to the following conclusion: a piece from the pk dissecting forceps instrument fell into the patient. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3279816 | 3250927 | 3147482 | 2955842201302793 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-06-26 | 2013-07-29 | Yes | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-06-26 | Yes | No | Reuse | 2013
IS3000 A60P6 | 2013-06-25 | Death | Death | It was reported that after completion of a da vinci si myomectomy procedure, while undocking the system from the patient, the patient required resuscitation. The anesthesiologist performed chest compressions on the patient; however, the resuscitation attempts failed and the patient expired. | On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the surgeon who performed the surgical procedure. The surgeon was not available; however, a nurse in his office provided the following additional information: she was not present during the surgical procedure. The patient underwent a routine da vinci myomectomy procedure. After completion of the surgical procedure, the patient's blood pressure dropped due to an unknown reason. The patient coded and resuscitation of the patient was attempted without success. There was no indication that a malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure. There is no video recording of the surgical procedure available. The patient's autopsy report is currently pending. The nurse was unable to provide any other details regarding the reported event. Based on the information provided, isi is unable to determine the root cause of the patient's demise. A follow-up medwatch report will be submitted to the fda if additional information is received. Review of the site's system logs found that there were no system errors generated during the surgical procedure that would have caused or contributed to the patient's demise. This complaint is being reported due to the following conclusion: the patient expired post-op a da vinci surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3279830 | 3250944 | 3147499 | 2955842201302794 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-06-27 | 2013-07-29 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-06-27 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 |  | Injury | Other | It was reported that during a da vinci si surgical procedure, the patient was injured by a patient side manipulator (psm) arm on the patient side cart. The injury sustained by the patient was located outside of the patient's body at the iliac crest. No other information was provided. | Based on the limited information provided regarding the reported event, at this time intuitive surgical is unable to determine the root cause of the injury sustained by the patient. Isi has made multiple attempts to contact the site to verify additional information regarding the reported event. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. This complaint is being reported due to the following conclusion: it was reported that a patient injury occurred during a da vinci surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3279886 | 3251002 | 3147557 | 2955842201302788 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-06-26 | 2013-07-29 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-06-26 | Device Not Retu | No | Reuse | 2013
| 2012-09-18 | Injury | Required Intervention | On (b)(4) 2013, intuitive surgical, inc. (isi) received a legal complaint indicating that a patient who underwent a da vinci surgical procedure on (b)(6) 2012, sustained an intra-operative thermal burn to the bladder, resulting in a laceration that required the placement of a foley catheter and further hospitalization. The legal complaint also alleges that the patient required an intra-operative placement of a transobturator tape (bladder sling) which caused erosion of the patient's vaginal mesh, requiring the patient to undergo two subsequent surgical procedures. Per the information indicated in the legal complaint, the patient continues to suffer from chronic vaginal pain and incontinence, thus requiring the patient to wear bladder pads. The patient also requires further bladder reconstructive surgery and ureteral re-implantation. Reportedly due to the injury sustained by the patient during the surgical procedure, the patient has undergone and continues to undergo multiple painful additional medical tests and procedures, physician consultations and additional surgeries. The patient has suffered and continues to suffer pain, loss of function, emotional distress and permanent injury. | Based on the limited information provided in the legal complaint, intuitive surgical, inc. Is unable to determine the root cause of the injury sustained by the patient. A follow-up medwatch report will be submitted to the fda if additional information is received. This complaint is being reported due to the following conclusion: the patient sustained an injury during a da vinci surgical procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 3279846 | 3251003 | 3147558 | 2955842201302803 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-06-28 | 2013-07-29 | Yes | No | Service Personnel |  | No | No | 2013-06-28 | Device Not Retu | No | Unkown | 2013
IS3000 A60P8 | 2013-06-27 | Other | Other | It was reported that during a da vinci si surgical procedure, the surgical staff indicated that the valley lab electrosurgical (esu) unit would not activate bipolar instruments. The surgeon then made the decision to convert the surgical procedure to open surgical techniques. After the surgery was converted to open, the surgical staff contacted an intuitive surgical inc. (isi) technical support engineer (tse) for assistance. According to the tse, the site did not have a backup esu cable. The tse also indicated that the surgeon did not want to use the external foot pedals and chose to convert to open surgical techniques. | On (b)(4) 2013, an isi field service engineer (fse) performed a field evaluation at the site. The fse discovered that the blue energy cable had a short where the valley lab adapter attaches to the esu cable. The cable short prevented monopolar energy from working. The fse was able to repair the system by replacing the blue energy cable. The fse tested the system and made sure that monopolar and bipolar energy, and the surgeon side console (ssc) pedals were working. On (b)(4) 2013, isi contacted the fse and obtained additional information regarding the reported event. According to the fse, the surgical staff attempted to contact isi for troubleshooting when the reported issue occurred. However, for an unknown reason, the fse indicated that the surgical staff could not get through to contact isi's customer support. The surgeon then made the decision to convert the surgical procedure to open surgical techniques. The fse did not know how long the surgical procedure was in progress before the case was converted to open surgical techniques. The fse verified on (b)(4) 2013, that the site had ordered a replacement energy cable and the issue had not reoccurred. Intuitive surgical inc. (isi) has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2013. A review of the system logs did not reveal any system errors that occurred during the surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon converted the da vinci si surgical procedure to open surgery after the surgical staff encountered an issue with the esu cable. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3279897 | 3251012 | 3147567 | 2955842201302805 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-06-27 | 2013-07-29 | Yes | Yes | Health Professional | IS3000 A60P8 | No | No | 2013-06-27 | Yes | No | Reuse | 2013
42030901 | 2013-03-28 | Injury | Required Intervention | On (b)(4) 2013, intuitive surgical received (b)(4). The event details are provided below: during a robotic-assisted laparoscopic multiple myomectomy, the da vinci mega suturecut needle driver's wire came off inside the patient's abdomen. X-ray taken per hospital policy showed no metallic radiopaque foreign object but very tiny foreign bodies may not be radiographically visible. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, intuitive surgical, inc. Contacted the materials manager regarding the reported event. The materials manager indicated that the fragment from the megasuturecut needle driver instrument fell into the patient and was retrieved from the patient laparoscopically. He was unable to provide any other details regarding the reported event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3279999 | 3251120 | 3147675 | 2955842201302806 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2013-04-29 | 2013-07-29 | Yes | Yes | Health Professional | 42030901 |  | No | 2013-04-29 | Device Not Retu | No | Unkown | 2013
| 2013-07-18 | Malfunction |  | During the case the instrument failed to deliver bipolar energy. The cord was replaced but the problem continued. A new instrument was obtained and the procedure continued without further incident. What was the original intended procedure?robotic assisted hysterectomy. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST | SYSTEMSURGICALCOMPUTER CONTROL | 3302477 | 3273426 | 3170021 | 3273426 | 1 | NAY | User Facility |  |  | Initial | 2013-07-29 | 2013-07-29 | No | Yes | UNKNOWN |  | No | No Answer Provided |  |  |  |  | 2013
IS3000 A60P3 | 2013-06-25 | Other | Other | It was reported that during a da vinci si surgical procedure, when the site installed different instruments onto a patient side manipulator (psm) arm, the instruments were not recognized. Prior to contacting intuitive surgical, inc. For technical support assistance, the surgeon made the decision to complete the planned surgical procedure using semi open techniques. Review of the site's system logs showed that system error code 31009 occurred during the surgical procedure. No patient harm was reported. | The investigation conducted by field service engineering (fse) was unable to replicate the instrument recognition issue experienced by the site. The instruments used during the surgical procedure were not available for the fse to evaluate with the system. However, during functional testing of the site's system using test instruments found that the system functioned within specification. The fse reviewed the site's system logs and found that the site experienced several occurrences of error code 31009 and also experienced system error code 282. System error code 31009 occurs when a tool was fully detected, and then one or more of the tool sensors were not detected for 3 seconds. System error code occurs when the da vinci surgical system detects that the installed instrument on a patient side manipulator (psm) is invalid. The fse concluded that the sterile adapter used during the surgical procedure was likely not seated properly, thus generating system error code 31009. System error code 282 was generated when the site installed an instrument that had no remaining uses. On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the nurse manager who reported the complaint. The nurse manager indicated that there was no injury to the patient as a result of the reported event. No other information regarding the patient was provided. Isi has contacted the surgeon who performed the surgical procedure several times to obtain additional information regarding the reported event. No additional information has been provided. A follow-up medwatch report will be submitted to the fda if additional information is received. As of (b)(4) 2013 there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3276971 | 3248309 | 3144864 | 2955842201302752 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-06-25 | 2013-07-25 | Yes | Yes | Health Professional | IS3000 A60P3 | Yes | No | 2013-06-25 | Yes | No | Reuse | 2013
| 2013-07-15 | Malfunction |  | Intuitive surgical robotic instrument noted to have frayed cable on standard pre-procedure inspection. Instrument did not have patient use. Pulled immediately. Multi-use 10x instrument. |  | DAVINCI ROBOT | PROGRASP FORCEPS | 3292907 | 3262526 | 3159081 | 3262526 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-07-17 | 2013-07-25 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2013
42030901 | 2013-06-24 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the surgical staff indicated that a cable broke on a megasuture cut needle driver instrument and a 2cm fragment fell into the patient. | On (b)(4) 2013, intuitive surgical inc. (isi) contacted the nurse who initially reported this complaint and obtained additional information regarding the reported event. The nurse indicated that the instrument was inspected prior to the surgical procedure and no issues were observed. She estimated that the hysterectomy procedure took about 2 hours to complete and the instrument was in use during most of the procedure. Near the end of the surgical procedure, the surgical staff noticed that the wires near the wrist of the instrument were broken and frayed. According to the nurse, the surgical staff did not see any fragments fall into the patient. Contrary to the initial report, no fragment was detected to have fallen into the patient. However, due to hospital policy, the surgical staff performed an x-ray to look for potential fragments that might have fallen into the patient. The x-ray did not find any fallen fragments. The nurse indicated that the patient did not experience any intra-operative or post-operative complications. She denied that there were any instrument collisions during the surgical procedure. The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Based on the information provided, this complaint is being reported due to the following conclusion: the surgical staff indicated that a cable on the instrument broke and an x-ray was performed to search for potential fragments that might have fallen into the patient. The x-ray performed on the patient did not find any fallen fragments. | MEGASUTURE CUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3274558 | 3246180 | 3142737 | 2955842201302718 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-06-24 | 2013-07-24 | Yes | Yes | Health Professional | 42030901 | No | Yes | 2013-06-24 | Device Not Retu | No | Unkown | 2013
41032204 | 2013-06-21 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the surgeon indicated that the endowrist one vessel sealer instrument was not sealing. Due to the reported issue, the surgeon made the decision to convert the surgical procedure to open surgical techniques. | On (b)(6) 2013, intuitive surgical inc. (isi) contacted the surgeon who performed the da vinci si hysterectomy procedure involved with this complaint and obtained additional information regarding the reported event. The surgeon indicated that surgical procedure was in progress for about 10-15 minutes before he decided to convert to open surgical techniques. The surgeon indicated that he decided to convert to open surgical techniques because the endowrist one vessel sealer instrument was not sealing and the patient had a huge uterus. The instrument was returned to isi and evaluated. Engineering evaluation was unable to replicate the reported customer complaint that the instrument would not seal. A visual inspection of the instrument showed no damage to the conductor wires. The instrument passed an electrical continuity test from pins to grips. The instrument was also placed on an in-house system for a functional test. The instrument passed an energy test. The system logs were reviewed by isi and 3 attempts at sealing were observed. The seal pedal was pressed for approximately 10, 2, and 4 seconds. The system logs did not show any sealing related errors. No instrument damage was found. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon converted the da vinci si hysterectomy procedure to open surgical techniques. However, at this time, there is no evidence that an instrument malfunction occurred during the da vinci si surgical procedure. The instrument involved with this complaint was returned to isi and no trouble was found with the evaluation of the device. | ENDOWRIST ONE VESSEL SEALER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3274560 | 3246183 | 3134907 | 2955842201302720 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-06-24 | 2013-07-24 | Yes | No | Health Professional | 41032204 | Device Returned To Manufacturer | No | 2013-06-24 | Yes | Yes | Initial | 2013
SH1531 | 2013-07-01 | Malfunction |  | During a procedure, the robot malfunctioned with a non'-recoverable" fault. Intuitive biomed 800# called. The robot hand was clamped onto the patients bladder and would not release. They suggested to retrieve a particular wrench to open instrument jaws. It took a few minutes to find this "special" sterile allen wrench. When located, staff were able to use the tool to release the clamp. No patient harm. What was the original intended procedure?prostatectomy. Device #1is this a laboratory device or laboratory test?no. |  | DA VINCI SI SURGICAL SYSTEM | SYSTEMSURGICALCOMPUTER CONTROL | 3277495 | 3248803 | 3145358 | 3248803 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-07-24 | 2013-07-24 | No | Yes | Invalid Data | SH1531 | Yes | No Answer Provided |  |  |  |  | 2013
SH1531 | 2013-07-01 | Malfunction |  | During a procedure, the robot malfunctioned with a non'-recoverable" fault. Intuitive biomed 800# called. The robot hand was clamped onto the patients bladder and would not release. They suggested to retrieve a particular wrench to open instrument jaws. It took a few minutes to find this "special" sterile allen wrench. When located, staff were able to use the tool to release the clamp. No patient harm. What was the original intended procedure?prostatectomy. Device #1is this a laboratory device or laboratory test?no. |  | DA VINCI SI SURGICAL SYSTEM | SYSTEMSURGICALCOMPUTER CONTROL | 3277495 | 3248803 | 3145358 | 3248803 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-07-24 | 2013-07-24 | No | Yes | Invalid Data | SH1531 | Yes | No Answer Provided |  |  |  |  | 2013
VER06 | 2013-07-22 | Malfunction |  | A portion of the brown circumferential plastic piece below the spatula of the da vinci permanent cautery spatula broke off of the instrument in the pt. The broken piece was recovered from the pt in its entirety. Diagnosis or reason for use: laparoscopic robotic total hysterectomy. |  | DAVINCI PERMANENT CAUTERY SPATULA 8MM | DAVINCI SPATULA | 3281294 | 3249266 | 3145821 | MW5031058 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-07-22 | 2013-07-22 | No | Yes | Health Professional | VER06 | Yes | Yes |  |  |  |  | 2013
IS2000 A51P8 | 2010-08-19 | Injury | Other | On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci s hysterectomy procedure on (b)(6) 2010. The legal complaint alleged that the da vinci si surgical system malfunctioned and the patient sustained severe, serious, and permanent injuries. In addition, the legal complaint alleged that the patient suffered permanent injuries, severe incapacities and mental anguish. | Intuitive surgical inc. (isi) has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2010. A review of the system logs revealed that no system errors occurred during reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the legal complaint claimed that the da vinci s surgical system malfunctioned and the patient sustained severe and permanent injuries. However, at this time, it is unknown if the da vinci s surgical system caused or contributed to the patient's injuries. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3266833 | 3237918 | 3134474 | 2955842201302649 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-06-24 | 2013-07-19 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2013-06-24 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-06-19 | Other | Other | It was reported that during a da vinci si prostatectomy procedure, the surgical staff encountered repeated system errors 23013 and 23008. The surgical staff contacted an intuitive surgical inc. (isi) technical support engineer (tse) for troubleshooting. The surgical staff used an emergency wrench to remove a cobra prograsp instrument and the system was rebooted. At that point, the system faulted with an error 23008. The tse reviewed the system logs and found repeated 23013 and 23008 errors referencing the left master tool manipulator (mtml), axis 6. The surgical staff emergency powered off the surgeon side console (ssc) and repeated exercising the mtml. However, the reported issue was not resolved and the system could not complete homing. Due to the unresolved system error message issues, the surgeon made the decision to convert the surgical procedure to open surgery. | On (b)(6) 2013, an isi field service engineer (fse) performed a field evaluation at the site and repaired the system by replacing the left gimbal on the mtml. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was then tested and verified to manufacturer specifications. The gimbal was returned to isi and evaluated. Engineering evaluation indicated that a visual inspection of the gimbal revealed a broken-off screw head on the axis 6 bevel gear that would have caused the reported error. On (b)(6) 2013, isi contacted the initial reporter of this complaint and obtained additional information regarding the reported event. The initial reporter indicated that the system errors began to occur around 4 hours into the surgical procedure. She stated that they were already near the end of the da vinci si prostatectomy procedure. Since the system error message issues could not be resolved with troubleshooting, the surgeon made the decision to complete the surgical procedure via open surgical techniques. The initial reporter denied that the patient experienced any intra-operative or post-operative complications. She also indicated that since the repair the da vinci si system was functioning properly and there had been no recurrences of the system errors. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon converted the da vinci prostatectomy procedure to open surgical techniques after encountering repeated system errors. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3267229 | 3238333 | 3134888 | 2955842201302652 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-06-19 | 2013-07-19 | Yes | Yes | Health Professional | IS3000 A60P8 | Device Returned To Manufacturer | Yes | 2013-06-19 | Yes | No | Reuse | 2013
41032204 | 2013-06-18 | Other | Other | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the surgeon indicated that the endowrist vessel sealer instrument would not cut. According to the initial reporter, the instrument would seal. It was reported that when the cut pedal would be pressed, the surgeon side console (ssc) would just give a warning beep. Due to the reported instrument issue, the surgeon made the decision to convert the surgical procedure to open surgical techniques. | Manufacturer Narrative: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) contacted the surgeon who performed the da vinci si hysterectomy procedure involved with this complaint and obtained additional information regarding the reported event. The surgeon indicated that the reported issue occurred on (b)(6) 2013. --/-- Additional Manufacturer Narrative: --/-- Intuitive surgical inc. (isi) has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. On (b)(4) 2013, an isi field service engineer (fse) performed a field evaluation at the site. The fse checked for vessel sealer instrument function by installing a test instrument on the system. The fse did not observe any issues with sealing or cutting with the test instrument. The fse tested and verified the system to manufacturer specifications. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2013. No system errors were found to have occurred during the reported surgical procedure. On (b)(4) 2013, isi contacted the clinical sales representative (csr) who reported this complaint and obtained additional information regarding the reported event. The csr indicated that the surgeon encountered the cutting issue with the endowrist vessel sealer instrument about 10 minutes into the surgical procedure. The csr indicated that the surgical staff did not attempt to contact isi for troubleshooting the reported issue. Due to the instrument issue, the surgeon made the decision to convert the surgical procedure to open surgical techniques. According to the csr, the patient did not experience any intra-operative or post-operative complications. The csr stated that he consulted with the site and determined that the surgeon was not properly using the pedal for cutting on the surgeon side console (ssc). The csr instructed the surgical staff and surgeon on proper use of the ssc cut pedal. As of the date of this report, there have been no reported recurrences of the instrument cutting issue. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon encountered a cutting issue with the endowrist vessel sealer instrument and made the decision to convert the surgical procedure to open surgical techniques. However, there is no evidence that an instrument malfunction occurred during the da vinci si surgical procedure. | ENDOWRIST ONE VESSEL SEALER INSTRUMENT | Unkown | 3267253 | 3238351 | 3134907 | 2955842201302653 | 1 | NAY | Manufacturer | Company RepresentativeHealth Professiona | Other | InitialFollowup | 2013-06-21 | 2013-07-19 | Yes | No | Health Professional | 41032204 | No | No | 2013-07-22 | No | Yes |  | 2013
IS3000 A60P8 | 2013-06-19 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, when the surgeon took control of the master tool manipulators (mtm) at the surgeon side console, the patient side manipulator arm 1 moved to the side. The site contacted intuitive surgical, inc. (isi) for technical support assistance. While the surgical staff was attempting to trouble-shoot the issue with the technical support engineer (tse), the surgeon made the decision to complete the surgical procedure using open surgical techniques. The tse was told by the site that the patient had slipped on the patient table approximately 2 inches. The tse explained to the surgical staff that the pressure placed on the cannula due to movement of the patient caused the psm to move when the surgeon manipulated the mtms. The tse advised the site that pressing the clutch buttons on the psm and moving the psm would have resolved the issue. No patient harm or injury was reported. | The investigation conducted by the field service engineer (fse) concluded that the surgeon was unable to take control of the master tool manipulators (mtms) during the surgical procedure due to the torque placed on the patient side manipulator (psm) arm when the patient moved. The fse re-trained the site on how press the clutch buttons on the psm and move the psm to relieve the pressure on the psms should the issue recur during future da vinci surgical procedures. On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the initial reporter of this complaint. The initial reporter indicated that the surgical procedure was in progress approximately 20 minutes when the reported issue occurred. Per the initial reporter only the patient's abdomen moved. The patient had a heavy abdominal pannus. The patient was at a level position when the port site incisions were created. After creation of the port incisions, the patient was placed in the trendelenburg position, thus causing the patient's abdomen to shift. Per the initial reporter, there have been no other occurrences of the issue at hospital since the surgical staff was re-trained on how to relieve pressure on the psm. The da vinci si user manual specifically states: warning: once the da vinci si system is positioned, cannulae are placed in the patient and patient cart arms are attached to the cannulae, the surgical table must not be moved in any way. Serious injury could result. If intraoperative or table movement is necessary, undock the da vinci si system, move the or table and re-dock the system. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon made the decision to convert to open surgical techniques after observing movement of the psm due to adjustment of the patient position during the surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3267377 | 3238476 | 3135032 | 2955842201302662 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-06-19 | 2013-07-19 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-06-19 | Device Not Retu | No | Reuse | 2013
IS3000 A60P5 | 2013-06-19 | Injury | Other | It was reported that during a da vinci si cholecystectomy procedure, the cannula mount on the patient side manipulator (psm) arm would not stay closed. It was also reported that the patient's gallbladder, which was to be removed as part of the surgical procedure was nicked. The planned surgical procedure was completed; however; due to the cannula mount not staying closed, the surgical time was lengthened approximately 2 hours. | The investigation conducted by the field service engineer (fse) confirmed the customer reported failure mode. The fse observed that the cannula mount on the patient side manipulator arm was loose, thus the clamps on the cannula mount would not stay closed. The fse found that the cannula mount did not exhibit any physical damage and concluded that the cannula mount became loose due to wear and tear. The cannula mount accessory is attached to the psm and is designed to hold the cannula securely in place, outside of the patient's cavity. The system was repaired by replacing the affected psm cannula mount accessory. It is unknown if the cannula mount issue is related to the reported nicking of the patient's gallbladder. Intuitive surgical, inc. Has contacted the site several times to verify additional information. A follow-up mdr will be submitted if additional information is received. Based on the information provided, this complaint is being reported due to the following conclusion: the patient's gall bladder was nicked during a da vinci surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3267380 | 3238488 | 3135043 | 2955842201302663 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-06-19 | 2013-07-19 | Yes | Yes | Health Professional | IS3000 A60P5 | Yes | No | 2013-06-19 | Yes | No | Reuse | 2013
IS3000 A60P8 | 2013-06-19 | Other | Other | It was reported that during a da vinci si pyeloplasty procedure, the surgeon was unable to take control of the instrument installed on patient side manipulator (psm) arm 2 while manipulating the master tool manipulators (mtm) on the surgeon side cart. The site replaced the installed instrument; however, the issue recurred. The surgeon switched to psm 3; however, the issue persisted. Unable to resolve the issue, the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The investigation conducted by the intuitive surgical, inc. Field service engineer (fse) was unable to replicate the customer reported failure mode. The cable tension on patient side manipulator (psm) arm 2 and psm 3 were found to be within specification and the system passed functional testing with no issues noted. The fse reviewed the site's system logs and found that no system errors were generated that would have caused or contributed to the issue experienced by the site. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. On (b)(4) 2013, isi contacted the initial reporter of this complaint. The initial reporter indicated that there were no system error codes or warnings noted when the issue occurred. The surgical staff contacted their isi clinical sales representative (csr) for trouble shooting assistance to resolve the issue. The site was advised by the csr to contact isi's technical support engineering for trouble shooting assistance. The initial reporter indicated that the surgeon declined to contact technical support and made the decision to convert to an open procedure. There was no injury or harm to the patient and the patient tolerated the open procedure well. This complaint is being reported due to the following conclusion. The surgeon was unable to gain control the psms, causing the surgeon to convert to an open procedure. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3267392 | 3238536 | 3135092 | 2955842201302665 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-06-19 | 2013-07-19 | Yes | Yes | Health Professional | IS3000 A60P8 | Yes | No | 2013-06-20 | Yes | No | Reuse | 2013
41032204 | 2013-06-21 | Injury | HospitalizationOther | It was reported that during a da vinci si hemicolectomy procedure, after sealing and cutting the patient's right colic artery using the endowrist one vessel sealer instrument, after the surgeon opened the jaws on the instrument, he observed that the patient's vessel was not sealed. The patient bled from both of the cut ends and the surgeon made 3 to 4 additional attempts to seal the affected vessel; however, his attempts were unsuccessful. The surgeon made the decision to convert the planned surgical procedure to an open procedure to control the bleeding experienced by the patient and to complete the planned surgical procedure. | The endowrist one vessel sealer instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the isi clinical sales representative (csr) who reported this complaint and was present during the surgical procedure. The csr indicated that the instrument was inspected prior to use and there were no issues noted. There was no indication that the vessel was not sealed, as audible tones were noted to have occurred while the surgeon was activating cautery energy and after the sealing cycle was completed. The csr indicated that the surgeon was able to control the patient's bleeding using two traditional laparoscopic grasper instruments. The patient lost 500ml of blood. On (b)(4) 2013, isi contacted the surgeon who performed the surgical procedure. The surgeon indicated that there was no indication to him that the patient's right colic artery was not sealed, since he observed tissue being cauterized as he activated cautery energy. Audible tones from the electrosurgical generator were also noted during the cautery and cut cycles. After opening the jaws on the instrument he observed that the patient's vessel was partially sealed. The surgeon indicated that the patient's right colic artery was 3mm in diameter, did not exhibit calcification and was found to be normal. The surgeon indicated that after controlling the patient's bleeding, he undocked the da vinci surgical system and performed an open laparotomy to complete the planned surgical procedure. Per the surgeon, the patient tolerated the open procedure perfectly fine and did not require a blood transfusion. The patient was discharged to a nursing home facility 5 days post-op. The patient's hospitalization was extended for 1 day. The surgeon indicated that it is his belief that the patient's vessel was not properly sealed due to the instrument was defective. The surgeon also indicated that the patient returned to the hospital on (b)(6) 2013, due to experiencing abdominal pain. Per the surgeon, the patient developed a wound infection in the laparotomy incision. The patient is currently being treated with antibiotics. This complaint is being reported due to the following conclusion: the planned surgical procedure was converted to open to control bleeding experienced by the patient, after the surgeon cut the patient's vessel that was not completely sealed during use of the endowrist one vessel sealer instrument. | ENDOWRIST ONE VESSEL SEALER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3267442 | 3238551 | 3135106 | 2955842201302669 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-06-21 | 2013-07-19 | Yes | Yes | Health Professional | 41032204 | No | No | 2013-06-21 | Device Not Retu | No | Initial | 2013
| 2013-03-06 | Injury | DisabilityHospitalization | Please note the 5 cases being reported concurrently by our facility have the d'vinci robot as a commonality; we are not implying that it is the cause. Our facility has two d'vinci robots. We are uncertain which of the two was used in this case. On (b)(6) 2013, pt underwent d'vinci pelviscopy with laser destruction of endometriosis, also had cystoscopy and proctoscopy. Discharged. On (b)(6) 2013, pt had ct-guided aspiration of pelvic abscess adjacent to r ovary. Discharged to home. On (b)(6) 2013, readmitted, went to operating room: exploratory laparoscopy converted to open procedure with finding of inflammation and adhesions, dense involvement of small intestine within pelvis affiliated with abscess. Abscess drained, lysis of adhesions done, and small bowel resected with reanastomosis. Discharged in stable condition on (b)(6) 2013. |  | DVINCI SI | DVINCI ROBOT | 3270923 | 3240434 | 3136989 | 3240434 | 1 | NAY | User Facility |  | Nurse | Initial | 2013-07-16 | 2013-07-16 | Yes | No | Health Professional |  | Yes | Yes |  |  |  |  | 2013
| 2013-05-10 | Death | Death | Please note the 5 cases being reported concurrently by our facility have the d'vinci robot as a commonality; we are not implying that it is the cause. Our facility has two d'vinci robots. We are uncertain which of the two was used in this case. On (b)(6) 2013, pt underwent a d'vinci combined total hysterectomy, salpingo-oophorectomy and node dissection. Pt also had a liver biopsy. By (b)(6) 2013, pt in septic shock with multi-organ failure. Her draining abdominal wound cultured candida, cn staph, anaerobes and diphtheroids. She had a staph. Positive blood culture. Source of sepsis not clear. Pt was on multiple gtts to support her hemodynamics and multiple antibiotics. Pt expired in icu on (b)(6) 2013. |  | DVINCI SI | DVINCI ROBOT | 3270944 | 3240436 | 3136991 | 3240436 | 1 | NAY | User Facility |  | Nurse | Initial | 2013-07-16 | 2013-07-16 | Yes | No | Health Professional |  | Yes | Yes |  |  |  |  | 2013
| 2013-04-10 | Injury | DisabilityHospitalization | Please note the 5 cases being reported concurrently by our facility have the d'vinci robot as a commonality; we are not implying that it is the cause. Our facility has two d'vinci robots. We are uncertain which of the two was used in this case. On (b)(6) 2013, pt underwent d'vinci pelviscopy with excision of endometriosis. During procedure, it was noticed that pt had adhesions, so surgeon was called in to do lysis of abdominal pelvic adhesions. Pt was discharged to home (b)(6) 2013. On (b)(6) 2013, pt was readmitted with abdominal pain. Ct of abd and pelvis fairly unremarkable; pt observed in hospital. With no fevers, well-controlled pain, abd wound clean and dry, pt was discharged on (b)(6) 2013. Pt was readmitted the same day, (b)(6) 2013, thru er with abdominal and l shoulder pain, and fever. Ct of pelvis and abd showed significant changes. On (b)(6) 2013, pt underwent open laparoscopic exploration with repair of sigmoid colon perforation. Pt was discharged on (b)(6) 2013, afebrile and in stable condition. |  | DVINCI SI | DVINCI ROBOT | 3271022 | 3240437 | 3210884 | 3240437 | 1 | NAY | User Facility |  | Nurse | Initial | 2013-07-16 | 2013-07-16 | Yes | No | Health Professional |  | Yes | Yes |  |  |  |  | 2013
| 2013-05-22 | Injury | DisabilityHospitalization | Please note the 5 cases being reported concurrently by our facility have the d'vinci robot as a commonality; we are not implying that it is the cause. Our facility has two d'vinci robots. We are uncertain which one was used in this case. On (b)(6) 2013, pt underwent a d'vinci pelviscopy, laser excision of endometriosis, d and c, hysteroscopy and cystoscopy. Discharged in stable condition. She later developed a temperature of 102 degrees f, nausea, vomiting, and diarrhea and presented to another hospital on (b)(46 2013. There she had a ct that showed on abscess on her uterus and possibly a small tear on her rectum, was treated with iv antibiotics and admitted for 2 days. Rectal tear evaluated by gi. She was discharged on oral antibiotics. Pt had outpt ct which showed abdominal abscess was draining on its own, thus pt did not require surgery. However, pt continued to have constant diarrhea, nausea, vomiting, fever, and pelvic pain, and was readmitted on (b)(6) 2013. Found to have ileus. Was npo and on antibiotics. Ileus and pelvic infection resolved, and pt discharged on (b)(6) 2013 in stable condition. |  | DVINCI SI | DVINCI ROBOT | 3271068 | 3240445 | 3137001 | 3240445 | 1 | NAY | User Facility |  | Nurse | Initial | 2013-07-16 | 2013-07-16 | Yes | No | Health Professional |  | Yes | Yes |  |  |  |  | 2013
| 2013-04-30 | Injury | DisabilityHospitalization | Please note the 5 cases being reported concurrently by our facility have the d'vinci robot as a commonality; we are not implying that it is the cause. Our facility has two d'vinci robots. We are uncertain which one was used in this case. On (b)(6) 2013, pt had d'vinci prostatectomy for treatment of prostate cancer. While undergoing this, adhesions noticed in area of apparent old appendectomy. Surgeon brought in to do laparoscopic lysis of adhesions. Pt discharged (b)(6) 2013. On (b)(6) 2013, pt presented to ed for abdominal pain. Ct done in er consistent with findings related to the pt's recent surgery; pt sent home. Pt had worsening generalized weakness, right lower abdominal pain, fevers and nausea and had a ct at another facility which returned as findings consistent with intraabdominal abscess. On (b)(6) 2013, admitted thru er for fevers and new ct findings which showed multiple abd and pelvic abscess collections. In interven radiology, had drainage of abscesses and jps placed. Cultures of drainage grew out lactobacillus. Blood cx neg. Pt improved clinically, and was discharged on iv antibiotics per picc line, and orders for weekly labs; discharged on (b)(6) 2013 in stable condition. |  | DVINCI SI | DVINCI ROBOT | 3271133 | 3240447 | 3137003 | 3240447 | 1 | NAY | User Facility |  | Nurse | Initial | 2013-07-16 | 2013-07-16 | Yes | No | Health Professional |  | Yes | Yes |  |  |  |  | 2013
42017912 | 2013-06-11 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, the monopolar curved scissors (mcs) instrument collided with the maryland bipolar forceps instrument causing a piece of the blade to break off and fall into the patient. The surgeon searched for the broken instrument piece; however, it was not located. The planned surgical procedure was completed. No patient harm was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported complaint. A microscopic inspection of the instrument's blade found that the blade had a chipped surface that was unrelated to normal wear and tear. The surface was no longer straight, the surface was uneven and material was missing. The instrument failed a cut test due to the damaged blade. Engineering concluded that the damage was likely due to misuse/mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3251673 | 3222809 | 3119364 | 2955842201302523 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-06-11 | 2013-07-12 | Yes | Yes | Health Professional | 42017912 | Device Returned To Manufacturer | No | 2013-07-17 | No | No | Initial | 2013
IS3000 A60P8 | 2011-06-09 | Injury | Required InterventionHospitalization | On (b)(6) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci si prostatectomy procedure on (b)(6) 2011. The legal complaint claimed that the robotic device is defective and that during the course of the surgery, the patient sustained significant injuries to this intestines and small bowels by the use of the da vinci surgical robot. In addition, the legal complaint indicated that the robotic surgery was converted to an open laparotomy surgery with repair of lacerations and tears of the intestines and small bowels. The legal complaint also claimed that due to the lacerations and tears to his small bowels by the da vinci robotic surgery, the patient would spend the next forty-four days in the hospital. According to the legal complaint, the patient developed fevers, confusion, abdominal distension and tenderness due to his injuries, and also bouts of sepsis, suffered abscesses and giant fistula. As a result of the injuries, the legal complaint also indicated that the patient had to have multiple additional operations and has suffered pain, loss of function, emotional distress, and permanent injury. | Intuitive surgical inc. (isi) has made several attempts to contact the risk management department of the site and the surgeon who performed the da vinci si prostatectomy procedure to obtain additional information concerning the reported event; however, no additional information has been provided by the risk management department or the surgeon as of the date of this report. A follow-up mdr will be submitted if additional information is received. On (b)(4) 2013, isi obtained additional information regarding the reported event from the practice administrator of the surgeon's office. The practice administrator indicated that she read the operative report for the da vinci si surgical procedure involved with this complaint and did not see any documentation that a malfunction of the da vinci surgical system occurred. She read one part of the operative report where the surgeon noted, i opened a piece of small intestine, while taking down an adhesion. According to the practice administrator, the operative report did not explain how or why the patient's small intestine was opened. The small intestine opening was immediately repaired by the surgeon. According to the patient's medical records, the patient underwent the following procedures on 6/13/2011: exploratory laparotomy, extensive lysis of adhesions, a small bowel resection, and insertion of a negative pressure dressing. On (b)(6) 2011, the patient also underwent the following procedures via open surgery: a second look re-exploratory procedure, drainage of enterotomy, placement of pelvic drains, and insertion of a negative pressure dressing. The practice administrator did not know the patient's current status. The patient's medical records indicated that his last visit to the hospital was on (b)(6) 2011. However, the practice administrator did not provide information regarding the patient's hospital visit on (b)(6) 2011. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2011 and no related system errors were found to have occurred during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the legal complaint alleged that the patient sustained injuries to his small bowel and intestines while undergoing a da vinci si surgical procedure. However, at this time, it is unknown if a malfunction of the da vinci si system, an instrument, or an accessory caused or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3253745 | 3224947 | 3121502 | 2955842201302539 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-06-14 | 2013-07-12 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-06-14 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-06-11 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, the surgical staff indicated that the tip of the permanent cautery spatula instrument broke off and fell into the patient. The initial reporter of this complaint indicated that the surgical staff was able to retrieve both pieces that fell into the patient. The planned surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation confirmed that a piece of the instrument's spatula tip measuring roughly about 0. 134 broke off but the fragment was not returned. The instrument passed a electrical continuity test. Engineering evaluation concluded that the damage was most likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Based on the information provided, this complaint is being reported due to the following conclusion: the surgical staff claimed that the tip of permanent cautery spatula instrument broke off and fell into the patient. However, at this time, it is unknown how the fragment(s) were retrieved. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3253816 | 3224994 | 3121549 | 2955842201302540 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-06-14 | 2013-07-12 | Yes | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-06-14 | Yes | No | Reuse | 2013
42022703 | 2013-04-26 | Injury | Other | On (b)(6) 2013, intuitive surgical, inc. (isi) received medwatch report with the following information: j. H. Underwent the following surgical procedure - robot-assisted total laparoscopic hysterectomy bilateral salpingo-oophorectomy and laparoscopic appendectomy. The pk da vinci dissecting forceps grasper portion of the instrument broke on the jaw of the grasper and one side of the grasper fell free into the abdomen. The 1 cm x 5 mm piece was successfully retrieved. A post-operative abdominal x-ray was performed to confirm foreign body was removed completely. This device was on its 1st of ten lives. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Based on the information provided, this complaint is being reported due to the following allegation: one of the jaws on the pk dissecting forceps instrument broke off and fell into the patient. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3253964 | 3225165 | 3121720 | 2955842201302542 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2013-06-14 | 2013-07-12 | Yes | Yes | Health Professional | 42022703 | No | No | 2013-06-14 | Device Not Retu | No | Initial | 2013
IS3000 A60P8 | 2013-06-09 | Injury | Required Intervention | It was reported that during a da vinci si prostatectomy procedure, with the use of dual surgeon side consoles (sscs), during dissection the patient's urethra, the surgeon experienced resistance on patient side manipulator (psm) arms 1 and 2 while manipulating the master tool manipulators (mtms) on the surgeon side carts (ssc). The surgeon continued to manipulate the mtms when psm 1 suddenly moved, causing the installed monopolar curved scissors instrument to cut the patient's urethra. The surgeon reconstructed the affected area to repair the damage to the patient's urethra and the planned surgical procedure was completed with the da vinci surgical system. | The field investigation conducted by the intuitive surgical, inc. (isi) field service engineer (fse) was unable to replicate the issue experienced by the site during functional testing of the site's system using test instruments, a test drape and cannula. However, after powering off the system, and moving axis 1 of patient side manipulator (psm) arm 1 the fse noted that axis 1 brake would not move, even after pressing the clutch button on the psm. The fse burnished the axis, however, the issue persisted. Axis 1 on psm 2 and psm 3 were also burnished; however, the fse was able to move the psms. Review of the site's system logs by the fse found no system error codes were generated that caused or contributed to the issue experienced by the site. Psm 1 was replaced to repair the system. As a precaution, psm 2 was also replaced. Psm 1 and psm 2 have been returned to isi for failure analysis evaluation. At this time, the root cause of the customer reported failure mode has not been determined, as evaluation of the affected parts is currently underway. A follow-up medwatch report will be submitted to the fda after the evaluations have been completed or if additional information is received. The monopolar curved scissors (mcs) instrument installed on psm 1 during the surgical procedure was returned to isi for failure analysis evaluation. Engineering evaluation was unable to replicate the sudden movement experienced by the site. The instrument passed recognition and engagement testing while installed on an is3000 in-house system. The instrument moved intuitively with a full range of motion in all directions and the grips opened and closed properly. On (b)(4) 2013 isi contacted the initial reporter of this event. The initial reporter indicated that the patient was in stable condition and had not yet been released from the hospital. No additional information concerning the patient's status has been provided. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3244097 | 3215235 | 3111790 | 2955842201302484 | 1 | NAY | Manufacturer | DistributorOther | Other | InitialFollowup | 2013-06-09 | 2013-07-10 | Yes | Yes | Health Professional | IS3000 A60P8 | Device Returned To Manufacturer | No | 2013-07-16 | No | No | Reuse | 2013
IS3000 A60P5 | 2013-06-04 | Injury | Required InterventionHospitalization | It was reported that 2 days post op a successful da vinci si left salpingo-oophorectomy procedure, performed on (b)(6) 2013, the patient returned to the hospital's emergency room on (b)(6) 2013 with symptoms of abdominal pain. The patient underwent an exploratory laparotomy and it was discovered that the patient had free air in her abdomen and her sigmoid colon was damaged at the site of dissection during the initial surgical procedure. | On (b)(4) 2013, intuitive surgical, inc. Contacted the surgeon who performed the surgical for additional information. The surgeon indicated that there were no intra-operative complications experienced by the patient during the da vinci left salpingo-oophorectomy procedure. The patient was discharged from the hospital on (b)(6) 2013 and prior to the patient's release from the hospital the patient had not experienced any post-surgical complications. The surgeon indicated that no malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure and that it was her belief that patient's injury was unrelated to the da vinci surgical system, instruments or accessories, but rather due to dense pelvic adhesions. During the (b)(6) 2013 surgical procedure, it was discovered that the patient's ovary was stuck to her left sigmoid colon. The surgeon indicated that the damage to the patient's sigmoid colon was a tiny perforation that likely occurred while taking down the patient's adhesions and ovary. The surgeon stated that it was her belief that the due to the patient's extensive adhesions, the injury sustained by the patient, may have occurred if she had performed the surgical procedure using traditional laparoscopic techniques or traditional open surgical techniques. A partial sigmoid resection with primary re-anastomosis was performed to repair the damage to the patient's sigmoid colon. The patient was discharged from the hospital on (b)(6) 2013 and is recovering slowly. Isi reviewed the site's system logs for the reported procedure date. No system error codes were generated that would have caused or contributed to the injury sustained by the patient. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3244188 | 3215389 | 3111944 | 2955842201302487 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2013-06-10 | 2013-07-10 | Yes | No | Health Professional | IS3000 A60P5 | No | Yes | 2013-06-10 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2010-10-14 | Injury | Other | On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci si bilateral salpingo oophorectomy and hysterectomy procedure on (b)(6) 2010. The legal complaint alleged that the patient sustained damage to her ureters and suffered severe and permanent injuries including severe internal injuries. In addition, the legal complaint alleged that she sustained pain and suffering, disability, physical impairment, mental impairment, disfigurement, mental anguish, inconvenience and loss of capacity for the enjoyment of life, all of which are continuing and are permanent in nature. She has incurred medical and related expenses and will continue to incur such expenses in the future. She has lost wages in the past and has suffered a loss of earning capacity which is continuing and permanent in nature. | On (b)(4) 2013, intuitive surgical inc. (isi) contacted the risk management department at the site and obtained additional information regarding the reported event. The risk manager indicated that no malfunction of the da vinci system, an instrument, or an accessory occurred during the surgical procedure. The risk manager mentioned that the patient had subsequent admissions and unspecified surgical procedures performed at one or more different hospitals. However, she was unable to provide any additional information regarding the reported event or the patient's subsequent admissions to the other hospital(s). Isi has reviewed the operative report of the patient and there was no complaint that a malfunction of the da vinci si surgical system occurred during the reported surgical procedure. No intra-operative complications related to the patient's ureters were reported. The operative report indicated that at the end of the surgical procedure, irrigation was carried out. All the counts were correct. Hemostasis was adequate. All the robotic instruments were removed under direct visualization. All trocars were removed under direct visualization. Also, after completion of the surgical procedure, the patient was taken to the recovery room in satisfactory condition. Isi has also reviewed the system logs for the site with a procedure date of (b)(6) 2013, and no system errors were found to have occurred during 3 da vinci surgical procedures performed that day. In addition, all instruments used during the surgical procedures that day were found to have been reprocessed and used again in subsequent surgeries with no complaints. Based on the information provided, this complaint is being reported due to the following conclusion: the patient claimed that she underwent a da vinci si surgical procedure and sustained damage to her ureters. However, at this time, it is unknown if the da vinci si surgical system caused or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3246305 | 3217393 | 3113948 | 2955842201302493 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-03 | 2013-07-10 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-01-03 | Device Not Retu | No | Unkown | 2013
IS3000 A60P8 | 2013-06-11 | Other | Other | It was reported that during a da vinci si nephrectomy procedure, the surgeon was unable to focus via the camera head or the surgeon side console (ssc). The surgeon was unable to focus by pressing the focus button on the camera head or ssc. The lens of the camera head was cleaned and the endoscope was removed by the surgical staff but the issue persisted. At the time the reported issue occurred, the patient was under anesthesia and port incisions had been placed. Due to the reported issue, the surgeon made the decision to abort the surgical procedure. | On (b)(6) 2013, an intuitive surgical inc. (isi) field service engineer contacted the site and instructed the robotics coordinator to order a replacement camera head. That same day, the fse verified with the robotics coordinator that the reported issue was resolved with replacement of the camera head. On (b)(6) 2013, isi contacted the robotics coordinator at the site and obtained additional information regarding the reported event. The robotics coordinator indicated that the patient side cart (psc) was docked to the patient and the surgeon was attempting to remove adhesions when the reported camera head issue occurred. Since the surgeon was unable to focus the camera head, he made the decision to abort the surgical procedure and rescheduled the procedure for the following week. The robotics coordinator did not know how long the da vinci si surgical procedure was in progress before the case was aborted. She indicated that the rescheduled procedure was completed with the da vinci si surgical system the following week with no reported issues. She denied that the patient experienced any complications. She also stated that the camera head issue was resolved. The camera was returned to isi and evaluated. Engineering evaluation concluded that mechanical shock resulted with damage to the camera's focus stage assembly. The camera was found to be out of alignment and needed adjustment/repair. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon made the decision to abort the da vinci si surgical procedure after experiencing issues with focusing the camera head. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3246805 | 3217836 | 3114391 | 2955842201302497 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-06-11 | 2013-07-10 | Yes | Yes | Health Professional | IS3000 A60P8 | Device Returned To Manufacturer | No | 2013-06-11 | Yes | No | Unkown | 2013
IS2000 A51P8 | 2013-06-10 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, while docking to the patient, the site experienced system error code 23003. With the assistance of an intuitive surgical, inc. (isi) technical support engineer (tse) the site removed the instruments from the patient side manipulator (psm) arms and cycle powered the system. The site advised the tse that the issue was resolved and that they were going to continue on with the case. The tse reviewed the site's system logs and confirmed that system error code 23003 occurred and was generated when psm 2 was bumped. 3 hours later the site contacted isi technical support engineering and reported that after they restarted the system, a flashing camera icon and a question mark were displayed on the touchscreen monitors. Unable to resolve the issue the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. | The investigation performed by the intuitive surgical, inc. (isi) field service engineer concluded that the flashing camera icons experienced by the site occurred due to the site had not selected the endoscope angle. The fse instructed the site on how to select the endoscope angle to resolve the issue. On (b)(4) 2013, isi contacted the initial reporter concerning the reported event. The initial reporter indicated that the system was docked to the patient when the flashing camera and question mark appeared on the system monitors and that the surgical staff was unable to resolve the issue. Prior to contacting isi for technical support assistance the surgeon made the decision to complete the procedure using open surgical techniques due to not being able to resolve the camera issue and due to the patient's anatomy. The patient had a large uterus and fibroid. The initial reporter indicated the size of the patient's anatomy affected the port site placements, making it difficult for the surgeon to access the surgical site. There was no harm to the patient and the patient tolerated the open procedure well. The initial reporter indicated that she was told by the surgeon that the patient is doing well. As of (b)(4) 2013, there have been no reported recurrences of the reported issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3247033 | 3218051 | 3114606 | 2955842201302502 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-06-10 | 2013-07-10 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2013-06-10 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-06-05 | Injury | Other | It was reported that during a da vinci si nissen fundoplication procedure, the coating around the tip of the permanent cautery spatula instrument broke off and fell into the patient. The broken instrument piece was not retrieved. The planned surgical procedure was completed and no patient harm was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported complaint with the following clarification that a small piece approximately. 0785 of the ceramic sleeve was broken off. The broken piece was not returned to isi. Engineering concluded that the damage was likely due to mishandling. The instrument also passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the surgeon who performed the surgical procedure. The surgeon indicated that he does not recall the specific surgical task that he was performing when the instrument broke and that it was one of the surgical staff members that observed that the instrument was broken. The surgeon indicated that he spent 15 minutes searching for the broken instrument piece; however, he was unable to locate it. No additional intervention was required to search for instrument piece. The surgeon indicated that there was no harm to the patient and the patient has not returned to the hospital due to retaining a foreign object. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3247023 | 3218057 | 3114612 | 2955842201302504 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-06-10 | 2013-07-10 | Yes | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-06-10 | Yes | No | Reuse | 2013
IS1200 A43 P5 | 2013-06-28 | Other | Other | It was reported that during a da vinci standard (is 1200) surgical procedure, the surgical staff encountered a 22003 system error message when homing the system. The nurse contacted an intuitive surgical inc. (isi) technical support engineer (tse) and reported an error pointing to the red arm that was not correctly docked. Tse noted that there was no red arm on this particular system. Tse requested the site check the connections and emergency power off (epo) the system; however, the system error 22003 returned. Tse found that the error message was pointing to the setup joint (suj3) and axis 1 of the endoscope camera manipulator (ecm) arm. Patient was under anesthesia during the system error message and surgeon decided to abort the surgical procedure. | Field service engineer (fse) went to the site on (b)(4) 2013 and observed error message 22003 which was pointing to the set up joint (sujc) axis 2 potentiometer (pot). Fse discovered that the pot cable from the remote interface adapter (ria) board to pots was defective. There was a faulty contact on the cable. When the fse touched the cable at the shocked point, it triggered the error 22003. The red arm not correctly docked was also observed. A potentiometer cable connected to setup joint for camera arm board (sjc) was replaced to resolve the fault contact that triggered the error 22003. The system verification was performed and passed testing and was within the manufacturer's specifications. Next surgery using this system is scheduled for (b)(6) 2013. Isi sent follow up questions and the initial reporter was unwilling to answer some of the additional clinical information. The following information was provided by the nurse: the system error message occurred on (b)(6) 2013. The system was pre-inspected prior to the surgical procedure with no issues. Nurse mentioned right after the system error message, the surgeon decided to abort the case. She did not provide information of how long the patient was under anesthesia prior to the system error message, what type of surgical procedure the patient was having with the da vinci system or whether any ports placed on the patient. Per nurse, the patient did not suffer any injury during the da vinci surgical procedure. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3238098 | 3209228 | 3105783 | 2955842201302451 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2013-06-28 | 2013-07-08 | Yes | Yes | Health Professional | IS1200 A43 P5 | Device Returned To Manufacturer | Yes | 2013-06-28 | Yes | No | Reuse | 2013
42017910 | 2011-06-10 | Injury | Required InterventionHospitalization | It was reported that post op a da vinci si hysterectomy procedure performed on (b)(6) 2011, the patient was re-admitted into the hospital due to sepsis as a result of bowel injury that occurred during the surgical procedure. The patient was hospitalized for 40 days. | The monopolar curved scissor (mcs) instrument and tip cover accessory have not been returned for evaluation; therefore, the root cause of the injury sustained by the patient cannot be determined. A follow-up mdr will be submitted if the instrument or tip cover are returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013 intuitive surgical contacted the initial reporter regarding the reported event. Isi received a response on (b)(4) 2013. The initial reporter indicated that the patient post op a da vinci si hysterectomy procedure, the patient was re-admitted into the hospital on (b)(6) 2011 presenting symptoms of sepsis. During a surgical procedure of the patient's abdomen, it was discovered that the patient's large intestine had a lesion. The patient underwent hartmann's procedure for repair of the affected area. The patient required extracorporeal membrane oxygenation, nova lung, dialysis, parenteral nutrition, blood transfusion components and other treatments due to multiple organ failure. The patient was in the critical patient unit until (b)(6) 2012 and was discharged from the hospital on (b)(6) 2012. It is the hospital's belief that the sepsis developed by the patient originated from the lesion on the patient's large intestinal wall. The hospital's pathologist told the initial reporter that the lesion was thermal. The site believes that the burn sustained by the patient may have been caused by arcing from the monopolar curved scissors (mcs) instrument that was used during the surgical procedure. No malfunction of the mcs instrument occurred during the surgical procedure and arcing was not observed. No injury to the patient was noted during the surgical procedure. The mcs instrument, mcs tip cover and cannula were not inspected prior to use. The initial reporter indicated that a video recording of the surgical procedure is available; however, it has not been provided to isi for review. Isi performed a review of the site's system logs and found that there was no procedure performed on the reported procedure date. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3234312 | 3205737 | 3102292 | 2955842201302422 | 1 | NAY | Manufacturer | Distributor | Other | Initial | 2013-06-06 | 2013-07-03 | Yes | No | Health Professional | 42017910 | No | No | 2013-06-06 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2011-06-02 | Injury | Other | On (b)(4) 2013, intuitive surgical, inc. (isi) received a legal complaint indicating that a patient who underwent a da vinci si prostatectomy procedure sustained bladder and pelvic wall injuries. Reportedly the injuries sustained by the patient caused a small bowel obstruction with fistula. | On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the surgeon who performed the surgical procedure. The attorney representing the surgeon responded indicating that at this time the surgeon is unable to provide additional information regarding the reported event due to legal constraints. A review of the site's system log for the reported procedure date found no system errors were generated that would have caused or contributed to the injuries sustained by the patient. This complaint is being reported due to an allegation that a patient sustained injuries during da vinci surgical procedure. Due to the limited information provided, at this time, isi is unable to determine the root cause of the injuries sustained by the patient. A follow-up medwatch report will be submitted to the fda if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3234429 | 3205845 | 3102400 | 2955842201302424 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-06-03 | 2013-07-03 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-06-03 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2013-06-04 | Other | Other | It was reported that during a da vinci si partial nephrectomy procedure, the site experienced a non-recoverable system error code 252. The site contacted intuitive surgical for troubleshooting assistance. Review of site's system logs by a technical support engineer (tse) confirmed that system error code 252 occurred and system error code 307 occurred and were associated with the double camera control unit (doco). With the assistance of the tse, the site examined the boom system for power to the doco. The site found that the boom system was plugged in and turned on. The tse also had the site switch the power plugs, however, the issue persisted. The patient was under anesthesia and port site incision had been created when the surgeon made the decision to abort the planned procedure and rescheduled it for a later date. No patient harm or injury was reported. | The investigation conducted by the field service engineer concluded that the non-recoverable system error code 252 and system error code 307 experienced by the site were associated with the double camera control unit (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right side image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the initial reporter regarding the reported event. The initial reporter indicated that the da vinci surgical system was docked to the patient when the system error code 252 occurred. The patient's port incisions were closed after abortion of the surgical procedure. The patient was discharged from the hospital on (b)(6) 2013 and the patient has not returned to the hospital due to experiencing any complications. On (b)(6) 2013, the initial reporter reported that there was no injury to the patient as a result of the reported event. The system alarm (system generated fault codes) functioned as designed and there was no injury to the patient. System error code 252 appears when the master supervisory controller does not receive an expected message within a specified time. Upon determining this condition, the safety system puts da vinci si in a non recoverable safe state. The doco was returned to isi for evaluation. Engineering was unable to reproduce the system error codes experienced by the site. The doco installed on an isi in-house pca system showed that the doco passed a running video test, white balance testing and calibration. The doco was also power cycled (b)(4) times and a was set idle for (b)(4) and was found to function within specifications. On (b)(6) 2013, the isi clinical sales representative for the site reported that the surgical procedure was scheduled to be performed on (b)(6) 2013. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3234459 | 3205868 | 3102424 | 2955842201302426 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-06-04 | 2013-07-03 | Yes | Yes | Health Professional | IS3000 A60P7 | Device Returned To Manufacturer | No | 2013-06-04 | Yes | No | Reuse | 2013
IS2000 A51P7 | 2013-06-04 | Injury | Other | It was reported that during a da vinci s mitral valve repair procedure, the instrument installed on patient side manipulator (psm) arm 1 was not rotating. With the assistance of an intuitive surgical, inc. (isi) technical support engineer (tse) the site installed different instruments and reseated the sterile adapter, however, the issue persisted. The site indicated that the surgeon encountered the issue every time he changed psms. The site advised the tse that they were continuing with the case. Approximately 1 hour later, the site called isi back for technical support assistance. The site indicated to the tse that the system was not working correctly and that the psm was shaking. The site declined to troubleshoot the issue with the tse, as they were in the middle of the procedure. The site requested that the tse dispatch an isi field service engineer to evaluate the system. | During the field investigation performed by the intuitive surgical, inc. (isi) field service engineer (fse), the fse learned that the surgeon made the decision to convert the planned surgical procedure to open, due to the patient's heart having been punctured. The fse was also told by the site's nurse manager that a power fluctuation or loss occurred during the surgical procedure. However, the nurse manager indicated that it is unknown if the power issue contributed to the issue experienced by the site. The fse was unable to replicate shaking on patient side manipulator (psm) 1 during functional testing of the system using test instruments. Isi contacted the initial reporter and she indicated that the psm was not shaking, as initially reported, but while retracting tissue, the force of the was not as strong as normal and the patient's tissue was slipping. Isi engineering also the site's system logs and confirmed that the system lost ac power during the surgical procedure and the ac power was restored after 2 minutes. No other system error codes were generated that would have caused or contributed to the issue experienced by the site. During evaluation of the site's system, the fse found that the friction was high for axis 4 on psm 1 and the friction was high for axis 1 on psm 2. Psm 1 and psm 2 were replaced to repair the system. Psm 1 and psm 2 were returned for evaluation. Engineering evaluation found that psm 1 failed friction testing for axis 1 and axis 2. The harmonic drives for axis 1 and 2 were replaced to repair psm 1. Engineering evaluation of psm 2 found that axis 4 failed friction testing. The upper pulleys for link 5 and 6 were sliding off the bearing, adding unwanted friction. The pulley assemblies were replaced to repair psm 2. Engineering was unable to replicate the issue experienced by the site, as psm 2 passed software testing, and normal mode testing using isi in-house test sterile adapters and instruments. The psm's cannula mount was found to be current and within specification. As a precaution, an embedded serializer for the instrument interface (esii) board was replaced. On (b)(6) 2013, isi contacted the surgeon who performed the surgical procedure. The surgeon indicated that he was performing a left dynamic atrial retraction using an atrial retractor instrument when the puncture to the patient's heart occurred. The surgeon indicated that it was his belief that the damage to the patient's heart was due to a malfunction of the psm. The surgeon also believes that the malfunction of the psm occurred as a result of wear and tear to the system over time. The surgeon indicated that the surgical staff should have inspected the system prior to his use of the system to ensure that it was functioning properly. The surgeon indicated that the patient tolerated the open procedure well and was discharged from the hospital. The patient is recovering well and has not returned to the hospital due to experiencing any post-surgical complications as a result of the reported event. This complaint is being reported due to the patient sustained an injury during a da vinci s mitral valve repair procedure. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3234655 | 3206001 | 3102557 | 2955842201302431 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-06-04 | 2013-07-03 | Yes | Yes | Health Professional | IS2000 A51P7 | Device Returned To Manufacturer | Yes | 2013-06-04 | Yes | No | Reuse | 2013
420179 |  | Injury | OtherRequired Intervention | It was reported that a patient who underwent a da vinci s gynecological procedure approximately one year ago, sustained a small burn on the bowel. It is the surgeon's belief that the burn injury sustained by the patient may have been caused by arcing from the monopolar curved scissors (mcs) instrument used during the surgical procedure. The mcs instrument was discarded. No other details regarding the reported event were provided. | On (b)(6) 2013, intuitive surgical, inc. (isi) contacted the surgeon who performed the surgical procedure. The surgeon indicated that the patient underwent a da vinci s gynecological procedure approximately 1 year ago and that 3 days post op, the patient was diagnosed with peritonitis and it was discovered that the patient's abdominal cavity contained fecal matter. The patient underwent a surgical procedure to have her belly flushed. The surgeon indicated that arcing during the surgical procedure was not observed; however, it is her belief that the injury sustained by the patient could have occurred due to arcing from the monopolar curved scissors (mcs) instrument that was used during the surgical procedure. The surgeon indicated that the patient is now doing fine. This complaint is being reported due to post-surgical complications experienced by the patient as a result of a bowel injury that occurred during a da vinci surgical procedure. Based on the information provided, isi is unable to determine the root cause of the injury sustained by the patient. Per the information provided by the initial reporter, the monopolar curved scissors instrument was discarded. A follow-up medwatch report will be submitted to the fda if additional information is received regarding the reported event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3234639 | 3206027 | 3102583 | 2955842201302435 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2013-06-05 | 2013-07-03 | Yes | Yes | Health Professional | 420179 | No | No | 2013-07-10 | No | No | Reuse | 2013
| 2010-08-31 |  | HospitalizationLife ThreateningDisabilityRequired Intervention | On (b)(6) 2010, my wife had hysterectomy davinci robot surgery at (b)(6), dr (b)(6) (obgyn). She was discharged the next day. That night i made her dinner and gave her a glass of water. She ate then finished off the glass of water. Minutes after that. Extreme pain! we went to the emergency room by ambulance. A medical doctor at the hospital took over her care after the robot doc said there was nothing more she could do. She went back into surgery. The doctor said her bowel was "pierced/burned" outside of the operating field during the surgery with the robot. Ten (10) open surgeries later. She lives today (most don't). Made a complaint to the (b)(6) medical board - they said it was caused by the "robot" not the doctor. They say its not uncommon using the robot. I went to the local media to investigate and warn the public and asked them why my wife's incident is not listed on the maude database by the doctor or the hospital. (b)(4) from intuitive surgical inc. Called my phone 05/29/2013, and asked me what happened, who the doctor was, etc. Without looking at my wife's file he insisted that it was the doctor's fault and not a single injury has been reported and that he would add the incident to the maude database. We had no idea we had to report the "incident" to the fda. We assumed the hospital or doctor would be mandated to do so. At the time were under the impression that it was the doctor's fault. Now with the 05/08/2013 "urgent medical device notification" issued by intuitive surgical inc for the 2009 and 2010 monopolar parts and now i have seen the fda 483 applications investigation done, i feel mr (b)(4) is trying to do what he should have already done. Unplug the "robot!". |  | DAVINCI SURGICAL ROBOT | DAVINCI SURGICAL ROBOT | 3246467 | 3212779 | 3109334 | MW5030788 | 1 | NAY | Voluntary |  | Other | Initial | 2013-06-27 | 2013-07-03 | Yes | Yes | Health Professional | [7 empty] | 2013
IS3000 A60P7 | 2013-05-10 | Death |  | It was reported that post op a davinci si gynecological procedure, the patient expired due to sepsis as result of a suspected bowel injury. | Intuitive surgical, inc. (is) has contacted the hospital several times to obtain additional information regarding the reported event. On (b)(4) 2013, the initial reporter of this complaint reported that this event will be investigated under the hospital's peer review process. At this time the hospital does not believe that the patient's demise was related to the da vinci surgical system, instrument or accessories; however, this cannot be confirmed, as the hospital has not completed their investigation. The initial reporter was unable to provide any further details concerning the reported event. A follow-up medwatch report will be submitted to the fda if additional information is received. On (b)(4) 2013, the isi clinical sales representative for the hospital spoke to the initial reporter regarding the event. The initial reporter indicated to the csr that the patient underwent a davinci si hysterectomy procedure for gyn oncology. At 13 days post op, the patient developed sepsis. The hospital's internal assessment determined that the complications and subsequent demise of the patient was not related to the da vinci surgical system, instruments or accessories. No other information was provided. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3226237 | 3197620 | 3094175 | 2955842201302372 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-05-31 | 2013-07-01 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2013-07-29 | No | No | Reuse | 2013
IS3000 A60P6 | 2013-05-14 | Death | Death | It was reported that several days post a da vinci single site cholecystectomy procedure the patient expired at home. | On (b)(6) 2013, intuitive surgical, inc. (isi) contacted the surgeon who performed the surgical procedure. The surgeon indicated that the planned surgical procedure was performed on (b)(6) 2013. Per the surgeon, he was not aware that the patient had experienced any intra-operative complications and the procedure seemed to go fine. No malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure. The surgeon indicated that he does not know the date that the patient expired and that at this, he was unable to provide the root cause of the patient's death, since he has not yet received a copy of the coroner's report. On (b)(6) 2013, isi received additional information from the hospital's risk manager regarding the patient. The risk maf the site's system logs for the reported procedure date found no system errors were generated that would have caused or nager indicated that the patient expired on (b)(6) 2013 and that the patient was not re-admitted into the hospital between the date of discharge and the date of death. No other information was provided. A review ocontributed to the patient's demise. This complaint is being reported due the patient's demise post a da vinci si single site cholecystectomy procedure. Based on the information provided, isi is unable to determine the root cause of the reported event. A follow-up medwatch report will be submitted to the fda if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3225981 | 3197419 | 3053543 | 2955842201302370 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-30 | 2013-06-28 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-05-30 | No | No | Reuse | 2013
| 2013-04-24 | Malfunction |  | During procedure, pk dissecting forceps cable broke after only three uses. The device was removed from the field, and another device was utilized. It is unknown if the new device was of the same or different lot. ======================manufacturer response for dissecting forceps, pk dissecting forceps (per site reporter). ======================none at this time. What was the original intended procedure?laparoscopic bilateral salpingo-oophorectomy and extensive adhesiolysis. Device #1is this a laboratory device or laboratory test?no. |  | PK DISSECTING FORCEPS | SYSTEM SURGICAL COMPUTER CONTR | 3258292 | 3229243 | 3125798 | 3229243 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-06-24 | 2013-06-27 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
| 2013-06-18 | Malfunction |  | When small clip applier endowrist was being loaded with the small clip, one of the tips on the endowrist broke. Instrument was not inside of the patient. What was the original intended procedure?robotic thymectomy. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST SMALL CLIP APPLIER DA VINCI | SYSTEMSURGICALCOMPUTER CONTROL | 3283298 | 3254025 | 3150580 | 3254025 | 1 | NAY | User Facility |  |  | Initial | 2013-06-26 | 2013-06-26 | No | Yes | UNKNOWN |  | No | No Answer Provided |  |  |  |  | 2013
| 2013-05-28 | Malfunction |  | Robotic needle holder noted to fray apart at tip of instrument. No pieces noted to come off, instrument removed from patient. No pieces noted in operative site. X-ray taken at end of procedure to verify. What was the original intended procedure?robot-assisted laparoscopic prostatectomy, bilateral nerve sparing. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST | ELECTROSURGICAL | 3227005 | 3198354 | 3173108 | 3198354 | 1 | NAY | User Facility |  |  | Initial | 2013-06-25 | 2013-06-25 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
428080 | 2013-06-05 | Injury | Required Intervention | Davinci single site maryland bipolar instrument began smoking in an unusual manner, so it was replaced. The "2 no" maryland also began smoking unusually as the first, but it also melted the plastic at the joint/hinge. The melted portion prevented the instrument from being removed through the trocar/port. The entire port had to be removed. Both episodes occurred while the instrument was in use in the pt. |  | DAVINCI | SINGLE SITE BIPOLAR MARYLAND F | 3233983 | 3202280 | 3098835 | MW5030708 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-06-25 | 2013-06-25 | No | Yes | Health Professional | 428080 |  | Yes |  |  |  |  | 2013
| 2013-06-04 | Malfunction |  | Robotic cardiere forcep reported by surgeon as not properly closing. Instrument passed off field and new forceps opened. |  | DAVINCI CARDIERE | SYSTEM SURGICAL COMPUTER CONTR | 3258322 | 3229280 | 3125835 | 3229280 | 1 | NAY | User Facility |  |  | Initial | 2013-06-25 | 2013-06-25 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
IS2000 A51P8 | 2013-06-11 | Other | Other | It was reported that during a da vinci s surgical procedure the surgical staff encountered a 22003 system error message. The nurse contacted an intuitive surgical inc. (isi) technical support engineer (tse) and found the error message was pointing to the setup joint (suj3), and axis 3 of the endoscope camera manipulator (ecm) arm. The tse requested that the site cycle the ecm suj in several different positions and this resolved the error message. However,when the site attempted to re-dock the patient side cart to the patient, the system error 22002 returned. The site emergency power off (epo) the patient side cart (psc), and the system error 2203 appeared after epo. A port placement had already been placed on the patient, the psc had not been docked to the patient when the surgeon decided to convert the procedure to traditional open surgical techniques. | The investigation conducted by the field service engineer concluded that system error code 23008 was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart, which provides the sterile interface for the 3d endoscope. The system was repaired by replacing the potentiometer on the suj axis 3. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. The ecm was returned to isi for failure analysis investigation. Engineering determined that an axis motor encoder had malfunctioned, thus generating the system error code experienced by the customer. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. Isi has been unsuccessful in getting a hold of the site to answer any additional clinical information. It would have been helpful to know how long the patient was under anesthesia prior to the alleged error message, whether the system was checked prior to the procedure and whether at that time they noticed any error messages, what type of surgical procedure the patient was having using the da vinci s system and the patient's status post-surgery. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3212418 | 3182918 | 3079473 | 2955842201302259 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-06-11 | 2013-06-21 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2013-06-11 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2010-05-21 | Injury | Required InterventionLife Threatening | On (b)(6) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci s total hysterectomy with bilateral salpingo-oophorectomy (bso) on (b)(6) 2010. The legal complaint alleged that the patient experienced complications including but not limited to severe bowel injuries, excessive bleeding, injury to the sigmoid colon and intestines, plus post operative complications and infection including pain, bleeding, fever, chills, nausea and other symptoms indicative of a bowel issue, infection and complications as early as her post-operative visit on (b)(6) 2010, with a doctor resulting in further suffering and complications to the plaintiff including but not limited to constipation, upper abdominal pain, vomiting, an abscess, a massive abdominal infection and the need for further surgery including a bowel resection on (b)(6) 2013. | Intuitive surgical has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. At this time, there is no allegation that a malfunction of the da vinci s system, an instrument, or an accessory occurred during the surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the legal complaint alleged that a patient, who underwent a da vinci surgical procedure, sustained post-operative complications and required a secondary surgical procedure. However, at this time, it is unknown if the da vinci s surgical system caused or contributed to the patient's injury. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3212648 | 3183159 | 3079714 | 2955842201302264 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-05-22 | 2013-06-21 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2013-05-22 | No | No | Reuse | 2013
42000606 | 2013-05-03 | Injury | Required Intervention | It was reported that during a da vinci si surgical procedure, the shaft of the large needle driver instrument was accidentally morcellated while the surgical staff was morcellating a fibroid in the patient's pelvic cavity. Two fragments from the large needle driver instrument fell into the patient but were retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The large needle driver instrument was returned and evaluated. Based on a visual inspection of the instrument, the distal end of the main tube exhibited a 1. 04 long section with material removed on one side of the tube, leaving a rough surface finish. Engineering evaluation concluded that the damage was likely due to mishandling and consistent with the reported accidental morcellation. No other instrument damage was found. On (b)(4) 2013, intuitive surgical inc. (isi) obtained additional information regarding the reported event from the initial reporter. The initial reporter indicated that all fragments were retrieved and an x-ray was performed as an attempt to look for possible retained fragments. The x-ray results were negative. The initial reporter indicated that there was no adverse outcome to the patient due to the reported event. Based on the information provided, this complaint is being reported due to the following conclusion: two fragments from the shaft of the large needle driver instrument fell into the patient and were retrieved. In addition, an x-ray was performed on the patient in an attempt to look for possible retained fragments. However, there is no evidence that a malfunction of the instrument occurred during the surgical procedure. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3212674 | 3183187 | 3079742 | 2955842201302265 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-23 | 2013-06-21 | Yes | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-05-23 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2013-05-23 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the surgical staff encountered a 23020 system error message. It was indicated that the error reoccurred each time the recover button was pressed and the site power cycled and emergency powered off (epo) the patient side cart (psc). However, the error message continued to occur. Due to the recurring error message, the surgeon made the decision to convert to traditional open surgical techniques. No patient harm, adverse outcome, or injury was reported. | On (b)(4) 2013, an intuitive surgical inc. (isi) technical support engineer (tse) reviewed the site's system logs and found 23020 and 23034 switch errors related to the endoscopic camera manipulator (ecm). That same day, an isi field support engineer (fse) performed a field evaluation at the site. The fse resolved the error messages by replacing the ecm. The system was then tested and found to be within specifications. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23020 appears when the da vinci s safety system determines one of the switches in a specific manipulator is showing inconsistent signals on two switch leads. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. Upon determining this condition, the safety system put da vinci in a recoverable safe state. On (b)(4) 2013, isi contacted the da vinci coordinator at the site and obtained additional information regarding the reported event. The da vinci coordinator indicated that the error message began at the start of the surgical procedure when the psc was being docked to the patient. The patient was under anesthesia and the incision ports were already placed when the reported issue began. Since the error message could not be resolved with troubleshooting, the surgeon made the decision to convert to open surgical techniques. The da vinci coordinator indicated that the open surgical procedure was completed. She did not know the patient's current status. To her knowledge, no patient harm, adverse outcome, or injury was reported. The ecm was returned and evaluated. Engineering evaluation found an camera clutch switch fault. While running the ecm in normal mode, an error 23034 was triggered as soon as the camera clutch was pressed. The arm was repaired by replacing the circuit assembly containing the switch. A visual inspection of the ecm revealed that the fault was being triggered due to a configuration (cfg) cap assembly switch that was found to be not making contact properly. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon made the decision to convert to open surgical techniques after encountering a recurring 23020 system error with the da vinci s surgical system. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3212680 | 3183191 | 3079746 | 2955842201302266 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-23 | 2013-06-21 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2013-05-23 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2012-09-25 | Injury | Other | On (b)(4) 2013, intuitive surgical inc. (isi) received a legal demand indicating that a patient who underwent a da vinci s hysterectomy with bilateral salpingo-oophorectomy procedure performed on (b)(6) 2012, experienced post-surgical complications, requiring the patient to undergo subsequent treatment. The patient's operative reports and other related medical records were also provided to isi. The patient's medical records indicated the following: the patient's medical history as indicated in the operative report regarding the da vinci hysterectomy with bilateral salpingo-oophorectomy procedure indicated that the patient experienced continued and worsening pain with her menstrual cycle. The operative report did not indicate that any intra-operative complications were experienced by the patient. The patient's uterus was removed vaginally and excellent hemostasis was noted. The patient's abdominal cavity was irrigated, inspected and there was no bleeding from any sites. There was no indication in the patient's operative report that any malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure. On (b)(6) 2013 the patient underwent a laparotomy procedure at a different hospital. The patient's operative summary indicated that the patient's pre-operative diagnosis was peritonitis with intra-abdominal abscess. The patient subsequently developed worsening abdominal pain and signs of multiorgan dysfunction including renal failure after undergoing a robotic hysterectomy. The patient was tachycardic into the 130's - 140's upon arrival to the hospital and appeared to be very ill. The patient experienced intermittent episodes of hypotension that were all consistent with septic shock and that the patient's abdominal examination was very concerning for peritonitis. The patient was resuscitated and given broad spectrum antibiotics and was taken to the operating room with a presumed diagnosis of bowel perforation believed to have been due to an intraoperative injury that possibly occurred during the patient's initial surgery. During the laparotomy procedure, the surgeon encountered a pocket of frank pus near one of the patient's previous port site incisions. Upon opening the patient's fascia, further purulence within the patient's abdominal cavity was noted. The patient was morbidly obese, a thick abdominal wall and deep pelvis, which made visualization of the surgical area very difficult. Purulent material and fibrinous exudate was also found to be scattered throughout the patient's abdomen on various sections of the small bowel as well as the entire colon and up over the liver. It was noted that the fluid did not appear to be frank stool or succus. During the procedure, it was discovered that the patient's sigmoid colon had a very small serosal tear on the anterior surface of the proximal rectum. The tear did not appear to be a full-thickness injury and the surgeon repaired the affected area using 3-0 interrupted silk sutures. Examination of the patient's colon, splenic flexture and transverse colon found no evidence of damage. Examination of the anterior surface of the patient's stomach and gastrointestinal (gi) tract found no evidence of perforation and a proctoscopy performed found no evidence of blood within the patient's rectum and there were no abnormalities within the rectal mucosa visualized. The patient's colon was insufflated and pinched off in the descending colon and the insufflated portion was submerged under water; however, no evidence of leakage was found. The patient's operative summary indicated that the surgeon was fairly satisfied that there was no continued perforation within the patient's intra-abdominal gi tract. The surgeon believes that the purulent-appearing fluid was suggestive that there was not an ongoing leak from the patient's gi tract. The patient's abdomen was irrigated with warm bacitracin saline and then closed using 1 looped pds with several 2 vicryl internal retention sutures placed within the fascia. The patient tolerated the procedure well and there were no complications experienced by the patient. The patient's operative summary indicated that an inspection of the patient's vaginal cuff found that it appeared to be intact. A cystourethroscopy performed found that the patient's bladder was intact and efflux of indigo carmine from both ureteral orifices was observed. | On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the hospital where the patient's da vinci s hysterectomy with bso (bilateral salpingo-oophorectomy) procedure was performed. The hospital's risk manager indicated that there was no record of the patient being re-admitted into the hospital after being discharged. There was no report by the surgical staff that an injury to the patient occurred during the surgical procedure and no intra-operative complications were reported to have been experienced by the patient. The risk manager indicated that there was no report by the surgical staff that a malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure. Due to hipaa regulations, the risk manager declined to provide any additional information regarding the patient. This complaint is being reported due to post surgical complications experienced by the patient as a result of infection. Based on the information provided isi is unable to determine the root cause of the peritonitis and intraabdominal abscess that the patient developed. A follow-up mdr will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3212701 | 3183214 | 3079769 | 2955842201302267 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-05-23 | 2013-06-21 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2013-05-23 | No | No | Reuse | 2013
IS3000 A60P6 | 2013-05-24 | Death |  | It was reported that a few days post op a da vinci si endometriosis resection procedure, performed on (b)(6) 2013, the patient expired at home. During the surgical procedure, extensive resection was completed on the patient's left pelvis surrounding the ureter and colon. | On (b)(6) 2013, isi contacted the hospital's risk manager. The risk manager indicated that there was no report by the surgical staff that any malfunction of the da vinci surgical system, instrument or accessories occurred during the surgical procedure and there were no issues noted during the surgical procedure. The hospital's investigation is currently ongoing and the root cause of the patient's demise has not been determined, as the hospital has not yet received the final report from the coroner's office. This complaint is being reported due the patient's death, post-op a da vinci surgical procedure. Based on the information provided, isi is unable to determine the root cause of the reported event. A follow-up medwatch report will be submitted if additional information is provided. A review of the site's system log for the reported procedure date found no system errors were generated that would have caused or contributed to the reported event. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3212736 | 3183229 | 3079784 | 2955842201302272 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-24 | 2013-06-21 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-05-24 | No | No | Reuse | 2013
42009307 | 2013-05-23 | Injury | Required Intervention | It was reported that during a da vinci si surgical procedure, slivers or hair like fibers broke off of an unknown da vinci surgical instrument and fell into the patient. The fragments were retrieved; however, the site is uncertain if all of the fragments were retrieved. No patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. (isi) has made multiple attempts to obtain additional information regarding the reported event and the instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3212738 | 3183233 | 3079788 | 2955842201302273 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-05-24 | 2013-06-21 | Yes | Yes | Health Professional | 42009307 | No | No | 2013-07-12 | No | No | Reuse | 2013
IS3000 A60P4 | 2011-05-05 | Injury | Other | On (b)(6) 2013, intuitive surgical received a telephone message from a patient who underwent a da vinci si ovarian cyst removal procedure on (b)(6) 2011. The patient reported that during the surgical procedure she sustained injuries and that post-operatively, she has continued to experience pain. The patient indicated that she is only able to get out of bed to go to the hospital or doctor's appointments. The patient underwent a post-operative check up with the surgeon; however, she has not had any subsequent care by the surgeon since that time. During a follow-up with an associate physician, the patient was told that more damage would occur if she underwent another surgical procedure. Per the information provided by the patient, for the last two years since the surgery, she has not had a life. | On (b)(4) 2013 intuitive surgical, inc. (isi) contacted the patient to obtain additional information regarding this report. The patient indicated that she was discharged from the hospital on (b)(6) 2011, and that for approximately 5 to 8 days post op, she experienced pain and bleeding. The patient indicated that since the procedure, she has experienced a continuity of pain in her rectum, vagina and bladder and has undergone multiple physical examinations due to her pain. The patient has been prescribed approximately 15 different medications to manage her pain. She is currently taking morphine. The patient also indicated that post-operatively she experienced vaginal cuff and rectum evisceration; however, she has not undergone any additional surgical procedure to repair these conditions. The patient indicated at that time that she would mail copies of her medical records to isi. On (b)(4) 2013 intuitive surgical contacted the surgeon who performed the surgical procedure for additional information regarding the patient's report. On (b)(4) 2013, a response was received from the hospital's risk manager indicating that they were unable to share any patient history information without a signed release authorization from the patient. On (b)(4) 2013 isi contacted the patient regarding the status of her medical records. The patient indicated that she has made the decision not to provide isi with copies of her medical records. Isi performed a review of the site's system logs for the reported procedure date and found that there were no system errors generated that would have caused or contributed to the post-surgical complications that the patient has experienced. This complaint is being reported due the patient's report of injury during a da vinci si ovarian cystectomy procedure. Based on the information provided, isi is unable to determine the root cause of the injury sustained by the patient since no other information was provided. A follow-up medwatch report will be submitted to the fda if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3212717 | 3183244 | 3079799 | 2955842201302270 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-05-28 | 2013-06-21 | Yes | No | Health Professional | IS3000 A60P4 | No | No | 2013-05-28 | No | No | Reuse | 2013
|  | Injury | Disability | On (b)(6) 2012, dr (b)(6) performed surgery of the removal of my left ovary. The da vinci robotic assist the surgeon. Soon after surgery, i became very ill with surgical menopause. After i requested a u. S. /pelvic transvaginal ultrasound, it was discovered the right ovary is not identified. It was removed in error during surgery. |  | DA VINCI ROBOTIC | NONE | 3225917 | 3193425 | 3089980 | MW5030655 | 1 | NAY | Voluntary |  | Patient | Initial |  | 2013-06-21 | Yes | No | Service Personnel |  |  | No |  |  |  |  | 2013
420179 | 2013-06-19 | Malfunction |  | During a robotics surgery case the surgeon noticed a burn hole on the monopolar tip cover. The instrument was removed and there was no harm to the pt. |  | INTUITIVE SURGICAL MONOPOLAR CURVED SCISSORS | HOT SHEARS | 3224543 | 3193618 | 3090173 | MW5030687 | 1 | GEI | Voluntary |  | RISK MANAGER | Initial | 2013-06-21 | 2013-06-21 | No | Yes | Health Professional | 420179 | Yes | Yes |  |  |  |  | 2013
IS3000 A60P8 | 2013-05-01 | Injury | Required Intervention | It was reported that a patient, who underwent a da vinci si hysterectomy procedure, sustained an unspecified bowel injury. | In relation to the complaint, the site returned to intuitive surgical inc. (isi) a monopolar curved scissors (mcs) instrument that was used during the surgical procedure. The instrument was evaluated and found to be fully functional with no damage or trouble observed. On (b)(4) 2013, isi contacted the risk manager at the site and obtained additional information regarding the reported event. She indicated that the patient's injury occurred due to the difficulty of the case in relation to the patient's size, girth, and anatomy. She was unable to provide details as to how the patient's injury occurred. To her knowledge, there was no malfunction of the da vinci si system, an instrument, or an accessory during the surgical procedure. The patient did not experience any intra-operative complications and was discharged on an unspecified date/time. According to the risk manager, a couple of days later after the surgery was completed, the patient experienced unspecified symptoms and returned to the hospital. The risk manager indicated that the site evaluated the patient and discovered that the patient had sustained a bowel injury. The site believes the patient's bowel was somehow perforated during the da vinci si hysterectomy procedure. She was unable to provide a description of the perforation. According to the risk manager, the bowel injury did not appear to be a thermal injury. A secondary open surgical procedure was performed on an unspecified date/time to repair the patient's bowel injury. Since the secondary surgical procedure was performed, the patient has been in and out of the hospital. Intuitive surgical inc. (isi) has made several attempts to contact the surgeon to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. At this time, there is no allegation that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the surgical procedure. The initial reporter of this complaint indicated that the surgeon did not think the patient's injury was due to the monopolar curved scissors instrument. Based on the information provided, this complaint is being reported due to the following conclusion: the surgical staff claimed that a patient who underwent a da vinci si hysterectomy procedure sustained a bowel injury. However, it is unknown if the da vinci si surgical system caused or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3210289 | 3180781 | 3077336 | 2955842201302243 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-21 | 2013-06-20 | Yes | No | Health Professional | IS3000 A60P8 | No | Yes | 2013-05-21 | No | No | Reuse | 2013
| 2013-06-13 | Malfunction |  | While in the middle of robotic case, vessel sealer alarmed and monitor displayed, "vessel sealer mulfunction". New vessel sealer opened to continue case. Used one sequestered. ======================manufacturer response for vessel sealer for da vinci s 1, endowrist one (per site reporter). ======================this is the 6th reported event at our facility with this device. There have been others not reported. Previous devices have been returned to mfg with no response as yet. Our facility has decided that until further notice and to see if this is user or device error, the device rep will be in all procedures using the device. What was the original intended procedure?gyn procedure. |  | ENDOWRIST ONE | INSTRUMENT SURGICAL COMPUTER C | 3232527 | 3203851 | 3100406 | 3203851 | 1 | NAY | User Facility |  |  | Initial | 2013-06-20 | 2013-06-20 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
40007704 | 2013-05-20 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the surgical staff observed that the 8 mm cannula seal accessory was missing from an 8 mm cannula. An x-ray was ordered since the surgical staff was unsure if the missing accessory had fallen into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | On (b)(4) 2013, intuitive surgical inc. (isi) contacted the isi customer sales representative (csr) assigned to the site and obtained additional information regarding the reported event. The csr indicated that an x-ray was performed on the patient and the missing cannula seal accessory was not found. The csr also indicated that the surgical staff searched the operating room (or) floor and the cannula seal accessory was not found. According to the csr, the site performed an internal investigation regarding the reported event and the site concluded that the cannula seal accessory was most likely not installed on the cannula at the start of the surgical procedure and also during the surgical procedure. On (b)(4) 2013, isi also contacted the initial reporter of this complaint and the site's risk management department. The initial reporter and risk manager confirmed that the site performed an internal investigation and the conclusion was made that the cannula seal accessory did not fall into the patient. The risk manager believed that the surgical staff inspected the cannula prior to the start of the surgical procedure but did not notice that the cannula seal accessory was not connect to the cannula. The risk manager denied that the patient experienced any intra-operative or post-operative complications. On (b)(4) 2013, isi contacted a nurse at the site to get clarification regarding the reported event. The nurse indicated that the clear cannula seal trap of the green cannula seal accessory was observed to be missing during the surgical procedure. She explained that the green cannula seal accessory was installed on the cannula during the surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the patient was exposed to an x-ray performed by the site during an attempt to find the missing cannula seal trap of the cannula seal accessory. However, there is no evidence that the cannula seal accessory or cannula seal trap fell into the patient. | 8 MM CANNULA SEAL ACCESSORY | ENDOSCOPIC ACCESSORY | 3208049 | 3178498 | 3075053 | 2955842201302227 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-20 | 2013-06-19 | Yes | Yes | Health Professional | 40007704 | No | No | 2013-05-20 | No | No | Unkown | 2013
428059 | 2013-06-18 | Malfunction |  | During robotic surgical procedure the surgeon noted that grasper forceps handpiece would not retract from the curved cannula assembly. After several attempts the complete grasping forceps and cannula assembly was removed from the robotic arm and exchanged with a new grasper/forceps set. The surgery was completed without further incident. The surgeon noted that the cannula had cut through the grasper forceps insulation coating and prevented the forceps from being removed from cannula. Surgeon noted that all of the forcep material remained intact and did not enter patient surgical field. Both cannula and attached forceps sent to biomedical for evaluation and reporting to manufacturer. ======================manufacturer response for robotic grasping forceps, crocodile grasper (per site reporter). ======================manufacturer issued rma for return and evaluation. What was the original intended procedure?robotic surgical procedure. Device #1is this a laboratory device or laboratory test?no. Device #2is this a laboratory device or laboratory test?no. |  | SINGLESITE | SYSTEMSURGICALCOMPUTER CONTROL | 3224364 | 3195395 | 3091950 | 3195395 | 1 | NAY | User Facility |  |  | Initial | 2013-06-19 | 2013-06-19 | No | Yes | Invalid Data | 428059 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
| 2013-06-13 | Malfunction |  | Snapfit scalpel which attaches to the instrument came off of instrument when surgeon was robotically bifurcating the uterus. Abdomen was searched robotically and the scalpel was recovered and noted to be bent but intact. Per surgeon, may have been bent from stress and long use with a very large fibroid in uterus. What was the original intended procedure?patient with history of fiboids and endometriosis. Underwent robotic laparoscopic hysterectomy removed with morcellation, bilateral salpingectomy, and appendectomy. Patient had a 16-week size multilobular fibroid uterus with saprophytic connections of her uterus and fibroids to the bowel. Device #1is this a laboratory device or laboratory test?no. |  | SNAPFIT | SYSTEMSURGICALCOMPUTER CONTROL | 3219226 | 3190310 | 3086865 | 3190310 | 1 | NAY | User Facility |  |  | Initial | 2013-06-18 | 2013-06-18 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
IS 3000 | 2013-05-22 | Other | Other | On (b)(6) 2013 a surgical technical assistant contacted isi technical support for a non-recoverable error message 40006 that happened during system start using da vinci si surgical procedure. An error 40006 means that there was a communication error internally in the system. The surgical assistance mentioned that the patient was under anesthesia with no incisions placed. Isi technical support had the site emergency power off (epo), system restart and had the site exercise master tool manipulator right (mtmr). The surgeon decided to reschedule the surgery for the following day. Isi technical support looked into the system log and noticed several error 40006 and 25326 pointing toward the mtmr. Error 25326 means communication timed- out and is waiting for a response from a remote compute engine (rce). Isi regulatory complaint analyst sent follow up questions and obtained the following information: a male patient was going in for a robotic assisted radical prostatectomy procedure on (b)(6) 2013. The patient was under anesthesia and approximately 2 hours and 4 minutes into the surgery the da vinci si system displayed a non-recoverable error message 40006. It was confirmed by the surgical technical support that there was no ports placed on the patient. There were no complications on the patient due to the alleged issue. The surgeon decided to stop the procedure and start the procedure again the following morning due to the alleged error messages. The field service engineer (fse) replaced the remote arm controller boards (rac) and performed a system verification test and test was ok. An electrical safety test was also performed and the test was ok. Issue with the alleged error messages was resolved. The alleged error messages were caused from the remote arm controller boards (rac). The patient had the surgery the following day using the da vinci si system with no complications. The patient status was in good condition after the procedure. The complaint is being reported due to the fact that the surgeon aborted the da vinci si system due to the alleged error messages. | On (b)(4) 2013, product came back for investigation and failure analysis (fa) was unable to reproduce error 40006 by installing the board into pca system and running 10 minutes since cycle, x 30 power cycles and sitting idle for an hour. Based on engineering's recommendation, the unit needed to be replaced. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3202066 | 3172872 | 3069427 | 2955842201302184 | 1 | NAY | Manufacturer | OtherHealth Professional | Other | Initial | 2013-06-17 | 2013-06-17 | Yes | Yes | Health Professional | IS 3000 | Device Returned To Manufacturer | No | 2013-05-22 | Device Not Retu | No | Reuse | 2013
VER 03 | 2013-06-03 | Malfunction |  | During davinci, an instrument got stuck in the davinci arm and wouldn't come out. After a few minutes of manipulation the pk finally came out. Upon inspection of instrument a wire had snapped which displaced it and caused the broken end to stick out. No harm to patient. A new pk was opened. Broken instrument was sent back to company. Device #1is this a laboratory device or laboratory test?no. |  | PK DISSECTING FORCEPS | SYSTEMSURGICALCOMPUTER CONTROL | 3232648 | 3203952 | 3100507 | 3203952 | 1 | NAY | User Facility |  |  | Initial | 2013-06-17 | 2013-06-17 | No | Yes | Physician | VER 03 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
IS3000 A60P8 | 2013-05-15 | Injury | Required Intervention | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure, the surgical staff indicated that the patient side manipulator (psm1) arm 1 moved inadvertently and the patient sustained a bowel injury. The surgeon repaired the bowel injury and the surgical procedure was completed. | Manufacturer Narrative: --/-- On (b)(4) 2013, an intuitive surgical (isi) field service engineer (fse) performed a field evaluation of the da vinci si surgical system at the site. Based on the fse's evaluation, the system tested to manufacturer specifications and the fse observed a case that day and no errors were observed. On (b)(4) 2013, isi contacted the clinical sales representative (csr) who was present during the surgical procedure and obtained the following information regarding the event: the surgical procedure was not recorded. An unspecified scissor instrument moved forward which resulted with the patient's bowel sustaining a perforation injury. The csr was unable to explain how or why the instrument moved forward. He also did not know if the surgeon's head was within the high resolution stereo viewer (hrsv) or if he was grasping the master tool manipulators (mtm) when the event occurred. The csr indicated that the surgeon was able to repair the patient's bowel injury using the da vinci system and the prostatectomy procedure was completed with no further issues reported. The csr stated the surgeon had gotten multiple informational messages during the procedure stating relax your grasp on the master so that it can self-align and the surgeon ignored them. The csr instructed the surgeon not to ignore these system messages in future surgical procedures. The surgeon informed the csr that the patient side manipulator (psm) arm moved as he was working at the end of the cannula. The csr instructed the surgeon on not working at the end of the cannula to have full range of motion with instruments. On (b)(4) 2013, isi contacted the risk management department of the site and obtained the following information: the patient was discharged from the hospital on (b)(6) 2013, and the patient did not sustain any post-operative complications. The risk management department read the operative notes and the surgeon noted, the arms got away from me and he could not control the process. Intuitive surgical has attempted to contact the surgeon to obtain additional information concerning the event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Isi has reviewed the system logs for the site with procedure date of (b)(6) 2013. The system logs revealed that an error 100 occurred 56 minutes into the surgical procedure. The error 100 signified that the setup joint 1 (suj1) on the patient side manipulator 1 (psm1, arm 1) was moved or bumped. It is unclear if the error 100 is associated with the reported event of psm1 moving inadvertently and contributing to the patient's injury. Based on the information provided, this complaint is being reported due to the following conclusion: the patient sustained a bowel injury during a da vinci si prostatectomy procedure. However, at this time, it is unknown if a malfunction of the da vinci si system, an instrument, or an accessory caused or contributed to the patient's injury. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) obtained additional information regarding the reported event. The surgeon contacted isi and provided the following information: the surgeon indicated that his head was within the high resolution stereo viewer (hrsv) and he was grasping both master tool manipulators (mtm's) when the patient side manipulator (psm1; right arm) jerked away from him. He denied moving the psm or the mtm's at the time of the event. He also stated that the psm1 was not touching any other instrument when the event occurred. Due to the inadvertent movement, the surgeon indicated that a monopolar curved scissors (mcs) instrument installed on psm1 perforated the patient's bowel. The surgeon indicated that the bowel injury was 1 cm in size. He repaired the injury by using the da vinci si surgical system and by suturing. The surgeon indicated that after the event occurred, he continued to use psm1 and completed the surgical procedure with no further issues reported. The surgeon denied that the patient experienced any post-operative complications. | DA VINCI SI SURGICAL SYSTEM | Reuse | 3199683 | 3170123 | 3066678 | 2955842201302170 | 1 | NAY | Manufacturer | Health ProfessionalCompany Representativ | Other | InitialFollowup | 2013-05-15 | 2013-06-14 | Yes | No | Health Professional | IS3000 A60P8 | Yes | No | 2013-05-15 | Device Not Retu | No |  | 2013
42017910 | 2013-05-16 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, the surgical staff began to smell something burning. The surgeon then seen granular fragments fall from the monopolar curved scissors instrument's shaft (proximal to the tip cover) into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. On (b)(4) 2013, isi spoke with the surgeon involved with the reported event. The surgeon indicated that he retrieved as much of the fragments as he could; however, his method of retrieval was not provided. As the fragments were disintegrated, he was unable to guarantee that all of them were retrieved. He stated that the patient is fine and has not reported any post-operative complications. To his knowledge, the monopolar curved scissors instrument and tip cover accessory were inspected prior to use. He estimated that the reported issue occurred 8 minutes into the procedure during the colpotomy process but also indicated that the instrument was used for a total of less than 8 minutes. It is unknown what the electrosurgical unit setting was at the time of the reported issue. The surgeon declined to answer further questions. Based on the provided information, this complaint is being reported due to fragments of the monopolar curved scissor instrument having fallen into the patient and due to the surgeon being unable to confirm that all of the fragments were retrieved. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3199740 | 3170194 | 3066749 | 2955842201302171 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-16 | 2013-06-14 | Yes | Yes | Health Professional | 42017910 | No | No | 2013-05-16 | No | No | Unkown | 2013
IS3000 A60P8 | 2013-05-13 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the patient's ureter was transected. At the end of the hysterectomy procedure, the surgeon had performed a cystoscopy and consulted with urology residents to verify that the patient had not been harmed. No further clinical information was provided. | On (b)(4) 2013, the intuitive surgical inc. (isi) obtained additional information from the site regarding the reported event. A urologist informed the initial reporter of this complaint, an isi representative, that the patient's ureter was transected inadvertently and was subsequently repaired. The patient was discharged on an unspecified date/time. The initial reporter did not know the patient's current condition and was unable to provide any additional information regarding the reported event. On (b)(4) 2013, isi contacted the urologist at the site and obtained additional information regarding the reported event. The urologist indicated that the reported event was completely unrelated to the da vinci system and no malfunction occurred during surgical procedure. According to the urologist, the event is an internal issue and is related to training. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. Isi has reviewed the system logs for the site with a procedure date of (b)(4) 2013, and no system errors were found to have occurred during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the patient's ureter was inadvertently transected during a da vinci si hysterectomy procedure. However, at this time, there is no allegation that a malfunction of the da vinci si system, an instrument, or an accessory caused or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3200043 | 3170522 | 3053543 | 2955842201302173 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-05-13 | 2013-06-14 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-05-13 | No | No | Reuse | 2013
IS2000 A51P8 | 2013-05-15 | Injury | Other | Event Description: --/-- It was reported that during a da vinci s surgical procedure, the patient sustained a ureteral injury that was due to surgeon error and not an issue with an instrument. | Manufacturer Narrative: --/-- On (b)(6) 2013, intuitive surgical inc. (isi) contacted the risk management department of the site. The risk management department was unable to provide any information regarding the reported event. Intuitive surgical has made several attempts to contact the surgeon to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. It is unknown what surgeon error contributed to the patient's ureteral injury. Isi has reviewed the system logs from the site with a procedure date of (b)(6) 2013, and no system errors were found to have occurred during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon indicated that the patient sustained a ureteral injury that was due to surgeon error. However, at this time, there is no evidence or allegation that a malfunction of the da vinci s system, an instrument, or an accessory caused or contributed to the patient's injury. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) obtained additional information regarding the reported event from the surgeon who performed the surgical procedure. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) intuitive surgical inc. (isi) obtained additional information regarding the reported event from the surgeon who performed the surgical procedure. The surgeon indicated that the patient was female and in her late twenties. The patient had pelvic pain and dyspareunia. She underwent a robotic laparoscopy with excision of endometriosis, lysis of adhesions, and chromotubation. An excisional biopsy showed endometriosis. Intraoperatively, the right uterosacral ligament was found to be very prominent and band-like. The uterus was fixed. The right uterosacral ligament was partially transected with the monopolar curved scissors (mcs) instrument. The surgeon denied that there was a malfunction of the instrument or the da vinci s surgical system. According to the surgeon, at the time of surgery, there was no apparent ureteral injury. Three to four weeks post-operatively, the patient complained of abdominal pain and back pain to her pcp. A ct was ordered and showed large pelvic fluid collection and hydronephrosis. A urologist performed a retrograde pyelogram that showed the injury to the distal ureter. A percutaneous nephrostomy tube was placed. The surgeon believed that the injury was thermal in nature and due to thermal spread or by direct contact with the ureter while using the mcs instrument. The surgeon indicated, i don't believe it was due to instrument failure or malfunction. | DA VINCI S SURGICAL SYSTEM | Reuse | 3200449 | 3171060 | 3067615 | 2955842201302176 | 1 | NAY | Manufacturer | Health Professional | Physician | InitialFollowup | 2013-05-15 | 2013-06-14 | Yes | No | Health Professional | IS2000 A51P8 | No | Yes | 2013-05-15 | Device Not Retu | No |  | 2013
IS3000 A60P8 | 2013-05-15 | Other | Other | It was reported that approximately 20 to 30 minutes during a da vinci si myomectomy procedure, the surgeon experienced instrument engagement issues on patient side manipulator (psm) arms 1 and 3. Due to the reoccurrence of the instrument engagement issue during the surgical procedure, the surgeon made the decision to convert to traditional open surgical techniques at approximately one hour and forty-six minutes. No patient harm, adverse outcome, or injury was reported. | On (b)(4) 2013, an isi technical service engineer (tse) reviewed the da vinci si system logs and found error 's 31009, 31010, and 31011 which signify intermittent connection between sterile adapter, instrument, and the manipulator arm. On (b)(4) 2013, an isi field service engineer (fse) performed a field evaluation of the site's da vinci si surgical system and found that the system operated within normal parameters. The fse observed a surgical procedure that day with the da vinci si surgical system and no problems were found during his observation of the case. The fse reviewed troubleshooting and correction techniques for instrument engagement issues with the site. On (b)(4) 2013, isi obtained additional information regarding the reported event from the risk management department at the site. The risk manager indicated that the instrument/sterile adapter engagement issues began 3 hours and 36 minutes into the surgical procedure. The surgical procedure was completed within 4 hours and 51 minutes. The patient did not experience any intra-operative complications. However, the patient experienced anemia, shoulder pain, and left hand numbness post-operatively that resolved within 1 day post-op. The patient did not receive or require additional procedures or post-operative treatment. The patient was discharged from the hospital and no further issues have been reported. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon made the decision to convert to open surgical techniques after experiencing reoccurring instrument engagement issues during a da vinci si surgical procedure. No further recurrences of this issue at this site have been reported as of the date of this report. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3200537 | 3171104 | 3067659 | 2955842201302177 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-05-15 | 2013-06-14 | Yes | Yes | Health Professional | IS3000 A60P8 | Yes | No | 2013-05-15 | Yes | No | Reuse | 2013
IS3000 A60P8 | 2012-10-03 | Injury | Required Intervention | On (b)(6) 2013, intuitive surgical (isi) received a legal complaint concerning a patient who underwent a da vinci si sigmoid colectomy procedure on (b)(6) 2012. The legal complaint alleged: immediately after the procedure, the patient suffered severe complications, including, but not limited to, an anastomosis leak requiring a second operation. In addition the legal complaint states that the surgical equipment manufactured by isi malfunctioned, causing and resulting in severe injuries and that the patient was hurt and injured in his health, strength and activity, sustaining injury to his body and person this was reported to have caused the patient great mental, physical, and nervous pain and suffering. | Intuitive surgical inc. (isi) has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. At this time, it is unknown if a malfunction of the da vinci si system, and instrument, or an accessory occurred during the reported surgical procedure. Isi has reviewed the system logs for the site with a procedure date of (b)(6) 2012, and no system errors were found to have occurred during the surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the legal complaint alleged that the patient experienced severe complications and required a second operation after undergoing a da vinci si surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3200545 | 3171112 | 3067667 | 2955842201302180 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-05-17 | 2013-06-14 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-05-17 | No | No | Reuse | 2013
| 2013-06-06 | Malfunction |  | Harmonic ace curved shears insert broke intraoperatively during robotic myomectomy procedure. Instrument was being used on fibroid tissue when it stopped working. Instrument removed from laparoscopic site and examined. Found the jaw was hanging, broken from the instrument. Harmonic ace curved shears insert and broken jawpiece sequestered. |  | HARMONIC ACE CURVED SHEARS INSERT | HARMONIC ACE DAVINCI INSERT | 3209213 | 3177515 | 3074070 | 3177515 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-06-10 | 2013-06-13 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2013
420227 | 2013-06-10 | Malfunction |  | During instrument pre-inspection, cable noted to be frayed. Staff verified. Not used on patient; no patient contact. |  | INTUITIVE SURGICAL | PK DISSECTING FORCEPS | 3209219 | 3177531 | 3074086 | 3177531 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-06-11 | 2013-06-13 | No | Yes | Health Professional | 420227 | Yes | Yes |  |  |  |  | 2013
IS3000 A60P8 | 2013-05-06 | Injury | Required Intervention | It was reported that during a da vinci si sacrocolpopexy procedure, the surgeon indicated that the right patient side manipulator (psm) on the patient side cart (psc) inadvertently moved and the patient's bowel/colon was pierced. At the time that the injury occurred, the psc was docked to the patient and instruments were inserted. The surgeon indicated that when she reached for the master tool manipulators (mtm), the right psm moved without her input. The patient's bowel was repaired via suturing and the surgical procedure was completed with the da vinci si surgical system. | On (b)(4) 2013, an intuitive surgical inc. (isi) field service engineer (fse) performed a field evaluation of the da vinci si surgical system at the site. The fse provided the following information: a gravity compensation test was performed on both master tool manipulators (mtm) of surgeon side console (ssc) 1. Both mtm's passed. A gravity calibration was also performed on both mtm's of ssc 2 which was used by a resident. The right mtm was found to be within specification. According to the fse, the left mtm needed recalibration. The fse performed a gravity calibration of the left mtm. The system passed a brake holding test on all psm's. A sine cycle was performed and all psm's were test driven on both consoles. In addition, gravity compensation was tested on both consoles. The fse indicated that no unusual drifting of the instruments was observed before and after the gravity calibration tests were performed. Isi has reviewed the system logs for the site regarding the reported surgical procedure. According to the system log, approximately 2 hours into the surgical procedure, an error 100 was recorded. The error 100 signified that at that moment, the setup joint (suj1, joint 4) of the right psm was bumped. It is unknown at this time if the error 100 occurred at the time the right psm inadvertently moved according to the surgeon. On (b)(4) 2013, isi contacted the isi clinical sales representative (csr) assigned to the site to obtain additional information regarding the reported event. The csr described the injury to the patient's bowel/colon as being more of a serosal tear. He denied that the patient's bowel/colon was perforated. To his knowledge, the patient did not experience any post-operative complications. The csr also indicated that the da vinci si surgical system was used in subsequent surgical procedure without any reported issues since the event occurred. The csr did not know if the surgeon's head was within the high resolution stereo viewer (hrsv) or if she was grasping the mtm's when the injury occurred. On (b)(4) 2013, isi obtained additional information from the surgeon who performed the surgical procedure involved with this complaint. The surgeon provided the following information: her head was in the high resolution stereo viewer (hrsv) and her fingers were through the velcro grips of the mtm's when the right psm of the patient side cart (psc) moved inadvertently. However, the surgeon denied that she moved any of the tools when the right psm inadvertently moved. The surgeon indicated that the patient's bowel/colon was perforated by a pk dissecting forceps instrument when the right psm moved. She was unable to provide the serial and lot of the instrument. After the injury occurred, the surgeon used the da vinci si surgical system to repair the perforated bowel/colon. The surgeon indicated that she continued to use the right psm installed on the psc and the pk dissecting forceps instrument during the remainder of the surgical procedure and no other issues were observed. The surgeon indicated that the patient did not experience any post-operative complications. According to the surgeon, the patient was currently stable. The surgeon indicated that the da vinci si surgical system involved with this complaint has been used in subsequent surgical procedures since the reported event occurred. In addition, to her knowledge, the psm installed on the psc involved with this complaint has been working properly. Based on the information provided, this complaint is being reported due to the following conclusion: the surgeon claimed that the right arm of the psc moved without her input and the patient's bowel/colon was pierced. However, at this time it is unknown if a malfunction of the da vinci si surgical system caused or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3195483 | 3166303 | 3062858 | 2955842201302131 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2013-05-13 | 2013-06-12 | Yes | No | Health Professional | IS3000 A60P8 | Yes | Yes | 2013-05-13 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-05-14 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the customer noted that the tenaculum forceps instrument had a broken pulley. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the broken cable if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3186110 | 3156907 | 3053442 | 2955842201302043 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-14 | 2013-06-08 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-05-14 | Device Not Retu | No | Reuse | 2013
40018012 | 2013-05-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, arcing from the monopolar curved scissors (mcs) instrument with the mcs tip cover installed was observed. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover and mcs instrument have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the initial reporter regarding the reported arcing event. The reporter confirmed that the installation tool was used to install the tip cover and that the instrument was inspected prior to use. The reporter indicated that the electrosurgical unit (esu) setting was set to 30, which meets the recommended power setting for the da vinci si system. The surgeon reportedly observed a lot of sparkles at the scissors while performing colpotomy on the patient. When the sparkles occurred, the instrument was not near any tissue or another instrument and it reportedly did not result in any patient injury. The procedure was completed as planned. Since the reporter was not present in the operating room when the alleged arcing event occurred, she was unable to report confirm if the energy was contained at the scissors tips or if it arced to another area. The reporter claimed she inspected the instrument after the procedure and did not see any visible physical damage to the mcs instrument or the tip cover. Isi attempted to contact the surgeon involved with this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. There was no indication that the device caused or contributed to an adverse event. However, this complaint is being reported due to the alleged arcing event. | 8MM TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3186164 | 3156942 | 3053477 | 2955842201302046 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-09 | 2013-06-08 | No | Yes | Health Professional | 40018012 | No | No | 2013-05-09 | No | Yes | Unkown | 2013
420179 | 2013-05-09 | Injury | Life Threatening | On (b)(6) 2013, it was reported that a patient sustained a possible ureteral burn during a da vinci si surgical procedure performed weeks earlier with the use of a monopolar curved scissors (mcs) instrument. The details regarding the surgical procedure such as the type of procedure, date/time of the procedure, patient information, and procedure outcome were not provided. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Intuitive surgical inc. (isi) contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, the initial reporter was unable to provide any information regarding the reported event. Isi has also attempted to contact the risk management department at the site. However, as of the date of this report, no additional information has been obtained. Based on the information provided, this complaint is being reported due to the following conclusion: the initial reporter claimed that a patient sustained a ureteral burn during a da vinci surgical procedure. However, at this time it is unknown if the da vinci si system, an instrument, and/or an accessory actually caused or contributed to the patient's reported injury. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3186185 | 3156951 | 3031892 | 2955842201302047 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-09 | 2013-06-08 | Yes | No | Health Professional | 420179 | No | Yes | 2013-05-09 | No | No | Unkown | 2013
42017910 | 2013-05-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff alleged that sparks were coming off the tip of the monopolar curved scissors (mcs) instrument. The surgical staff reportedly reused the mcs tip cover accessory on a different instrument and no further issues were observed. The planned surgical procedure was completed. There were no missing or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation was unable to confirm the alleged issue that sparks came off of the mcs instrument. The mcs instrument was placed on an in-house system for cautery testing. The mcs instrument performed with no arcing issues. Engineering evaluation did not find any signs of arcing, or burnt or char marks on the mcs instrument. Engineering evaluation also found a hairline crack at the distal end of the instrument's main tube. The hairline crack was found to be running in an axial direction and measured approximately 0. 03. On (b)(4) 2013, intuitive surgical inc. (isi) contacted the initial reporter of this complaint and obtained additional information regarding the reported event. The initial reporter indicated that the mcs instrument was used for about 5 minutes during the surgical procedure before a spark was first noticed on the distal end of the instrument's shaft. The surgical staff reportedly notified the surgeon. However, the surgeon continued with the surgical procedure. A minute later, the initial reporter indicated that a flash appeared around the same area where the spark was observed. At that point, the surgical staff claimed that electric energy arced to the patient's tissue and a small piece of floating black ash was observed. However, the initial reporter denied that there was any visual evidence of tissue damage. After the reported flash was observed, the surgical staff reportedly removed the mcs instrument. The surgical staff claimed that the tips of the mcs instrument were darkened. The mcs tip cover accessory was also removed and inspected by the surgical staff and no signs of damage were found. According to the surgical staff, the mcs tip cover accessory appeared to be correctly installed on the mcs instrument. The surgical staff reportedly re-installed the tip cover on a second, replacement mcs instrument and the surgical procedure was completed with no further issues reported. The initial reporter indicated that the surgical procedure was not recorded. The site was using a valley lab esu with 35 esu settings. The cannula was reportedly inspected with a pin gage and no issues were found. The surgical staff reportedly applied electro lube to the mcs instrument prior to installing the mcs tip cover accessory. The initial reporter indicated that the mcs instrument involved with the alleged arcing event was inspected prior to use and no issues were observed. She also denied that the mcs instrument was removed at any time before the alleged arcing event occurred. She indicated that there were no instrument collisions during the reported surgical procedure. Based on the information provided, this complaint is being reported due to the following conclusion: the surgical staff claimed that an arcing event occurred with the mcs instrument. However, an evaluation of the mcs instrument did not find evidence of arcing. In addition, there is no evidence that a serious injury occurred as a result of the alleged arcing event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3186213 | 3156957 | 3053492 | 2955842201302044 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-09 | 2013-06-08 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-05-09 | Device Not Retu | No | Reuse | 2013
40018012 | 2013-05-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff alleged that the monopolar curved scissor (mcs) tip cover accessory installed on the mcs instrument fell off and into the patient. The mcs tip cover accessory was retrieved and the surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | On (b)(4) 2013, intuitive surgical inc. (isi) contacted the initial reporter of this complaint, and the robotics coordinator at the site who was present during the reported surgical procedure to obtain additional information regarding the reported event. The robotics coordinator indicated that the mcs instrument and mcs tip cover accessory were inspected prior to the surgical procedure and no issues were observed. According to the robotics coordinator, the mcs tip cover accessory was installed with the installation tool and electrolube was used. The initial reporter and robotics coordinator did not know if the surgical procedure was recorded. The robotics coordinator indicated that the surgeon was able to use the mcs instrument during the surgical procedure without any issues before the mcs tip cover accessory was found to be missing. He also indicated that nobody actually witnessed the mcs tip cover accessory slip off of the mcs instrument. The robotics coordinator mentioned that the mcs instrument was not removed at any time prior to the incident and there were no instrument collisions during the surgical procedure. The robotics coordinator indicated that when the surgeon switched to the mcs instrument while the procedure was about 80% completed, the surgical staff allegedly noticed that the tip cover accessory was missing and the orange section of the mcs instrument was exposed. He indicated that the mcs tip cover accessory was retrieved laparoscopically via a laparoscopic assist port. He denied that any x-rays or additional surgical procedures were required to retrieve the mcs tip cover accessory. After the mcs tip cover accessory was retrieved, the surgical staff inspected the mcs tip cover accessory and the mcs instrument. The surgical staff reportedly did not find any issues with either device. The surgical staff installed a new mcs tip cover accessory on the same mcs instrument and the surgical procedure was completed as planned. According to the robotics coordinator and initial reporter, the patient did not experience any intra-operative or post-operative complications because of the reported issue. The robotics coordinator indicated that the mcs instrument involved in this complaint has been used in subsequent surgical procedures with no reported issues. The tip cover has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned (post engineering evaluation) or if additional information is received. Based on the information provided, this complaint is being reported due to the following conclusion: the surgical staff claimed that the mcs tip cover accessory installed on the mcs instrument unexpectedly came off during the surgical procedure. However, there is no evidence that the alleged issue caused or contributed to a serious injury. | MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3186199 | 3156961 | 3053496 | 2955842201302051 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-05-10 | 2013-06-08 | No | Yes | Health Professional | 40018012 | No | No | 2013-05-10 | No | Yes | Initial | 2013
42018306 | 2013-05-13 | Malfunction |  | On 05/15/2013,a clinical sales representative (csr) from intuitive surgical, inc. (isi) reported that the tip of the permanent cautery hook instrument allegedly started to melt during a da vinci si myomectomy procedure. There were no missing or fallen pieces reported. There was no allegation of patient injury as a result of the reported issue. | On (b)(4) 2013, intuitive surgical inc. (isi) contacted the csr to obtain additional information about the reported event. The csr confirmed with the user facility that the instrument was inspected prior to use. The instrument was used with a valley lab force fx electrosurgical unit with a setting of 35. The instrument was in use for about 15-20 minutes before the user facility observed a spark rising from the instrument's wrist. The patient did not suffer any injuries and the procedure was completed as planned in 2 hours. The patient's post operative condition was reportedly excellent and she did not experience any post operative complications or receive any post operative treatment. There was no indication that the product caused or contributed to an adverse event. However, this complaint is being reported because the instrument allegedly sparked at the instrument's wrist. The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3186201 | 3156963 | 3053556 | 2955842201302050 | 1 | NAY | Manufacturer | Company RepresentativeOther | Other | InitialFollowup | 2013-08-05 | 2013-06-08 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-08-05 | Yes | No | Unkown | 2013
IS2000 A51P8 | 2013-05-10 | Other | Other | It was reported that during a da vinci s transoral robotic surgery (tors) procedure, a non-intuitive movement issue occurred. According to the initial reporter of this complaint, the surgical procedure was reportedly converted to open surgical techniques. No patient injury was reported. No further clinical information was provided. | On (b)(4) 2013, the intuitive surgical inc. (isi) clinical sales representative (csr), who initially reported this complaint, indicated that she inspected the instruments that were used during the reported procedure. According to the csr, the instruments appeared as though they had not been cleaned properly. The csr suspected that the reported non-intuitive movement issue might have been related to a possible issue with the site's instrument sterilization process. In turn, the csr reportedly notified an isi sterile reprocessing specialist (srr) of her findings. Isi contacted the srr on (b)(4) 2013, to obtain additional information regarding the reported event. The srr was unable to provide any additional information since he was waiting for permission from the site to come in and perform an investigation of their instrument sterilization processes. On (b)(4) 2013, an isi field service engineer (fse) performed a field evaluation of the site's da vinci s system. The fse checked the cable tension on all 3 patient side manipulator (psm) arms. The cable tensions were reportedly found to be within specification. The system was test driven by the fse and a sine cycle was performed. According to the fse, the system performed normally and was found to be within specification. Proper operation of the system was verified by the fse. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. Based on the information provided, this complaint is being reported due to the following conclusion: the surgical staff indicated that a da vinci s tors procedure was converted to open surgical techniques after a non-intuitive movement issue occurred. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3186193 | 3156969 | 3053504 | 2955842201302052 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-05-10 | 2013-06-08 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-08-23 | Injury | Required Intervention | Event Description: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint with the following allegation related to a da vinci si hysterectomy procedure performed on (b)(6) 2012: the legal complaint alleged that the patient had both of her ureters burned and cut (one totally destroyed). She is still suffering. Additionally, she became septic and had a reconstruction surgery. | Manufacturer Narrative: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) contacted the isi clinical sales representative (csr). The csr reportedly spoke to the surgeon, who performed the hysterectomy procedure, the surgeon indicated that no incidents occurred during the hysterectomy procedure; the procedure was completed using the da vinci si robot system. Per the information provided by the surgeon, at the end of case a blue dye was used to inspect for any possible ureter leaks, no leaks were observed. He did not observe any malfunction of the da vinci si surgical system, instrument, and/or accessories during the surgical procedure. The surgeon indicated that on (b)(6) 2012, during the patient's follow up visit, the patient had a positive lymph node. He also indicated that he did not perform the patient's ureter repair, the patient went to the intense care unit to have the ureter reconstruction; surgeon doesn't have any other details. Intuitive surgical inc. (isi) contacted the risk management department and the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. Isi has reviewed the system logs with a date of (b)(6) 2012, the date the reported da vinci si hysterectomy procedure was performed. Based on the review of the system logs, there is no evidence that a system error occurred during the surgical procedure. --/-- Additional Manufacturer Narrative: --/-- Product problem was inadvertently selected on the initial mdr. | DA VINCI SI SURGICAL SYSTEM | Reuse | 3186216 | 3156976 | 3053511 | 2955842201302045 | 1 | NAY | Manufacturer | Other | Other | InitialFollowup | 2013-05-10 | 2013-06-08 | Yes | Yes | Health Professional | IS3000 A60P8 | No | No | 2013-05-10 | Device Not Retu | No |  | 2013
40015801 | 2013-05-13 | Injury | Life Threatening | Event Description: --/-- It was reported that during a da vinci si surgical procedure, the surgical staff alleged that the snap-fit paddle blade accessory installed on a snap-fit scalpel instrument broke off and fell into the patient. The broken fragment was reportedly not retrieved. No patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) obtained the following information from the site: the surgical procedure involved with this case was a total hysterectomy. The snap-fit scalpel instrument and snap-fit paddle blade instrument accessory were inspected prior to the surgical procedure and no issues were observed. The cleaning/sterilization process at the site had not recently changed. According to the surgeon, the snap-fit scalpel instrument was used in the procedure for about 10-15 minutes before the snap-fit paddle blade fell off. The surgeon indicated that at the time the blade fell off, the instrument was being introduced through the right side of the abdomen and the uterus was being held by two tenacula, both on the left side of the abdomen, port 2 and port 3. Also at that time, the surgeon indicated that the uterus had already been removed from it's attachment to the vagina and it was being morcellated because it weighed 930 gm and was too large to fit through the vagina. The surgeon denied that the snap-fit scalpel instrument collided with any other instruments or was removed at any time during the surgical procedure. The surgeon indicated that since the snap-fit paddle blade was getting dull, he pointed vertically to cut with the tip and that it probably went deep in the hard fibroid uterus and was suddenly pulled out rather than being broken off and was then deeply embedded in the tissue. The surgeon further indicated that when the uterus was pulled through the vagina, he believed that the snapfit then fell out onto the floor on got caught into the drapes. The surgeon also stated that x-rays of the patient's abdomen and the removed uterus did not reveal the snap-fit paddle blade. The surgeon denied that the surgical staff searched the operating room floor or between the patient's legs for the instrument accessory that fell off. The surgeon denied that the patient experienced any intra-operative or post-operative complications. He indicated that the patient was doing great at 4 weeks post-op. He also stated that the snap-fit scalpel instrument was inspected after the surgical procedure was completed and no damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned (post engineering evaluation) or if additional information is received. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. It is unknown if any attempts were made by the site to retrieve the broken fragment, if the broken fragment was retrieved since the reported event occurred, if the patient experienced any post-operative complications, and what the patient's current status is. Based on the information provided, this complaint is being reported due to the following conclusion: the surgical staff claimed that the instrument accessory broke off and fell into the patient. In addition, the instrument accessory was reportedly not retrieved. | SNAPFIT PADDLE BLADE ACCESSORY | Unkown | 3186204 | 3156987 | 3053522 | 2955842201302053 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-05-13 | 2013-06-08 | Yes | Yes | Health Professional | 40015801 | No | No | 2013-05-13 | Device Not Retu | Yes |  | 2013
IS3000 A60P8 | 2013-05-10 | Injury | Required Intervention | It was reported that during a da vinci si sacrocolpopexy procedure, the patient side manipulator (psm) arm 1 allegedly jumped and one of the patient's arteries was nicked by the instrument installed on psm 1. | On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the initial reporter of this incident to obtain additional information regarding the reported event. The initial reporter indicated that during the surgical procedure two monopolar curved scissors (mcs) instruments were used on patient side manipulator (psm) 1. During the surgical procedure, the first mcs instrument installed on psm 1 allegedly became stuck and was difficult to remove. The mcs instrument and mcs tip cover accessory were reportedly inspected prior to the procedure by the surgical staff and no damage was observed. The site reportedly decided to use a replacement mcs instrument due to the reported difficulty they had with removing the first mcs instrument from the cannula. The first mcs instrument was inspected upon removal from the cannula and no damage was found by the surgical staff. While using the second mcs instrument on psm 1, the surgeon and surgical staff alleged that psm 1 jerked forward, causing the mcs instrument to puncture one of the patient's arteries. At the time psm 1 allegedly jerked forward, the surgeon was reportedly performing a dissection of the patient's middle sacral ligament and was not cutting or cauterizing tissue. According to the initial reporter, the surgeon was reportedly using the mcs instrument to push the patient's anatomy back. The patient's artery was reportedly repaired using clips installed on an endowrist clip applier instrument. The second mcs instrument was inspected by the surgical staff after the incident occurred and since no damage was found, the surgeon made the decision to complete the surgical procedure with the same instrument installed on psm 3. The initial reporter indicated that the surgeon did not move his head from the high resolution stereo viewer (hrsv) or remove his hands from the master tool manipulators (mtm) when the reported injury occurred. The initial reporter denied that any system alarms or error messages occurred when psm 1 allegedly jerked. The surgical procedure was reportedly not recorded. According to the initial reporter, the patient did not require any other medical treatments or surgical intervention because of the alleged issue. The patient reportedly tolerated the surgical procedure well and was recovering. The patient was in stable condition. On (b)(4) 2013, a field service engineer (fse) performed a system verification at the site and no issues were found with psm 1. The fse also noted that three subsequent surgical procedures had been completed with the system since the reported event occurred. On (b)(4) 2013, isi contacted the fse to obtain additional information regarding the reported event. The fse indicated that during his field evaluation, he was denied access to the affected mcs instruments and sterile adapter by the site since the products were in quarantine. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3186240 | 3157008 | 3053543 | 2955842201302049 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-10 | 2013-06-08 | Yes | Yes | Health Professional | IS3000 A60P8 | No | Yes | 2013-05-10 | No | No | Reuse | 2013
40018012 | 2013-05-13 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, while the surgeon was using the monopolar curved scissors (mcs) instrument with the tip cover accessory installed, the surgeon observed an arc of energy near the end of the case and noticed a small burn near the rectum that could have been caused by the energy that arced through the tip cover accessory. She observed a small hole in the tip cover accessory. There was no clinical sales representative (csr) present during the case. The tip cover accessory was discarded. | The mcs tip cover was discarded; therefore, the root cause of the customer reported failure mode cannot be determined. On (b)(4) 2013, the isi clinical sales representative (csr) was contacted. The csr indicated he was not present during the surgical procedure when the injury to the patient occurred, however, the surgeon indicated that the tip cover accessory appeared to have a hole on it. Csr indicated that the planned surgical procedure was completed; the patient was stable and she didn't suffer any additional complications as a result of this incident. The patient had been discharged from the hospital. Isi has made several attempts to obtain additional information regarding the reported event; however, no response has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. | 8MM TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3186248 | 3157013 | 3053548 | 2955842201302054 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2013-05-15 | 2013-06-08 | Yes | Yes | Health Professional | 40018012 | No | Yes | 2013-05-15 | No | Yes | Unkown | 2013
42018306 | 2013-05-13 | Malfunction | Life Threatening | Event Description: --/-- On (b)(6) 2013, a clinical sales representative (csr) from intuitive surgical, inc. (isi) reported that the tip of the permanent cautery hook instrument reportedly melted and burned a part of the patient's uterus during a da vinci si hysterectomy procedure. There were no missing or fallen pieces reported. | Manufacturer Narrative: --/-- On (b)(6) 2013, intuitive surgical inc. (isi) contacted the csr to obtain additional information about the reported event. The csr confirmed with the user facility that the instrument was inspected prior to use. The instrument was used with a valley lab force fx electrosurgical unit (esu) with a setting of 35. The instrument was in use for about 15-20 minutes before the instrument's tip started to melt. Prior to the tip melting, the user facility observed the casing around the instrument's wrist melting. The instrument reportedly burned a superficial area of 1. 5 cm on the patient's uterus. The patient did not receive any medical intervention but the patient and family were counseled and no additional steps were taken. The procedure was completed as planned in 2 hours. The patient's post operative condition was reportedly excellent and she did not experience any post operative complications or receive any post operative treatment. This complaint is being reported because the patient reportedly suffered a uterus burn after the tip of the instrument melted. The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- This complaint was initially reported as an adverse event with a life-threatening injury in error on (b)(6) 2013 and should have been reported as a malfunction. Since the patient was undergoing a da vinci si hysterectomy procedure, the alleged uterus burn was not a life threatening injury. Corrections: unchecked adverse event, unchecked life threatening, unchecked serious injury, and checked malfunction. | PERMANENT CAUTERY HOOK INSTRUMENT | Unkown | 3186232 | 3157021 | 3053556 | 2955842201302048 | 1 | NAY | Manufacturer | Company RepresentativeOther | Other | InitialFollowup | 2013-08-05 | 2013-06-08 | Yes | Yes | Health Professional | 42018306 | No | No | 2013-08-05 | Yes | No |  | 2013
42027802 | 2013-05-10 | Malfunction |  | It was reported that during reprocessing that a cable was popping out of a grasping retractor instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with a pitch cable frayed and derailed at the distal clevis hub. No damage was found on the clevis. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3186290 | 3157088 | 3053623 | 2955842201302038 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-15 | 2013-06-08 | No | Yes | Health Professional | 42027802 | Device Returned To Manufacturer | No | 2013-05-15 | Device Not Retu | No | Reuse | 2013
IS3000A60P8 | 2010-08-01 | Injury | Required Intervention | On (b)(4) 2013, the initial reporter of this complaint contacted intuitive surgical inc. (isi) alleging that the patient was injured during a da vinci hysterectomy procedure performed in (b)(6) 2010 at (b)(6). Isi contacted the initial reporter to obtain additional information regarding the reported event. According to the initial reporter, the surgical procedure went fine. However, the initial reporter claimed that after the patient went home, she experienced acute abdominal pain. The patient was subsequently seen and treated by a physician at (b)(6) on an unspecified date/time. Approximately 10 days later, a cat scan was performed on the patient and it was discovered that the patient had sustained an injury to her bowel. The initial reporter stated that he was told that the injury to the bowel was an electrical burn. Additional surgeries to repair the patient's bowel were performed. The initial reporter indicated that the patient had recovered from the surgeries. Isi has no previous record of this complaint. If additional information can be obtained it will be provided in a supplemental report. | Intuitive surgical has attempted to contact the risk management department of the medical center involved with this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. Based on the information, this complaint is being reported due to the following conclusion: the initial reporter claimed that the patient sustained a bowel injury during a da vinci hysterectomy procedure. However, at this time it is unknown if a malfunction of the da vinci si system, an instrument, and/or an accessory occurred during the surgical procedure. It is also unknown what caused the reported injury to the patient's bowel. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3186322 | 3157108 | 3053643 | 2955842201302055 | 1 | NAY | Manufacturer | ConsumerOther | Other | FollowupFollowup | 2013-08-09 | 2013-06-08 | Yes | No | Health Professional | IS3000A60P8 | No | No | 2013-08-09 | No | No | Reuse | 2013
40027202 | 2013-04-17 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the harmonic ace curved shears insert worked for about the 20 minutes and then the generator would not engage. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint with the following clarification: the insert failed a self-test on the generator. A functional energy test was performed using a harmonic generator and handpiece. The insert failed a self-test, causing the generator to produce an instrument error message. A high pitched noise was heard during the self test. The insert was disassembled and a hairline crack was found on the blade. The crack was located adjacent to the clamp arm pin. A wear mark was also observed on the blade in the same area as the clamp arm pin, suggesting pin and blade made contact during harmonic energy activation, likely contributing to crack propagation. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3180036 | 3151266 | 3047801 | 2955842201302041 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-15 | 2013-06-07 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2013-05-15 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-05-15 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure a wire came out of distal end of the fenestrated bipolar forceps instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint with the following clarification, the nonconformance was a broken pitch cable. The pitch cable was broken at the instrument's wrist. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3180038 | 3151267 | 3047802 | 2955842201302042 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-15 | 2013-06-07 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-15 | Device Not Retu | No | Reuse | 2013
40027202 | 2013-04-17 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the harmonic ace curved shears insert worked for about the 20 minutes and then the generator would not engaged. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint with the following clarification: the insert failed a self-test on the generator. A functional energy test was performed using a harmonic generator and handpiece. The insert failed a self-test, causing the generator to produce an instrument error message. A high pitched noise was heard during the self test. The insert was visually inspected under magnification and a hairline crack was found on the blade. The crack was located. 400 below the distal tip. A gouge and scratch was also observed on the blade edge in the same area as the crack. Engineering concluded based on location of damage, the gouge was likely caused by mishandling/misuse. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | HARMONIC ACE CURVED SHEARS INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3179965 | 3151284 | 3047801 | 2955842201302040 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-15 | 2013-06-07 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2013-05-15 | Device Not Retu | No | Reuse | 2013
VER07 8MM | 2013-05-23 | Malfunction |  | During robotic case, the end of the prograsp forceps instrument (while inside pt) had shivers or shreds of wire separating from the actual instrument. There were also more slivers inside the pt not attached to the instrument. |  | DAVINCI PROGRASP FORCEPS | DAVINCI PROGRASP FORCEPS | 3199065 | 3167578 | 3064133 | MW5030512 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-06-07 | 2013-06-07 | No | Yes | Health Professional | VER07 8MM | Yes | Yes |  |  |  |  | 2013
42019408 | 2013-05-13 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure a cable broke on a mega needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken near the proximal pulleys and proximal clevis cable hole. The idler pulley spun freely and did not exhibit damage. The broken cable segment stuck out at the instrument's wrist. No major wear was found on the proximal clevis cable hole. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3178484 | 3149501 | 3046036 | 2955842201302012 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-14 | 2013-06-06 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2013-05-14 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-05-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument was noted to have a crack. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was cracked at the proximal clevis interface. The clevis was not dislodged from the tube extension. The evidence was inconclusive, but the tube extension likely broke due to excessive side loading or other mishandling. Engineering also found a bent banana plug at the back of the chassis. The evidence was inconclusive, but the damage was likely due to mishandling. The instrument passed electrical continuity testing. Additional observation that was not reported by the customer was a micro-crack found near the reinforcement ring. These types of micro-cracks will not lead to mechanical failure of the instrument, however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of these is confined to 2cm of the distal end of the instrument shaft. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event, in a fragment-causing failure (such as a cable failure, a crimp separating from a cable etc. ) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument; however, the damage to the main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SHEARS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3178782 | 3149901 | 3029401 | 2955842201302013 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-05-06 | Malfunction |  | It was reported that prior to starting a da vinci si procedure the monopolar curved scissors (mcs) instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. The instrument was checked for recognition on an in-house system and the instrument was successfully recognized on multiple attempts. The pogo pins were not stuck or contaminated. Engineering was unable to replicate recognition issue. The instrument was driven and moved intuitively with full range of motion in all directions. The instrument passed electrical continuity testing. Additional observation that was not reported by the customer were micro-cracks found at the distal end of the tube. The micro-cracks run in the axial direction and measures approximately 0. 334. These types of micro-cracks will not lead to mechanical failure of the instrument, however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of these is confined to 2cm of the distal end of the instrument shaft. | MONOPOLAR CURVED SHEARS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3178856 | 3150014 | 3046549 | 2955842201302014 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-05-06 | Malfunction |  | It was reported that prior to starting a da vinci s procedure the maryland bipolar forceps instrument wire was noted to come off track. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the nonconformance was loose pitch cables. The intuitive motion was not functioning due to loose pitch cables found at the distal clevis hub. One of the loose cable segments had become frayed as well. The clevis did not exhibit any damage or wear marks. Both cable crimps remained in clevis. The housing was removed and the pitch cables were found loose at the clamping pulley, but no loose clamping pulley screw was found. The housing interior also showed signs of excessive residue around the clamping pulley and corrosion to the bearings. Engineering also found the main tube insulation appeared to have excessive wear where material loss is evident. The surface of the main tube was no longer smooth but rough in nature. The evidence was inconclusive, but the damages to the instrument were likely due to improper cleaning. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3178917 | 3150155 | 3046690 | 2955842201302015 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-06 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-29 | Malfunction |  | It was reported that during a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument was noted to be dull. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the scissors were dull. The instrument was placed on an in-house system for a latex cut test and the scissors did not cut cleanly through (b)(4) latex. The blades were undamaged. The blades exhibited wear at the tips affecting cut performance. The instrument passed electrical continuity testing. Additional observation that was not reported by the customer were micro-cracks found at the distal end of the tube. The micro-cracks run in the axial direction. These types of micro-cracks will not lead to mechanical failure of the instrument, however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of these is confined to (b)(4) of the distal end of the instrument shaft. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SHEARS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179295 | 3150565 | 3047100 | 2955842201302016 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-10 | 2013-06-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42018907 | 2013-05-01 | Malfunction |  | It was reported that during a da vinci si surgical procedure the double fenestrated grasper instrument failed to work properly after putting the instrument on the robot arm. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. The instrument was installed on an in-house is3000 system and the system showed the instrument was expired. Engineering was unable to drive the instrument on the system, but grips could be opened and closed when input discs were manually rotated. Engineering also found tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 170 in length and not aligned with the tube axis. The evidence was inconclusive, but the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3179353 | 3150605 | 3047140 | 2955842201302017 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-10 | 2013-06-06 | No | Yes | Health Professional | 42018907 | Device Returned To Manufacturer | Yes | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-05-09 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure that wires broke on a mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the surface. The broken cable segment stuck out at the instrument's wrist. Other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3179380 | 3150639 | 3047174 | 2955842201302018 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-06 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-05-14 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure some shavings came off the shaft of a fenestrated bipolar forceps instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The main tube insulation had scratches and gouge marks with material removed at the distal end. Engineering concluded the damage was likely caused by excess force contact on the main tube surface. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179426 | 3150657 | 3047192 | 2955842201302020 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-14 | 2013-06-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-14 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-05-10 | Malfunction |  | It was reported that during a da vinci s surgical procedure the pk dissecting forceps instrument was noted to have broken wires. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179386 | 3150667 | 3047202 | 2955842201302019 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-06 | No | Yes | Health Professional | 42022703 | No | No | 2013-05-10 | No | No | Unkown | 2013
42014204 | 2013-05-10 | Malfunction |  | Through the product notification, the customer had taken initiative to inspect the monopolar cautery instrument through a high mag resolution and had identified a micro crack. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint through the product notification, the customer had taken initiative to inspect the item through a high mag resolution and had identified a micro-crack. The main tube had signs of abrasion/wear that exhibited damage to the insulation but no micro-cracks were found. Two hairline scratches and one small circular hole found near the distal end. No materials were removed or found missing. The instrument passed electrical continuity testing. No other damage was found. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179556 | 3150789 | 3047324 | 2955842201302022 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-06 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the fenestrated bipolar forceps instrument was noted to have a broken cable. The instrument was exchanged into a backup instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch up cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The pitch down cable was frayed at the clevis hub. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179527 | 3150790 | 3047325 | 2955842201302021 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-05-10 | 2013-06-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-05-09 | Malfunction |  | It was reported that during a da vinci si surgical procedure the mega suturecut needle driver instrument it was noted that a little wire popped out. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. There was an indentation at the edge of the distal pulley. There were visible scratches on the surface of the pulley. The evidence is inconclusive but the damage was most likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3179531 | 3150796 | 3047331 | 2955842201302023 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-06 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci s sacrocolpopexy procedure the mega suturecut needle driver instrument wire was noted to be exposed and broken. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. There were indentations at the edge of the distal pulley. There were visible scratches on the surface. The evidence is inconclusive but the damage was most likely due to mishandling. Engineering also found the distal end of the main tube had a scratch mark showing light material removal. The scratch was short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage may have been due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3179594 | 3150821 | 3047356 | 2955842201302024 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-06 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2010-02-03 | Injury | Disability | On (b)(6) 2013, intuitive surgical inc. (isi) received a legal complaint concerning a patient who underwent a da vinci si hysterectomy and cholecystectomy procedure on (b)(6) 2010. The legal filing claims that at some point during the da vinci si surgical procedure the patient's right ovary was removed without her consent or knowledge. The patient reportedly discovered that her right ovary was missing after undergoing a surgical procedure on (b)(6) 2013 to remove her left ovary. After completion of the surgical procedure performed on (b)(6) 2013, the surgeon informed the patient that there was scar tissue where her right ovary should have been, and that the right ovary was gone. The legal filing indicates that the only surgical procedure the patient had prior to (b)(6) 2013, was the da vinci procedure performed on (b)(6) 2010. | Based on the information provided, there is no evidence that a malfunction of the da vinci system, an instrument, or an accessory caused or contributed to the patient's reported injury. There is no allegation that a malfunction of the da vinci si surgical system occurred during the surgical procedure performed on (b)(6) 2010. At this time, it is unknown how or when the patient's reported injury actually occurred. According to the legal filing, it was noted in the patient's medical records for the surgical procedure performed on (b)(6) 2010, that her ovaries were normal and their vascular supply was uninterrupted. Intuitive surgical inc. (isi) has attempted to contact the risk management department of the site to obtain additional information regarding the reported event. However, as of the date of this report, no additional information has been provided. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3179604 | 3150836 | 3047371 | 2955842201302026 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-05-07 | 2013-06-06 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-05-07 | No | No | Reuse | 2013
42017910 | 2013-05-03 | Malfunction |  | It was reported that during a da vinci si surgical procedure the robot did not recognize the monopolar curved scissors (mcs) instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. The instrument was checked for recognition on an in-house is13000 system, the system successfully recognized the instrument, showing zero uses left, and displaying an instrument was expired message. The pogo pins were not stuck or contaminated. The dcp (dallas chip programmer) verification of the instrument passed. Engineering was unable to replicate the recognition issue. Engineering also found the distal end of the main tube had various scratch marks with light material removal. The scratches were. 090 -. 250 in length and not aligned with the tube axis. The evidence was inconclusive, but the damage may have been caused by mishandling/misuse. Additional observation that was not reported by the customer were micro-cracks found at the distal end of the tube. The micro-cracks run in the axial direction and measures less than 0. 350. These types of micro-cracks will not lead to mechanical failure of the instrument, however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of these is confined to 2cm of the distal end of the instrument shaft. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179598 | 3150845 | 3035223 | 2955842201302025 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-13 | 2013-06-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No | Reuse | 2013
42004906 | 2013-05-01 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the cadiere forceps instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument. The pogo pins were not stuck or contaminated. There was an indentation on the edge of the distal pulley. The evidence is inconclusive but most likely due to mishandling. There were visible scratches on the surface of the pulley. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3179644 | 3150880 | 3047415 | 2955842201302027 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-13 | 2013-06-06 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-05-08 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have something wrong with the cable; the cable was loose. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was frayed at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. The instrument passed electrical continuity testing. Engineering also found the one grip was bent, causing side-to-side misalignment of grips. There was a. 080 offset at the tips, indicating overloading at the tip. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was most likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179735 | 3150955 | 3042692 | 2955842201302028 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-13 | 2013-06-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No | Reuse | 2013
42020702 | 2013-05-09 | Malfunction |  | It was reported that during a da vinci s surgical procedure the tenaculum forceps instrument plastic part of the shaft near the tip was noted to be snapped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was broken at the proximal clevis to the main tube interface. The clevis was dislodged from the main tube as a result. Both the proximal clevis and the main tube were missing pieces. The instrument was built in 2006, prior to implementation of the metal proximal clevis. The evidence was inconclusive, but the breakage was likely due to overloading at the tip. Engineering also found main tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 100 -. 195 in length and not aligned with the tube axis. The evidence was inconclusive, but the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3179767 | 3151006 | 3047541 | 2955842201302029 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-13 | 2013-06-06 | No | Yes | Health Professional | 42020702 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-05-09 | Malfunction |  | It was reported that during a da vinci s surgical procedure the pk dissecting forceps instrument was noted to have a separated cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. There were indentations at the edge of the distal pulley. There were visible scratches on the surface of pulley. The evidence was inconclusive but most likely due to mishandling. Engineering also found the distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179779 | 3151017 | 3040128 | 2955842201302030 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-13 | 2013-06-06 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-05-13 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical fenestrated bipolar forceps instrument was noted to have a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the nonconformance was a broken pitch cable. The pitch cable was broken at the wrist. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179781 | 3151019 | 3047554 | 2955842201302031 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-13 | 2013-06-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-05-06 | Malfunction |  | It was reported that during a da vinci si inguinal hernia repair procedure a wire broke on a mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the surface. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3179783 | 3151021 | 3047556 | 2955842201302032 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-05-15 | 2013-06-06 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-15 | No | No | Reuse | 2013
42022703 | 2013-05-14 | Malfunction |  | It was reported that during a da vinci s surgical procedure the pk dissecting forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179791 | 3151026 | 3047570 | 2955842201302034 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-06-11 | 2013-06-06 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-06-11 | Yes | No | Reuse | 2013
42022703 | 2013-05-14 | Malfunction |  | It was reported that during a da vinci s surgical procedure the pk dissecting forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179796 | 3151031 | 3047570 | 2955842201302035 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-06-11 | 2013-06-06 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-06-11 | Yes | No | Reuse | 2013
42022703 | 2013-05-14 | Malfunction |  | It was reported that during a da vinci s surgical procedure the pk dissecting forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179799 | 3151035 | 3047570 | 2955842201302036 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-06-11 | 2013-06-06 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-06-11 | Yes | No | Reuse | 2013
42020706 | 2013-05-14 | Malfunction |  | It was reported that during a da vinci si surgical procedure the tenaculum forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the wrist. The cable segment that contained the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Engineering also found abnormal wear throughout the entire main tube. The surface of the main tube is rough in nature, discolored, and splintering off. Material loss is evident. The evidence was inconclusive, but the damage was most likely due to improper cleaning process. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3179803 | 3151039 | 3047574 | 2955842201302037 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-14 | 2013-06-06 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-05-14 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-05-10 | Malfunction |  | It was reported that during a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch up cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at wrist were undamaged. Engineering also found main tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 250 in length and not aligned with the tube axis. The evidence was inconclusive, but the damage may have been caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3179805 | 3151041 | 3047576 | 2955842201302039 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-13 | 2013-06-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No | Reuse | 2013
| 2013-05-30 | Malfunction |  | During surgery, the vessel sealer stopped working. The alarm on the monitor said, "blade may be exposed". No harm to patient. ======================manufacturer response for vessel sealer, endowrist one (per site reporter). ======================the local rep at this facility picked up previous instruments reported (4). He reports " user error" despite different users. I have yet to receive any written evaluation from the manufacturer. What was the original intended procedure?gyn procedure. Device #1is this a laboratory device or laboratory test?no. |  | DAVINCI ENDOWRIST ONE | INSTRUMENT COMPUTER CONTROLLED | 3196321 | 3167136 | 3063691 | 3167136 | 1 | NAY | User Facility |  |  | Initial | 2013-06-06 | 2013-06-06 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
| 2013-05-24 | Malfunction |  | none | none | DAVINCI MONOPOLAR CAUTERY | SYSTEM SURGICAL COMPUTER | 3202587 | 3173371 | 3069926 | 3173371 | 1 | NAY | User Facility |  |  | Initial | 2013-06-05 | 2013-06-06 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2013
| 2013-05-24 | Malfunction |  | Robotic total abdominal hysterectomy in progress when surgical team noticed a small black object in abdominal cavity, visible on monitor. Object was removed and inspected; black in color and hard to the touch. Tech removed instruments from trocars. 5mm monopolar cautery device removed from 8mm davinci trocar and examined. Outer, black covering of device had an approximately 1. 5 mm area on sheath that was frayed and damaged. Device removed from sterile field. Abdomen was irrigated and suctioned. What was the original intended procedure?total abdominal hysterectomy. Device #1is this a laboratory device or laboratory test?no. |  | DAVINCI MONOPOLAR CAUTERY | SYSTEM SURGICAL COMPUTER CONTR | 3206436 | 3176943 | 3073498 | 3176943 | 1 | NAY | User Facility |  | Invalid Data | Initial | 2013-06-05 | 2013-06-06 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2013
42009308 | 2013-04-05 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure that the prograsp forceps instrument was inserted into arm 1 and was noted not to work well. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. There was an indentation at the edge of the distal pulley and some scratches on the surface of the distal pulley. Engineering also found the distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3175924 | 3146992 | 3043527 | 2955842201301977 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-05-08 | 2013-06-05 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-29 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure the fenestrated bipolar forceps instrument was noted not to be cauterizing. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of does not cauterize. Visual inspection showed no damage to the conductor wires at the wrist. The instrument passed electrical continuity testing. A functional test was performed using an is3000 system, (b)(4) generator and a bipolar cord. A wet paper towel was placed in the grips, the bipolar foot pedal was pressed and towel began to smoke, indicating successful energy transfer from the generator to the tips. Engineering also found tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 150 in length and not aligned with the tube axis. The evidence was inconclusive, but the damage was likely caused by mishandling/misuse. No other damage was found. This report does not admit that the report or information submitted under the instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3175907 | 3147025 | 3043560 | 2955842201301981 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-06-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42020906 | 2013-04-19 | Malfunction |  | It was reported that during a da vinci si nissen fundoplication procedure, it was noted a cable was loose after attempts to use a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with a grip cable loose at the distal clevis. The grip cables were found broken at the idler block inside the housing. Engineering concluded the damage was likely due to mishandling during the cleaning process. Signs of corrosion were found on the clamping pulley. Additional findings were scratches on the surface of the distal pulley. The distal end of the main tube also had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Some scratches were towards the instrument tip area and one was at the end of housing area. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3176670 | 3147682 | 3044217 | 2955842201301991 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42020906 | Device Returned To Manufacturer | Yes | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-05-10 | Malfunction |  | It was reported that the customer identified a micro crack on a monopolar curved scissors instrument using a high magnified scope. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that two small micro cracks were found on opposites sides of the main tube near tube extension. Instrument passed electrical continuity test. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2cm of the distal end of the instrument shaft. Additional observation not reported by site was main tube scratches. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were short in length approximately 0. 0405 and they were not aligned with the tube axis. Engineering concluded that the main tube scratches were likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a mdr reportable event; however; the main tube scratches and cracks, found during failure analysis evaluation may cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3176847 | 3147846 | 3044381 | 2955842201301992 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-05 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-26 | Malfunction |  | It was reported that during a da vinci s low anterior resection procedure the monopolar curved scissors (mcs) instrument wire snapped at the tip of the instrument. The surgical staff exchanged it into a new instrument and the planned procedure was completed without problem. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the wrist. Any potential fragments would likely be retained by the use of the tip cover. The other cables at the wrist were undamaged. Engineering also found main tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 100 -. 140 in length and not aligned with the tube axis. The evidence was inconclusive, but the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event, in a fragment-causing failure (such as a cable failure, a crimp separating from a cable etc. ) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument; however, the scratches on the main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3176985 | 3147900 | 3044435 | 2955842201301993 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-05-03 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure the monopolar curved scissors (mcs) instrument was noted to be dull. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. The instrument was placed on an in-house is3000 system for a latex cut test and the scissors cut cleanly through (b)(4) latex. The blade edges were undamaged. Additional observation that was not reported by the customer was a cracked main tube. The distal end of the main tube exhibited a couple of hairline cracks in the axial direction. The cracks were less than. 450 in length. These types of micro-cracks will not lead to mechanical failure of the instrument, however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of these is confined to 2cm of the distal end of the instrument shaft. No other damage found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3177053 | 3147950 | 3044485 | 2955842201301994 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-04-29 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the tenaculum forceps instrument wires broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one pair of grip cables was loose, but not visibly broken at the wrist. The grip motion was not intuitive as a result of loose cables. The housing was removed to find one grip cable was broken at the clamping pulley cable groove. The clamping pulleys exhibited residue around the cables. Additional observations not reported by the customer were main tube damage and corroded clamping pulleys. The main tube exhibited wear with light material removal and a rough surface finish throughout the entire tube length. The epoxy layer on the main tube had been removed, exposing glass fibers. The clamping pulleys exhibited pitting corrosion upon closer inspection. The evidence was inconclusive, but the tube and the clamping pulley damage were likely due to improper cleaning. No other damage was found. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177105 | 3148020 | 3044555 | 2955842201301995 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-05-06 | Malfunction |  | It was reported that prior to starting a da vinci si procedure the pk dissecting forceps instrument would not engage. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. The instrument was checked for recognition and engagement on an in-house system. The system successfully recognized instrument and passed engagement. Additional damage found was a broken pitch cable at the distal end. The cable segment that contained the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. The instrument passed electrical continuity testing. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3177107 | 3148022 | 3040128 | 2955842201301996 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-05-09 | Malfunction |  | It was reported that prior to starting a da vinci si procedure the mega suturecut needle driver instrument cable at the tip was noted to be broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. Engineering also found indentations at the edge of the distal pulley. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive but most likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177193 | 3148107 | 3044642 | 2955842201301997 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-05-09 | Malfunction |  | It was reported that during reprocessing the da vinci s monopolar cautery instrument was noted to have a hole in the insulation. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the black tube insulation was damaged at the distal end. The insulation was gouged on one side and had a. 110 x. 065 piece missing roughly 3. 65 above the snake wrist. There were also a couple of sections of insulation lifted up below the missing section. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3177214 | 3148109 | 3044644 | 2955842201301998 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42014304 | 2013-05-09 | Malfunction |  | It was reported that during reprocessing a hole was found in the insulation. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. Black tube insulation was damaged at the distal end. The insulation was gouged on one side and had a. 180 x. 095 piece missing roughly 5 above the snake wrist. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177250 | 3148160 | 3044695 | 2955842201301999 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42011704 | 2013-05-09 | Malfunction |  | It was reported that during reprocessing a needle driver instrument, insulation was noted damaged. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. Black tube insulation was damaged at the distal end. The insulation was gouged on one side and had a. 040 x. 040 piece missing roughly 4 above the snake wrist. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177257 | 3148168 | 3044703 | 2955842201302000 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42805312 | 2013-05-01 | Malfunction |  | It was reported that during a da vinci si surgical procedure the hem-o-lok medium large clip applier instrument was allegedly not holding clips. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was placed on an in-house system and driven, recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. The grips opened, closed and grasped properly. The instrument was fully functional. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | HEMOLOK MEDIUM LARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177259 | 3148170 | 3044705 | 2955842201302001 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42805312 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-05-10 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the monopolar curved scissor's instrument broke under near the orange portion. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension was found to be broken and missing a piece at the main tube interface. The tube extension was also found to be dislodged from the main tube. Engineering concluded that the tube extension likely broke due to excessive side loading or other mishandling. Additional observations not reported by site were micro cracks and reinforcement ring tube damage. A small micro crack was found on the main tube near tube extension. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2cm of the distal end of the instrument shaft. Engineering also observed two indentations on tube reinforcement ring. 80 - last observation not reported by site was scratches on the main tube. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a mdr reportable event; however; the micro cracks, scratches on the main tube and reinforcement ring tube damage, found during failure analysis evaluation may cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3177266 | 3148177 | 3044712 | 2955842201302002 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-10 | 2013-06-05 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-05-01 | Malfunction |  | It was reported that during a da vinci si surgical procedure the bowel grasper instrument was noted to have insulation damage. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177273 | 3148183 | 3036923 | 2955842201302004 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42017704 | No | No | 2013-05-09 | No | No | Unkown | 2013
42017704 | 2013-05-01 | Malfunction |  | It was reported that during a da vinci si surgical procedure the bowel grasper instrument was noted to have insulation damage. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177276 | 3148186 | 3036923 | 2955842201302003 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-05 | No | Yes | Health Professional | 42017704 | No | No | 2013-05-09 | No | No | Unkown | 2013
42009308 | 2013-02-28 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure when retrieving the prograsp forceps instrument from the trocar, the pulley blocked the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of the pulley blocking the trocar when retrieving the instrument. Engineering also found scratches at the distal end of the main tube, approximately 0. 53 in length. The evidence was inconclusive, but the damage was likely due to mishandling. The instrument pitch cable was also found derailed at the idler block pulley, inside the housing, resulting in the pitch cable being loose at the distal clevis hub. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177313 | 3148229 | 3044764 | 2955842201302005 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-10 | 2013-06-05 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-17 | Malfunction |  | It was reported that during a da vinci si surgical procedure the mega suturecut needle driver instrument cable was noted to be in a loop. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found a broken grip cable. The broken grip cable stuck out at the wrist. The idler pulley spun freely and exhibited pulley rim and face damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177344 | 3148240 | 3044775 | 2955842201302008 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42805812 | 2013-04-18 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure the fundus grasper instrument was noted to have a bent shaft. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was found with the main tube bent near the distal end, approximately 3. 8 from the base of the grip. Additional observation found instrument main tube with material removed near the bending point. The roll gear was found broken at the base of the main tube. The evidence was inconclusive, but the damages were likely due to mishandling. General precautions and warnings - single-site instruments should be handled and operated by trained personnel. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Always have a backup instrument available to complete the surgical procedure in case of instrument failure. Clean the instruments immediately after each use. Do not allow debris to dry on or inside the instrument intraoperatively before instrument processing. In order to keep the instrument from drying when soiled, keep the instrument in water or an enzymatic bath between the surgical procedure and instrument processing. The instrument may also be flushed through the main flush port with sterile water during use to minimize buildup of internal deposits of bio-material. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | FUNDUS GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177392 | 3148304 | 3044839 | 2955842201302006 | 1 | GCJ | Manufacturer | User facility | Nurse | Initial | 2013-05-10 | 2013-06-05 | No | Yes | Health Professional | 42805812 | Device Returned To Manufacturer | Yes | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-03-28 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure there was an unspecified problem with a small grasping retractor instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. No reason for return was provided. A complaint cannot be confirmed or denied. Upon failure analysis, both pitch cables were observed to be broken at the wrist. Both cable segments that contain the crimp were still installed in clevis. Clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177394 | 3148306 | 3044841 | 2955842201302007 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-05 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-05-13 | Malfunction |  | It was reported that during reprocessing a mega suturecut needle driver instrument, fibers on the end of instrument were showing. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The customer initially reported fibers showing at the end of the instrument; however, for clarification, the nonconformance was a broken grip cable. The broken grip cable stuck out at the instrument's wrist. The idler pulley spun freely and exhibited pulley rim damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177406 | 3148323 | 3044858 | 2955842201302009 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-13 | 2013-06-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-23 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, a wire was in a loop on a prograsp forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Additional findings were scratches on the surface of the distal pulley. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177407 | 3148344 | 3044879 | 2955842201302010 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-13 | 2013-06-05 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-05-03 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure a broken cable was noted on a large suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. An additional finding was the distal end of main tube had scratch mark showing light material removal. The scratch was short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3177410 | 3148347 | 3044882 | 2955842201302011 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-06-05 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-17 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tips were misaligned on a maryland bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip was bent, causing a side to side misalignment of the grips. There was a. 0345 offset at the instrument's tips. Engineering concluded the damage was likely due to mishandling. Additional damage found were deep scratches. The distal end of the main tube had several scratch marks exhibiting light material removal and a rough surface finish. The scratch marks varied in size approximately from. 045 to. 2234 long. Engineering concluded the damage may be due to mishandling. Electrical continuity test passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3172776 | 3144215 | 3040750 | 2955842201301972 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-06-04 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
40011010 | 2013-04-22 | Malfunction |  | It was reported that during reprocessing the da vinci s surgical precise bipolar forceps instrument a little metallic piece was found. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The complaint of after sterilization, a little metallic piece was found was confirmed with the following clarification: the metal piece likely does not belong to the returned instrument. The instrument was returned with a. 185 long roll pin taped to the housing. The roll pin was similar in size to the quantity four roll pins installed in the chassis to hold release levers in place. Visual inspection showed all four roll pins were installed in the chassis. There were no other roll pins that were used on the instrument, suggesting the roll pin was from a different instrument. Engineering also found tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 060 -. 200 in length and not aligned with the tube axis. The evidence was inconclusive, but the damage was likely may have been caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3174152 | 3145519 | 3042054 | 2955842201301973 | 1 | NAY | Manufacturer | ForeignUser facility | PHARMACIST | Initial | 2013-05-06 | 2013-06-04 | No | Yes | Health Professional | 40011010 | Device Returned To Manufacturer | Yes | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-08 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure the prograsp forceps instrument was noted to have restricted movement. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis, which was likely the reason of lack of intuitive movement. The cable strand stuck out at the wrist. Engineering also found a derailed cable, deep scratches, corroded clamping pulleys and back idler pulleys. The one grip cable was derailed off the distal idler pulley. The cable was still tensioned but was also frayed because of contact from rubbing against the idler pulley. Engineering also found the distal end of main tube had several scratch marks exhibiting light material removal and a rough surface finish. The evidence was inconclusive, but the damage was likely due to mishandling. All three clamping pulleys exhibited corrosion/contamination around the grooves and screws. The back idler pulleys exhibited yellowish material on the surface of each. The evidence was inconclusive, but the damage was likely due to improper cleaning. No other damage was found. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3174314 | 3145666 | 3042201 | 2955842201301974 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-05-07 | 2013-06-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-07 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-29 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the large needle driver instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3174320 | 3145672 | 3042207 | 2955842201301975 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-06-24 | 2013-06-04 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-06-24 | Yes | No | Reuse | 2013
42009308 | 2013-05-07 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure that the prograsp forceps instrument, where the tip meets the shaft, the instrument was noted to have broken wires out of the sterile kit. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch up cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. Engineering also found tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 170 in length and not aligned with the tube axis. The evidence was inconclusive, but the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3174488 | 3145808 | 3042343 | 2955842201301976 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-07 | 2013-06-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-07 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-26 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical monopolar curved scissors (mcs) instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the wrist. Any potential fragments would likely be retained by use of tip cover. The other cables at the wrist were undamaged. Engineering also found main tube scratches and cracking. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 100 -. 160 in length and not aligned with the tube axis. The evidence was inconclusive, but the scratches were likely caused by mishandling/misuse. Additional observation that was not reported by the customer was micro-cracks found at the distal end of the tube. The micro-cracks run in the axial direction. These types of micro-cracks will not lead to mechanical failure of the instrument, however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of these is confined to 2cm of the distal end of the instrument shaft. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event, in a fragment-causing failure (such as a cable failure, a crimp separating from a cable etc. ) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument; however, the scratches on the main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3174713 | 3145998 | 3035278 | 2955842201301978 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-06-04 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
40011010 | 2013-04-19 | Malfunction |  | It was reported the isolation of the bipolar bottom was completely unbound on a precise bipolar forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, the bipolar insert was missing from the plug riser. Pins and conductor wires were sticking out of the plug riser. Electrical continuity test passed. This complaint case is non reportable related to the bipolar insert missing from the plug riser as this instrument and finding will be tracked as part of the field action number (b)(4). Additional damages found were deep scratches and a cracked main tube. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The main tube also exhibited small hairline cracks on the proximal end. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3174789 | 3146092 | 3042627 | 2955842201301979 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-08 | 2013-06-04 | No | Yes | Health Professional | 40011010 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-30 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the pk dissecting forceps instrument tips were noted to misaligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of the grips. There was a. 053 offset at the tips, indicating overloading at the tip. The instrument passed electrical continuity testing. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3174805 | 3146105 | 3042640 | 2955842201301980 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-06-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-05-07 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the prograsp forceps instrument jaws were noted not to close all the way and have frayed wires at the distal end. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was frayed at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 03 in length. The one grip close cable was found frayed as well at the distal clevis pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3174827 | 3146133 | 3042668 | 2955842201301983 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-06-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-05-14 | Malfunction |  | The user facility reportedly identified frayed cables at the distal end of the large suturecut needle driver instrument after it was cleaned and sterilized. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed there was a broken grip close cable at the distal idlers. The idler pulley spun freely and did not exhibit any damage. The cable segment was sticking out at wrist. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3174810 | 3146144 | 3042679 | 2955842201301982 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-14 | 2013-06-04 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-05-14 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-28 | Malfunction |  | A maryland bipolar forceps instrument was returned to intuitive surgical. No complaint was communicated to intuitive surgical regarding performance of the da vinci surgical system. No malfunction of the da vinci system was alleged. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. No failure was reported for this instrument. Engineering found bent grip tips, bent bipolar pins, and deep scratches on main tube. One grip was bent, causing a side to side misalignment of grips. There was a. 1235 offset at the instrument tips. Engineering concluded the damage was likely due to mishandling. The bottom bipolar pin was bent at the back of the chassis. Engineering concluded the damage was likely due to mishandling. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. The scratches were short in length approximately. 06 to 08. Engineering concluded the damage may be due to mishandling. Electrical continuity passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering's findings does not itself constitute a mdr reportable event; however; the tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3174837 | 3146150 | 2975602 | 2955842201301984 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-07 | 2013-06-04 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-07 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-05 | Malfunction |  | The user facility reportedly identified a broken cable at the tip of the fenestrated bipolar instrument prior to starting a da vinci si surgical procedure. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed there was a broken conductor wire at the yaw pulley exit. The wire was detached from its connection at the grip. Electrical continuity testing failed. The yaw pulley was missing a piece at the opposite area of the conductor break allowing the grip to move within the pulley and have a larger range of motion in yaw. Engineering was unable to conclude how the piece of yaw pulley broke. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3174857 | 3146156 | 3042691 | 2955842201301985 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-16 | 2013-06-04 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-16 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-05-08 | Malfunction |  | It was reported that during a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have wires tearing. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch up cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3174841 | 3146157 | 3042692 | 2955842201301986 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-06-04 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-18 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure a maryland bipolar forceps instrument jaws would not open. When the instrument was taken out of the patient, they noticed that the brown part looked like it melted. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. Both yaw pulleys exhibited charring and localized melting at the grip base, between the grips. Charring is a sign of an arcing event. Evidence not conclusive, but charring may be due to a breach in the insulation, carbonized tissue creating a conductive path, or high generator settings. The grips could not open with a normal range of motion due to thermal deformation on the sides of the yaw pulleys. Internal arcing in bipolar instruments results in a potential loss of product function, but no transfer of energy to the patient. Intuitive has concluded that the likelihood of injury due to this failure is remote. User are instructed to retain back-up instruments in case that an instrument fails to perform, and are also instructed in proper technique to minimize the likelihood of internal arcing additional observation not reported by site was tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 060 -. 160 in length and not aligned with the tube axis. Engineering concluded the damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3174859 | 3146160 | 3042695 | 2955842201301987 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-07 | 2013-06-04 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-07 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-25 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the pk dissecting forceps instrument was noted to have a separated, broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one conductor wire was broken at the yaw pulley exit. The wire was detached from the connection at the grip. The instrument failed electrical continuity testing. The yaw pulley showed no signs of arcing. No other damage was found. The yaw pulley was missing a. 152 x. 055 piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. There were indentations at the edge of the distal pulley. The evidence was inconclusive but most likely due to mishandling. There were visible scratches on the surface of the pulley. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3174899 | 3146193 | 3042728 | 2955842201301989 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-06-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-05-08 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the prograsp forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the grip cable was frayed. The frayed cable sections were in various lengths and stuck out at the wrist. The idler pulley exhibited pulley rim and face damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3174913 | 3146194 | 3042729 | 2955842201301988 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-06-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-03-28 | Malfunction |  | The grasping retractor instrument was returned without any allegations of malfunction. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. No failure was reported for this instrument. Engineering evaluation found deep scratches, the luer plate missing, and broken snaps off the housing. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. The scratches were short in length approximately 0. 0432 to 0. 1280. The instrument was found with the luer plate missing from the chassis. Three of four housing snap tabs and housing pins were broken off the housing. The housing could be easily removed as a result of the damage. The evidence was inconclusive, but the damage was likely due to mishandling. No other damage was found. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3174910 | 3146203 | 3042738 | 2955842201301990 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-04 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2013-05-02 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the lens was noted to be missing on the endoscope. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the endoscope was received with mechanical damage to the proximal end resulting in broken optical components and 3d optical misalignment. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3169778 | 3141033 | 3037568 | 2955842201301947 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-05-02 | 2013-06-03 | No | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-05-03 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the instrument was noted to have a frayed wire and was disconnected from the working end of the grasper. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip cable was frayed at the distal end and lacked intuitive movement. The frayed strands stuck out at the wrist. No damage was found on the distal idler pulley. The other cables at the wrist were undamaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3169788 | 3141046 | 3037581 | 2955842201301948 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-06-03 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
| 2013-03-21 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the patient's external iliac vein was damaged. The patient's vessel was repaired laparoscopically and the planned surgical procedure was completed. Reportedly, the surgeon believes that the damage to the patient's vessel may have occurred as result of thermal spread from a micro crack in the monopolar curved scissors instrument. The site discarded the instrument. | On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the isi clinical sales representative (csr) who initially reported this incident. The csr indicated that he was not present during the surgical procedure when the injury occurred to the patient. The csr indicated that he was told by the surgeon, that he noted that while working on another area of the patient's anatomy, the surgeon noted that the patient was bleeding from the iliac vein. The surgeon clamped the patient's vessel using a bulldog clamp and repaired the affected area with the application of sutures using an endowrist needle driver instrument. Examination of the patient's vessel by the surgeon determined the patient had sustained a thermal injury. The csr indicated that the surgeon initially believed that the injury to the patient was unrelated to a malfunction of the da vinci surgical system, instruments and accessories, but rather, he (surgeon) had inadvertently caused damage to the patient's vessel while performing lymph node dissection or while opening the patient's pelvic side wall to access the uterine artery. Upon receipt of the isi's product notification letter regarding the monopolar curved scissors (mcs) instrument, the surgeon notified the csr and indicated that the damage to the patient's vessel may have occurred as a result of cautery energy coming from a possible crack in the mcs instrument. The csr indicated that during the surgical procedure, arcing from the mcs instrument with the tip cover accessory installed was not observed. The csr does not know if the site performed an inspection of the mcs instrument or tip cover prior to use. An inspection of msc instrument and tip cover after use was not performed by the surgical staff, since it was initially believed that the injury to the patient occurred as a result of user error. The csr indicated that the planned surgical procedure was completed and to the best of his knowledge the patient recovered and there was no report to him by the site that the patient experienced any post-surgical complications as a result of the reported event. At this time, isi is unable to determine the root cause of the injury sustained by the patient as the instrument is not available for analysis. Isi has contacted the surgeon several times to obtain additional information regarding the reported event; however, as of the date of this report, no additional information has been provided. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3169802 | 3141065 | 3037620 | 2955842201301949 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-03-21 | 2013-06-03 | Yes | Yes | Health Professional |  | No | No | 2013-03-21 | No | No | Unkown | 2013
42017207 | 2013-05-03 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the (b)(4) bipolar forceps instrument wires were noted to be coming apart at the top. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the conductor wires were intact and undamaged, but the drive cable was frayed. The pitch down cable was frayed at the distal clevis hub. The frayed strands stuck out at the wrist. Engineering also found a derailed cable. The one grip close cable was derailed from the distal idler pulley. The grips still opened and closed, but movement may not be precise. The evidence was inconclusive, but the cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTR | ENDOSCOPIC ELECTROSURGICAL INS | 3169834 | 3141080 | 3037635 | 2955842201301950 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-05-29 | 2013-06-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-29 | Yes | No | Reuse | 2013
42011704 | 2013-05-03 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical instrument was noted to have a broken tip. No patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3169821 | 3141091 | 3037610 | 2955842201301951 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-06-03 | No | Yes | Health Professional | 42011704 | No | No | 2013-05-03 | No | No | Unkown | 2013
42019408 | 2013-05-09 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken cable on the mega needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed there was a broken grip cable that was sticking out at the distal idlers. The idler pulley spun freely but exhibited some wear. Other cables at the instrument's wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3169959 | 3141181 | 3037760 | 2955842201301952 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-03 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42018304 | 2013-05-09 | Malfunction |  | Prior to starting a da vinci s surgical procedure, the user facility reportedly identified a broken cable on the permanent cautery hook instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering did not confirm the allege broken cable complaint. All cables on the instrument were all still tensioned. However, engineering found a broken proximal clevis at the main tube interface. The clevis was dislodged from main tube and both the proximal clevis and main tube were missing pieces. Electrical continuity testing was performed and passed. No other damage found. The observed breakage is likely due to overloading at the tip or other type of mishandling. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3171326 | 3142824 | 3039359 | 2955842201301953 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-03 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-15 | Malfunction |  | It was reported that during a da vinci si surgical procedure the maryland bipolar forceps instrument was not working at all. The jaws were noted to be misaligned. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of the grips. There was a. 051 offset at the tips, indicating overloading at the tip. The instrument passed electrical continuity testing. There were scratches on the surface of the distal pulley. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3171614 | 3143044 | 3039579 | 2955842201301954 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-06 | 2013-06-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-24 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the maryland bipolar forceps instrument was not recognized by the robot. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of grips. There was a. 070 offset at the tips, indicating overloading at the tip. The instrument passed electrical continuity testing. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3171977 | 3143420 | 3039955 | 2955842201301956 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-06-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure the pk dissecting forceps instrument tips were noted to not approximate. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of the grips. There was a. 160 offset at the tips, indicating overloading at the tip. The instrument passed electrical continuity testing. Engineering also found scratch marks showing light material removal by the proximal clevis side. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3171976 | 3143435 | 3039974 | 2955842201301955 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-06-03 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42017105 | 2013-05-09 | Malfunction |  | During a da vinci si esophageal diverticulum repair procedure, the user facility reportedly identified the grips were not staying closed on the micro bipolar forceps instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. The customer initially reported that the grips will not stay closed. However for clarification, the nonconformance was non-intuitive motion. Based on this, the complaint was confirmed. Intuitive motion was not being experienced due to a broken pitch cable at the instrument's wrist. The cable segment that contains the crimp was missing. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MICRO BIPOLAR FORCEPS INSTRUME | ENDOSCOPIC ELECTROSURGICAL INS | 3172013 | 3143475 | 3040014 | 2955842201301957 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-15 | 2013-06-03 | No | Yes | Health Professional | 42017105 | Device Returned To Manufacturer | No | 2013-05-15 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-05-14 | Malfunction |  | During a da vinci si hysterectomy procedure, the surgeon reportedly noted that the prograsp forceps instrument was not responding. The instrument was removed from the patient and the user facility allegedly found broken cables on the instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed that the pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp is still installed in clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3172067 | 3143512 | 3040047 | 2955842201301959 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-16 | 2013-06-03 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-05-16 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-31 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the fenestrated bipolar forceps instrument grip was noted to not close properly. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated Per the customer reported complaint engineering evaluation found that the one grip was bent causing side to side misalignment of the grips There was a 148 offset at the tips indicating overloading at the tip The instrument electrical continuity passed Engineering also found the pitch cable was frayed at the distal idler No damage was found on the clevis There were scratches on the surface of the distal pulley The distal end of the main tube had various scratch marks showing light material removal The scratches were short in length and not axially aligned with the tube The evidence was inconclusive but the damage was likely due to mishandling The instruments and accessories user manual specifically states general precautions and warnings handle instruments with care Avoid mechanical shock or stress that can cause damage to the instruments Do not use an instrument to clean debris from another instrument intraoperatively This may result in damage to the instruments or other unintended consequences such as disconnection of the instrument tip The customer reported complaint does not itself constitute a mdr reportable event however the damage to the instruments main tube which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3172070 | 3143514 | 3040049 | 2955842201301958 | 1 | NAY | Manufacturer | DistributorHealth Professional | Nurse | Initial | 2013-05-06 | 2013-06-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-05-08 | Malfunction |  | During a da vinci si hysterectomy procedure, the mega needle driver reportedly was not moving intuitively with the surgeon's controls. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering did not confirm the alleged complaint. The instrument was placed on an in house da vinci system: recognition and engagement tests passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Additional observations not reported by the customer were main tube and pulley damages. The distal end of the main tube had a scratch mark showing light material removal. The scratch was short in length and not axially aligned with the main tube. There was also an indentation at the edge of the distal pulley and scratches on the surface of the pulley. Evidence not conclusive, but main tube and pulley damage may be due to mishandling. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the main tube scratches if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3172089 | 3143551 | 3040090 | 2955842201301960 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-16 | 2013-06-03 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-05-16 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-05-08 | Malfunction |  | During a da vinci si surgical procedure, the surgeon reportedly identified frayed wires on the grasping retractor instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering did not confirm the alleged frayed wires complaint. For clarification, the pitch cable was found broken at the distal end and was not frayed. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3172104 | 3143566 | 3040101 | 2955842201301962 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-20 | 2013-06-03 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-05-20 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-23 | Malfunction |  | It was reported that during a da vinci si gastric bypass procedure the monopolar curved scissors instrument was noted to be dull. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was placed on system for latex cut test and the scissors did not cut cleanly through. 006 latex. The latex was snagged at the scissor tips. The blade edges were undamaged, but edges exhibited wear at the tips, negatively affecting cut performance. The instrument passed electrical continuity testing. Deep scratches near the distal tip of the instrument approximately 0. 8 from the base of the tube extension. Additional observation that was not reported by the customer was micro-cracks found at the distal end of the tube. The micro-cracks run in the axial direction. These types of micro-cracks will not lead to mechanical failure of the instrument, however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of these is confined to 2cm of the distal end of the instrument shaft. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3172153 | 3143576 | 3040111 | 2955842201301965 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-06 | 2013-06-03 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-05-06 | Malfunction |  | It was reported that during reprocessing the da vinci si fenestrated bipolar forceps instrument was noted to have a broken tip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of the grips. There was a. 106 offset at the tips, indicating overloading at the tip. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. The distal end of main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3172137 | 3143579 | 3040114 | 2955842201301966 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-06-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-05-03 | Malfunction |  | It was reported that during a da vinci s total hysterectomy procedure, a cable broke at the jaw of a tenaculum forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3172122 | 3143585 | 3040124 | 2955842201301961 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-07 | 2013-06-03 | No | Yes | Health Professional | 42020706 | No | No | 2013-05-07 | No | No | Reuse | 2013
42022703 | 2013-05-07 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the pk dissecting forceps instrument was not being recognized by the da vinci robot system. Instrument was not used in the case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated Engineering was unable to confirm the reported failure mode The instrument was placed on inhouse da vinci robot system and the system successfully recognized the instrument The instrument passed electrical continuity tests Additional observations not reported by site were bent tips distal pulley damage and main tube damage First finding one grip was found to be bent causing side to side misalignment of grips There was a 0 088 offset at the tips indicating overloading at the tip Second finding scratches on the surface of the distal pulley Third finding the distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish The scratches were short in length and were not aligned with the tube axis Engineering concluded that damages were likely due to mishandling The endowrist instruments instructions for use ifu specifically states general precautions and warnings handle instruments with care Avoid mechanical shock or stress that can cause damage to the instruments The customer reported complaint does not itself constitute a reportable event however the reported malfunction of scratches on the main tube if to reoccur could cause or contribute to an adverse event | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3172127 | 3143589 | 3040128 | 2955842201301963 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-06-03 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-02-05 | Malfunction |  | It was reported that during a da vinci si surgical procedure the small grasping retractor instrument was noted to have restricted endowrist movement. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. Both cable crimps were missing. The clevis did not exhibit any damage or wear marks. There were scratches on the surface of the distal pulleys. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3172110 | 3143594 | 3040133 | 2955842201301964 | 1 | NAY | Manufacturer | DistributorHealth Professional | Nurse | Initial | 2013-05-06 | 2013-06-03 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | Yes | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-18 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the maryland bipolar forceps instrument wires broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one conductor wire was broken at the yaw pulley exit. The wire was detached from the connection at the grip. The instrument failed electrical continuity testing. The yaw pulley showed no signs of arcing. There were scratches on the surface of the distal pulley. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3172245 | 3143676 | 3040211 | 2955842201301967 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-06-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-05-02 | Malfunction |  | It was reported that during a da vinci si gastric bypass procedure the large suturecut needle driver instrument wire snapped when the jaws were opened. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken near the proximal pulleys and proximal clevis cable hole. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the wrist. No major wear was found on proximal clevis cable hole. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3172292 | 3143722 | 3040257 | 2955842201301968 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-06-03 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-05-06 | Malfunction |  | It was reported that prior to starting a da vinci s procedure the small clip applier instrument tip was noted to be broken off upon removal from peel pack. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip broken off at the base of the clip groove. The missing grip piece was roughly. 190 x. 045. The intact grip was bent just below the midpoint of the clip groove. The evidence was inconclusive, but the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3172338 | 3143762 | 3040297 | 2955842201301969 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-06 | 2013-06-03 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | Yes | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-05-08 | Malfunction |  | It was reported that prior to starting a da vinci si hysterectomy procedure, frayed wires were noticed at the distal end of a prograsp forceps instrument. The instrument was not used. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. A frayed pitch cable was found at the instrument's wrist. The frayed segments varied in length. No damage or wear marks were found at the clevis. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3172378 | 3143809 | 3040344 | 2955842201301970 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-06-03 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the jaws of a prograsp forceps instrument would not close. The instrument was not used in the procedure. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported the instrument was placed on a system and driven. Recognition and engagement test passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. The instrument was fully functional. Additional findings not reported were various scratch marks showing light material at the distal end of the main tube. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3172401 | 3143845 | 3040380 | 2955842201301971 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-06-03 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
410322 | 2013-05-30 | Malfunction |  | The vessel sealer coagulated but did not fully cut the tissue as expected. Another vessel sealer was opened and used with the same result. The surgeon involved uses these devices on a regular basis without incident. Reason for use: robotic prostatectomy. |  | ENDOWRIST ONE VESSEL SEALER | VESSEL SEALER | 3187438 | 3151763 | 3048298 | MW5030449 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-06-03 | 2013-06-03 | No | Yes | Health Professional | 410322 | Yes | Yes |  |  |  |  | 2013
| 2013-05-21 | Malfunction |  | During the robotic case, the harmonic insert jaw tip broke off. The piece was retrieved with no apparent injury to the patient. Device #1is this a laboratory device or laboratory test?no. |  | HARMONIC ACE CURVED SHEARS | SYSTEM SURGICAL COMPUTER CONTR | 3201299 | 3171948 | 3068503 | 3171948 | 1 | NAY | User Facility |  |  | Initial | 2013-06-03 | 2013-06-03 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
| 2013-05-31 | Malfunction |  | While in use the blade would not retract. The device was removed from pt. No patient harm. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST | SYSTEM SURGICAL COMPUTER CONTR | 3201333 | 3172013 | 3068568 | 3172013 | 1 | NAY | User Facility |  |  | Initial | 2013-06-03 | 2013-06-03 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2013
| 2013-05-23 | Malfunction |  | Instrument placed into robotic. Noticed the wires were broken, the device was removed and replaced. All pieces intact and the equipment saved. No harm to the patient. Device #1is this a laboratory device or laboratory test?no. |  | DAVINCI ENDOWRIST | SYSTEM SURGICAL COMPUTER CONTR | 3236055 | 3207155 | 3103710 | 3207155 | 1 | NAY | User Facility |  |  | Initial | 2013-06-03 | 2013-06-03 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
42030901 | 2013-05-06 | Malfunction |  | It was reported that during reprocessing a mega suturecut needle driver instrument, wires were coming out of the instrument at the distal end. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. There were also scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3167519 | 3139235 | 3035770 | 2955842201301932 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-31 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42011704 | 2013-05-06 | Malfunction |  | It was reported that during reprocessing a needle driver instrument, insulation was coming off the instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. Black tube insulation was damaged at the distal end. Insulation was gouged in various places and had pieces of insulation lifted up. The metal shaft was exposed as a result. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3167520 | 3139236 | 3035771 | 2955842201301933 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-31 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
41032204 | 2013-05-02 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, frayed wires at the distal end of a vessel sealer instrument were noted. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with a broken cable at the instruments snake wrist. No other cables were broken but some bio debris was found at the snake wrist and between the grips. The instrument blade was not exposed out of the garage. The broken cable segment stuck out approximately. 1161 at the snake wrist. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | VESSEL SEALER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3167521 | 3139237 | 3035772 | 2955842201301934 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-31 | No | Yes | Health Professional | 41032204 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-26 | Malfunction |  | It was reported that at the start of a da vinci si surgical procedure, one of the wires was broken at the distal end of a prograsp forceps instrument. A new instrument of the same type was used to complete the procedure. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. There were also scratches on the surface of the distal pulley. Additional findings not reported was the distal end of the main tube had various scratch marks with light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3167523 | 3139239 | 3035774 | 2955842201301935 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-31 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-30 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the mega suturecut needle driver instrument cable snapped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. There were multiple indentations at the edge of the distal pulley and visible scratches on the surface of the pulley. The evidence was inconclusive; however, the damage was most likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3168265 | 3139749 | 3036304 | 2955842201301919 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-02 | 2013-05-31 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-05-03 | Other | Other | It was reported that during a da vinci s sigmoid colectomy procedure, the insulation on the bowel grasper instrument scrapped off. No missing or fallen pieces were reported. Reportedly, the planned surgical procedure was completed using open surgical techniques. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the main tube insulation exhibited damage from the distal end until approximately 4. 5 inches below. The insulation was found with several gouge marks within that area. The marks were short in length and not axially aligned with the tube. The main tube also exhibited a couple of different damaged sections with tube insulation removed. The main tube also appeared with discoloration marks with no material loss throughout the entire main tube. Intuitive surgical, inc. Has made several attempts to obtain additional information concerning the reported event. As of the date of this report, no additional information has been provided. A follow-up medwatch report will be submitted to the fda if additional information is received. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3168933 | 3140314 | 3036923 | 2955842201301939 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-31 | Yes | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-24 | Malfunction |  | It was reported that prior to starting a da vinci si prostatectomy procedure, when the customer removed the monopolar curved scissors (mcs) instrument out of the box, they noted that the connection pin was missing from the instrument. No patient harm, adverse outcome or injury was reported. Sincerely,. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the connector pin was missing and no banana plug was found, but there were loose screws found inside housing. There was no additional damage found on neither the proximal nor the distal end. The complaint stated that the instrument was never used; however, the instrument was returned showing 9 uses remaining. The system error logs were reviewed and they showed that the mcs instrument was installed on the da vinci system on (b)(4) 2013. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the additional finding of main tube scratches if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3168982 | 3140317 | 3036816 | 2955842201301941 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-06 | 2013-05-31 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-05-02 | Injury | Required Intervention | It was reported that post op a da vinci hysterectomy procedure performed on (b)(6) 2013, the patient returned to the hospital the following week due to experiencing pain. During a surgical examination performed on the patient it was discovered that the patient's rectum was necrotic and had possible cauterized tissue. The patient underwent a reversible colostomy procedure. Reportedly, no system issues or patient injury was observed during the surgical procedure. | The mcs instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. It is unknown if the mcs instrument has been used in any subsequent surgical procedures after the reported event occurred. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation). On (b)(6) 2013, intuitive surgical, inc. (isi) contacted the hospital to obtain additional information concerning the reported event. On (b)(6) 2013, the site provided isi with a copy of their uf importer (b)(4). The uf importer report indicated that the patient underwent a robotically assisted total hysterectomy with bilateral salpingectomy procedure on (b)(6) 2012 and was discharged from the hospital on (b)(6) 2013. On (b)(6) 2013 the patient arrived at the hospital's emergency department with complaints of abdominal pain, back pain and nausea and vomiting. A ct scan performed on the patient's abdomen was inconclusive for small bowel obstruction or ileus. The patient was then admitted into the hospital for evaluation. An exploratory laparotomy, bowel resection and diverting colostomy were performed on the patient. A thermal burn to the patient's rectum was discovered during a post-surgical diagnosis. The mcs instrument was used during the surgical procedure. As of (b)(4) 2013, no additional information regarding the reported event has been provided by the site. A follow-up mdr will be submitted if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3169033 | 3140322 | 3036935 | 2955842201301945 | 1 | NAY | Manufacturer | Company RepresentativeOther | Other | InitialFollowup | 2013-05-09 | 2013-05-31 | Yes | Yes | Health Professional | 42017910 | No | No | 2013-08-22 | No | No | Unkown | 2013
42017704 | 2013-05-03 | Other | Other | It was reported that after completion of a da vinci s sigmoid colectomy procedure, the insulation on the bowel grasper instrument was noted to have a puncture. There were no missing or fallen pieces reported. The planned surgical procedure was converted to open. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the black tube insulation was damaged at the distal end. The insulation was gouged in various places and had a. 324 x. 089 1st piece of insulation lifted up, a 2nd piece measured approximately 0. 231 x 0. 115, and the last piece measured approximately 0. 306 x 0. 092. The metal shaft was exposed as a result. Engineering concluded that the damage was likely due to mishandling. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Intuitive surgical, inc. Has made several attempts to obtain additional information concerning the reported event. As of the date of this report, no additional information has been provided. A follow-up medwatch report will be submitted to the fda if additional information is received. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3168879 | 3140388 | 3036923 | 2955842201301937 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-31 | Yes | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
420179 | 2013-05-03 | Injury | Required Intervention | It was reported that during a da vinci hysterectomy procedure, the patient side assistant inserted an eea sizer into the patient's rectum, causing the surgeon to inadvertently cut the patient's rectum while using a monopolar curved scissors instrument. Reportedly, the patient's uterine anatomy was unique. The patient's uterus was described as being disjointed. The surgeon had initially installed a vcare uterine manipulator into the patient's vagina; however, he had difficulty with the surgical planes. The surgeon had the patient side assistant remove the vcare uterine manipulator and place an eea sizer into what he (surgeon) thought was the patient's vagina. The surgeon proceeded with the colpotomy, and after approximately 5 minutes the surgeon was still dissatisfied with the surgical plane. The surgeon scrubbed back in to evaluate the patient's anatomy at the patient bedside. The surgeon discovered that the patient side assistant had inserted the eea sizer into the patient's anus instead of the patient's vagina, thus making it difficult for the surgeon to delineate the surgical planes, causing the surgeon to inadvertently cut the patient's rectum. The surgeon made the decision to complete the planned surgical procedure using open techniques and then repaired the patient's rectum with suture. | On (b)(4) 2013 intuitive surgical, inc. Contacted the (isi) clinical sales representative who initially reported this incident. The csr indicated that he was present during the surgical procedure and that no malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure. The csr indicated that the surgeon was using cautery energy when the patient's rectum was damaged. The csr indicated that during his site visit the following week he was told by the surgical staff that the patient was doing well post-operatively. Based on the information provided, it has been determined that the injury sustained by the patient was not caused by a malfunction of the da vinci surgical system, instruments or accessories, but likely occurred as a result of the surgeon's difficulty delineating the surgical planes due to the patient's difficult anatomy. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3168918 | 3140400 | 3036935 | 2955842201301938 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-05-03 | 2013-05-31 | Yes | No | Health Professional | 420179 | No | No | 2013-05-03 | No | No | Reuse | 2013
IS3000 A60P8 | 2013-05-08 | Other | Other | It was reported that during a da vinci si hysterectomy procedure on (b)(6) 2013, the surgical staff indicated that a non-recoverable error 40006 issue occurred. The surgical staff rebooted the system but the error message reportedly persisted. The patient side cart (psc), surgeon side console (ssc), and vision side cart (vcs) were reportedly powered off and the system powered up and homed as expected. A clinical sales representative (csr) monitored the site's system log remotely and identified error 40006 pointing to the left master tool manipulator (mtml). The csr also identified the following errors in the system log: 1103, 55000, 55100, 31201, 255521, 25100, 25116, and 50000. Due to the unresolved and recurring error 40006 issue, the surgeon reportedly made the decision to convert the procedure to open surgical techniques. | On (b)(4) 2013, an intuitive surgical inc. (isi) field service engineer (fse) visited the site and reportedly resolved the error 40006 issue by replacing the a remote arm controller (rac) board. The fse also replaced the dual motor drive (dmd) board and the motor connector backplane (mcb) to resolve an error 31046 issue. According to the fse, the system was tested and passed to original equipment manufacturer specifications. The dmd was returned to isi and evaluated. Engineering evaluation was able to replicate the reported error 31046 issue while testing the dmd on an in-house pca system. The rac was returned to isi and evaluated. Engineering evaluation was unable to duplicate the reported error message issue. The rac was tested on an in-house patient cart controller (pcc) and found to have passed. Engineering evaluation did not find any issues with the rac. The mcb was also returned to isi and evaluated. Engineering evaluation was unable to replicate the reported error 31046 issue by testing the mcb on an in-house printed circuit assembly (pca). On (b)(4) 2013, isi contacted the csr to obtain additional information regarding the reported event. The csr indicated that the error 40006 issue began while the incision ports were being placed and before the psc was docked to the patient. At that time, the patient had been under anesthesia for approximately 30 minutes. The csr did not know if the patient experienced any intra-operative or post-operative complications. According to the csr, the patient was doing well. On (b)(4) 2013, isi contacted the site and obtained additional information regarding the reported event. The reporter indicated that the patient was under anesthesia for about 1 hour and the psc was docked to the patient. The reporter confirmed that the error 40006 issue occurred before the da vinci si hysterectomy procedure was started. However, the reporter indicated that while he and the csr were troubleshooting the error message issue, the surgeon attempted to use the da vinci si system for about 10-15 minutes. Due to the recurring error message issue, the reporter indicated that the surgeon made the decision to convert to open surgical techniques. To the reporter's knowledge, the patient did not experience any intra-operative or post-operative complications. He also verified that the da vinci si system was successfully repaired by the fse. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3168977 | 3140434 | 3036969 | 2955842201301940 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-05-31 | Yes | Yes | Health Professional | IS3000 A60P8 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-05-03 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure that a cable broke on a fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the instrument's wrist. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3168988 | 3140440 | 3036975 | 2955842201301942 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-05-06 | 2013-05-31 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-05-06 | Malfunction |  | It was reported that during a da vinci si thymectomy procedure, the insulation on the permanent cautery spatula instrument broke off and fell into the patient. The broken instrument piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation). On (b)(4) 2013 intuitive surgical contacted the initial reporter regarding the reported event. The initial reporter indicated the she was not present during the surgical procedure when the issue occurred; however, it was reported to her by the surgical staff that the permanent cautery spatula (pcs) instrument made contact with another instrument that was also being used during the surgical procedure, causing the pcs instrument to break. The broken instrument piece was removed from the patient laparoscopically using the da vinci surgical system and there was no harm to the patient. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3168970 | 3140445 | 3036980 | 2955842201301943 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-08 | 2013-05-31 | No | Yes | Health Professional | 42018406 | No | No | 2013-05-08 | No | No | Unkown | 2013
42017910 | 2013-05-10 | Injury | Required Intervention | It was reported that one week post op a da vinci s urological procedure performed on (b)(6) 2013, it was discovered that the patient had sustained a bowel injury. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. It is unknown if the mcs instrument has been used in any subsequent surgical procedures after the reported event occurred. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, the initial reporter provided additional information concerning the reported event. The initial reporter indicated that on (b)(6) 2013 the patient presented symptoms of pain and fever. A ct scan performed showed that the patient had air around the rectum and had a recto-urethral fistula. Open laparotomy techniques were utilized to perform a diverting colostomy to repair the patient's rectum. The patient was released from the hospital on (b)(6) 2013 and is reported to be healing well. The patient's fistula resolved with a urethral catheter and the colostomy will be reversed in 1 - 2 months. The initial reporter indicated that per the surgeon the root cause of the injury sustained by the patient is unknown, since the injury to the patient was unrecognized intra-operatively during the surgical procedure. The monopolar curved scissors (mcs) instrument and mcs tip cover accessory were inspected prior to use and a visual inspection of the cannula was performed. Arcing from the monopolar curved scissors (mcs) instrument with the mcs tip cover accessory installed was not observed during the surgical procedure. The mcs instrument did not make contact with any other instrument during the surgical procedure and there is no video recording of the procedure. An excalibur plus 36 electrosurgical unit (esu) was used during the surgical procedure. The esu settings were coag 30 and cut 30. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3169006 | 3140450 | 3036985 | 2955842201301944 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-10 | 2013-05-31 | Yes | No | Health Professional | 42017910 | No | No | 2013-05-10 | No | No | Reuse | 2013
IS3000 A60P8 | 2013-05-03 | Other | Other | It was reported that during a da vinci si small bowel resection procedure, the surgical staff claimed that they began to have instrument recognition issues with arm 1 and arm 3. According to the initial reporter, a suction irrigator instrument was not recognized by arm 3. In addition, the site allegedly had instrument engagement issues with arm 1. The initial reporter indicated that the patient was under anesthesia for approximately 2 hours. The patient side cart (psc) was reportedly undocked from the patient when the surgical staff contacted intuitive surgical inc. (isi) to report the instrument recognition issues. At the time of contact with isi, the surgical staff reportedly did not have time to troubleshoot the reported issues. The technical support engineer (tse) instructed the site to call isi back when available to troubleshoot. The tse reviewed the site's system logs and identified errors: 22008, 31009, 282, and 23020. According to the tse, error 22008 referenced the patient side manipulator (psm) 1 cannula mount. Error 31009 reportedly referenced the tool sensor missing on psm 1. Error 282 reportedly referenced a tool rejected on psm 3. Error 23020 reportedly referenced the cannula port clutch switch on psm 1. | On (b)(4) 2013, an isi field support engineer (fse) reportedly evaluated the reported issues at the site. The fse indicated that the site had completed 2 da vinci si surgical procedures since the reported event with no reported issues. In addition, the fse indicated that the suction irrigator instrument involved in the initial complaint had also been used since the reported event with no further issues reported. On (b)(4) 2013, isi again contacted the initial reporter of this complaint to obtain additional information regarding the reported event. The reporter indicated that the alleged instrument recognition issues began 2 hours after the da vinci si small bowel resection procedure began. According to the reporter, there were several factors as to why the surgeon made the decision to convert to open surgical techniques. The reporter indicated that one reason why the surgeon made the decision to convert to open surgical techniques was due to experiencing the instrument recognition issues. The reporter also indicated that the surgeon was already planning on converting to open surgical techniques since he needed to use a stapler instrument to complete the small bowel resection procedure. The reporter indicated that the patient tolerated the open surgical procedure well. To her knowledge, the patient did not experience any intra-operative or post-operative complications. The initial reporter indicated that the alleged instrument recognition issues were resolved after an fse replaced arm 1. She stated that the site had completed two subsequent da vinci si surgical procedures since the reported event with no reported issues. She also indicated that the suction irrigator instrument involved with this complaint had also been used subsequently with no reported issues. On (b)(4) 2013, isi contacted the fse to obtain additional information regarding the reported event. The fse stated that he did not replace any products or arm 1 as reported by the initial reporter. The fse indicated that he had contacted the site to follow-up on the reported issue and was informed that the alleged issue was resolved. The fse claimed that the site did not inform him of what actions were taken to resolve the reported issue. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3169233 | 3140580 | 3037115 | 2955842201301936 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-03 | 2013-05-31 | Yes | Yes | Health Professional | IS3000 A60P8 | No | No | 2013-05-03 | No | No | Reuse | 2013
42009308 | 2013-03-15 | Malfunction |  | It was reported that during a da vinci si pyeloplasty procedure the prograsp forceps instrument was noted to be frozen. The instrument was removed. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was derailed at the distal idler pulley. The grips still opened and closed, but movement may not be precise. The cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. The fleet angle between the proximal and the distal pulleys may have contributed to the derailment. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3169458 | 3140747 | 3037282 | 2955842201301946 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-02 | 2013-05-31 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-23 | Malfunction |  | It was reported that during a da vinci si thymectomy procedure, the surgeon could not grasp tissue safely and found the jaw of the instrument could not be closed completely. The instrument was pulled out and checked. There was no visible damage. No fragment fell into patient. No patient injury. The procedure was completed. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated Engineering confirmed the reported complaint One grip was bent causing a side to side misalignment of the grips There was a 122 offset at the instruments tips indicating overloading Electrical continuity test passed An additional finding was the distal end of the main tube had various scratch marks showing light material removal and a rough surface finish The scratches were short in length and not axially aligned with the tube Engineering concluded the damage may be due to mishandling The endowrist instruments instructions for use ifu specifically states general precautions and warnings handle instruments with care Avoid mechanical shock or stress that can cause damage to the instruments Do not use an instrument to clean debris from another instrument intraoperatively This may result in damage to the instruments or other unintended consequences such as disconnection of the instrument tip The customer reported complaint does not itself constitute a mdr reportable event however tube abrasions with material removal found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3164864 | 3136261 | 3032800 | 2955842201301900 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-05-02 | 2013-05-30 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-30 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure that one of the tungston cables broke on a fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch down cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3164865 | 3136262 | 3005896 | 2955842201301901 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-02 | 2013-05-30 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-28 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, wires broke at the distal end of a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the surface. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164869 | 3136266 | 3032805 | 2955842201301902 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-30 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-19 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted a broken wire on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3165698 | 3137682 | 3034221 | 2955842201301903 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-06-04 | 2013-05-30 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-06-04 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-05-01 | Malfunction |  | It was reported that upon removing the da vinci si thoracic grasper instrument from the box the instrument was noted to have a nick on the outer black sheath. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the black tube insulation was damaged at the distal end. The insulation was gouged on the one side and had a. 100 x. 150 piece missing roughly 1. 75 above the snake wrist. The main tube exhibited a slight bend near the midpoint. The distal tip wobbled when the tube was rotated through roll axis range of motion. It was unclear if the instrument was shipped to the customer in this condition. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 3166051 | 3137940 | 3034475 | 2955842201301904 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-01 | 2013-05-30 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-28 | Malfunction |  | A monopolar curved scissors instrument was returned without reporting a complaint. The instrument was never used. | The instrument was returned and evaluated. No failure was reported for this rma. Engineering found three hairline cracks at the distal end of the main tube. Electrical continuity test passed. The instrument was never used. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. No other damage was found. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3166406 | 3138215 | 3034770 | 2955842201301905 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-02 | 2013-05-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
41032204 | 2013-04-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted broken wires at wrist of the vessel sealing instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cables were found to be broken and stuck out at the wrist. The instrument knife blade was also exposed. No additional damage was found on snake wrist and no other cable damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction of broken cable if to reoccur could cause or contribute to an adverse event. | VESSEL SEALING INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3166428 | 3138248 | 3034783 | 2955842201301906 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-30 | No | Yes | Health Professional | 41032204 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Initial | 2013
42000606 | 2013-04-17 | Malfunction |  | It was reported that during a da vinci s surgical procedure the large needle driver instrument stopped responding to the system during suturing. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. There were indentations on the edge of the distal pulley. The evidence was inconclusive; however, the damage was most likely due to mishandling. There are also scratches on the surface of the pulley. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3166430 | 3138249 | 3034784 | 2955842201301907 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-02 | 2013-05-30 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-05-03 | Malfunction |  | It was reported that during reprocessing, the tips of a thoracic grasper instrument were not opening because of a possible broken wire. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. Visual inspection shows no broken or loose cables at the distal end. The instrument was placed on a is3000 system and driven. Recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Functional performance testing did not find any issues. Additional observation not reported by site was tube damage. Black tube insulation was damaged at the distal end. Insulation was gouged on one side and had a. 050 diameter piece missing roughly. 940 above the instruments snake wrist. The distal end of the tube also exhibited various scratch marks with no material removal. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; gouged insulation missing , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3166518 | 3138271 | 3034810 | 2955842201301909 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-30 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
40011010 | 2013-04-16 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the precise bipolar forceps instrument jaws were not aligned. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of the grips. There was a. 060 offset at the tips. The bent grip had separation of the yaw pulley and the pulley cover at the glue joint, indicating overloading at the tip. Engineering also found tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 140 -. 210 in length and not aligned with the tube axis. The evidence was inconclusive, but damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3166520 | 3138290 | 3034825 | 2955842201301908 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-02 | 2013-05-30 | No | Yes | Health Professional | 40011010 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-05-03 | Malfunction |  | It was reported that during reprocessing a thoracic grasper instrument, a broken wire near the tip was noted. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. A couple of drive cables were broken at the distal end of the snake wrist. The broken segments stuck out of the instruments wrist. The snake wrist could not properly straighten and motion was not intuitive due to the cable breakage. Additional observation not reported by the site was tube damage. Black tube insulation was damaged at the distal end. The insulation was gouged on one side and had a. 055 x. 045 piece missing roughly. 700 above the snake wrist. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3166534 | 3138297 | 3034810 | 2955842201301910 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-30 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-05-03 | Malfunction |  | It was reported that during reprocessing a thoracic grasper instrument, a wire was broken near the tip. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One drive cable was broken and a second cable was loose at the distal end of the snake wrist. The cable segments stuck out of the instruments wrist. The snake wrist could not properly straighten and motion was not intuitive due to the cable breakage. Additional observation not reported by site was tube damage. Black tube insulation was damaged at the distal end. Insulation was gouged on one side and had a. 075 x. 040 piece missing roughly 1. 31 above the snake wrist. The distal end of the tube also exhibited various scratch marks without material removal. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3166556 | 3138302 | 3034810 | 2955842201301911 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-30 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-19 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the monopolar curved scissors (mcs) instrument was noted to have broken wires. The instrument was replaced. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 04 in length. Engineering also found the instrument banana plug bent at the back end of the housing. The evidence was inconclusive, but the damage was likely due to mishandling. Engineering found a gap at the connection joint between the tube extension and the main tube. There was a deep scratch at the distal end of the main tube where it connects with the tube extension. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3166560 | 3138325 | 3034860 | 2955842201301912 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-02 | 2013-05-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-24 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure, the monopolar curved scissors (mcs) instrument was noted to be dull and there was a notch on the scissors. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was found with severe blade damage on the right blade, with corresponding and less severe damage on the left blade, in the form of indentations onto the blade. The evidence was inconclusive, but damage was likely due to mishandling. Engineering also found multiple hairline cracks on the distal end of the main tube. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event, in a fragment-causing failure (such as a cable failure, a crimp separating from a cable etc. ) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument; however, the cracks in the main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3166751 | 3138483 | 3035042 | 2955842201301913 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-02 | 2013-05-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
40027202 | 2013-04-08 | Injury | Required Intervention | On (b)(4) 2013, intuitive surgical inc. (isi) received uf/importer report (b)(4) with the following event description: during a laparoscopic myomectomy, the robotic harmonic scalpel metal tip broke. X-ray was taken and revealed no retained foreign body. | On (b)(4) 2013, isi contacted the site and obtained additional information regarding the reported event. The reporter indicated that an unspecified fragment from the instrument fell into the patient. According to the reporter, the fragment was retrieved; however, he was unable to provide details on how the fragment was retrieved. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. Based on the information provided, this complaint is being reported due to the following conclusion: a fragment from an instrument allegedly broke off and fell into a patient. In addition, although the fragment was retrieved, the patient was reportedly exposed to an x-ray as an attempt to search for retained fragments. No retained fragments were reportedly found with the x-ray. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3166942 | 3138663 | 3035198 | 2955842201301914 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-30 | 2013-05-30 | Yes | Yes | Health Professional | 40027202 | No | No | 2013-04-30 | No | Yes | Unkown | 2013
42017910 | 2013-04-17 | Malfunction |  | During a da vinci si prostatectomy procedure, the user facility reportedly identified a defective cable on the monopolar curved scissors instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a broken grip close cable at the distal idlers pulley. The broken cable segment was sticking out at wrist. The idler pulley spun freely and was not damaged. Other cables at the instrument's wrist were not damaged. Additional observations found not reported by the customer were multiple micro-cracks on the distal end of the main tube. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The customer reported complaint does not itself constitute a reportable event; however, the reported broken cable and main tube damage if to reoccur may cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3166955 | 3138680 | 3035239 | 2955842201301916 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-10 | 2013-05-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-05-02 | Malfunction |  | During a da vinci si esophageal diverticulum repair procedure, the user facility reportedly identified the orange portion of the monopolar curved scissors instrument's shaft was cracked. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed that the tube extension (orange portion of shaft) was broken and missing a piece at the distal end of the tube. The clevis was not dislodged from tube extension but the tube extension was found bent away and also pushed into the clevis. Electrical continuity testing was performed and passed. Additional observation that was not reported by the customer were micro-cracks found at the distal end of the tube. The micro-cracks run in the axial direction and measures approximately 0. 277. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. The tube extension likely broke due to excessive side loading or other type of mishandling. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a mdr reportable event; however; the cracks in the instrument's main tube, found during failure analysis evaluation may cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3166948 | 3138688 | 3035223 | 2955842201301915 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2013-05-09 | 2013-05-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-25 | Malfunction |  | The monopolar curved scissors instrument reportedly became broken during a da vinci si surgical procedure. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering found a damaged tube extension, broken cable, and micro-cracks on the main tube. The tube extension was found broken and was also missing a 0. 160 x 0. 200 piece at the proximal clevis interface. The clevis was dislodged from tube extension. The tube extension was fractured next to one of the keys that mates with the proximal clevis. One grip close cable was broken within the main tube near the tube extension breakage location. Other cables at the instrument's wrist were not damaged. Engineering was unable to conclude if the tube extension breakage contributed to cable breakage. Any potential fragments from the damaged tube extension or broken cable would likely be retained by use of tip cover. The clinical risk evaluation performed on the health hazard assessment for the monopolar curve scissors states: in a fragment-causing failure (such as a cable failure, a crimp separating from a cable, etc) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument. The distal end of main tube exhibited a couple of micro-cracks in the axial direction. Cracks were under 0. 400 in length. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. No other damage was found. Evidence not conclusive, but damages to the tube extension may be due to mishandling. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the micro-cracks found on the main tube found during failure analysis may cause or contribute to an adverse event, if they were to reoccur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3166981 | 3138699 | 3021040 | 2955842201301917 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-14 | 2013-05-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-14 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-05-09 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a cable sticking out at the distal end of the tenaculum forceps instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed there was a frayed grip cable at the distal clevis of the instrument. The idler pulley spun freely and was not damaged. The cable strand was sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. Additional observations that were not initially reported by the customer were main tube damage. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. The scratches were short in length and were not axially aligned with the main tube. No other damage found. Evidence not conclusive, but damages to the main tube may be due to mishandling. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3166958 | 3138705 | 3035260 | 2955842201301918 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-05-30 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of the system would not read the instrument. The instrument was checked for recognition on an in-house system and the instrument was successfully recognized. The pogo pins were not stuck or contaminated. The dcp (dallas chip programmer) verification of the instrument passed. The instrument was driven and moved intuitively and the scissors opened and closed. Engineering also found pad printing removed on the tube extension. Multiple scratches were found on the tube surface. The evidence was inconclusive, but damage was likely due to improper cleaning. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3167008 | 3138723 | 3035278 | 2955842201301921 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-05-03 | Malfunction |  | It was reported that during reprocessing the da vinci si mega needle driver instrument was noted to have a frayed cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was frayed at the grip hub. The frayed strands stuck out slightly above the hub. The other cables at the wrist were undamaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3166970 | 3138741 | 3035340 | 2955842201301925 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-30 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
42001103 | 2013-04-29 | Malfunction |  | During a da vinci si prostatectomy procedure, the user facility reportedly identified a bent cannula. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument accessory involved with the complaint. Engineering confirmed that the returned cannula was bent at the distal end, which was deforming the oval shape opening. The 5mm obturator would not fit properly as a result of the deformed opening. No other damage found. Evidence not conclusive but damage is likely due to mishandling. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | 5MM CANNULA | ENDOSCOPIC INSTRUMENT ACCESSOR | 3167009 | 3138744 | 3035279 | 2955842201301920 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-10 | 2013-05-30 | No | Yes | Health Professional | 42001103 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-05-08 | Malfunction |  | During a da vinci si myomectomy procedure, the user facility reportedly identified a broken cable on the tenaculum forceps instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed there was a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. Clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3167031 | 3138767 | 3035302 | 2955842201301922 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-10 | 2013-05-30 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-05-02 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the maryland bipolar forceps instrument was not articulating and when the instrument was removed it was noted that the cable was broken. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the conductor wire was broken. The drive cables at the wrist were intact; however, one conductor wire was broken at the yaw pulley exit. The wire was detached from its connection at the grip. The instrument failed electrical continuity testing. The yaw pulley showed no signs of arcing. Engineering also found tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 065 -. 125 in length and not aligned with the tube axis. The evidence was inconclusive, but damage may have been caused by mishandling/misuse. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3166967 | 3138769 | 3035304 | 2955842201301923 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-02 | 2013-05-30 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-05-01 | Malfunction |  | A pk dissecting forceps instrument was returned without a return material authorization. No complaint was communicated to intuitive surgical regarding performance of the da vinci surgical system. No malfunction of the da vinci system was alleged. | No reason for return was provided. A complaint cannot be confirmed or denied. The instrument was returned and evaluated. Engineering observed a frayed cable and tube damage. The pitch up cable was frayed at the distal clevis hub. The frayed strands stuck out at the instruments wrist. Other cables at the wrist were not damaged. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 150 -. 180 in length and not aligned with the tube axis. Engineering concluded the damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer returned instrument does not itself constitute a mdr reportable event; however; the frayed cable and or tube abrasions with material removal, if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3167068 | 3138794 | 3035329 | 2955842201301924 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-01 | 2013-05-30 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-28 | Malfunction |  | A fenestrated bipolar forceps instrument was returned. No complaint was communicated to intuitive surgical regarding the da vinci surgical system or instrument. No malfunction of the da vinci system was alleged. | The instrument was returned and evaluated. Engineering found broken snaps inside the housing. One of four housing snap tabs and all three adjacent housing pins were broken. The housing could be removed as a result of the damage. Instrument may have been mishandled. Additionally, the luer plate was missing. Due to the broken snaps and the housing was able to move, it may have resulted in the missing luer plate. The distal end of the main tube also had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The engineering findings does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3167081 | 3138817 | 3035372 | 2955842201301927 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-30 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42805312 | 2013-03-28 | Malfunction |  | A medium-large clip applier instrument was returned. No complaint was communicated to intuitive surgical regarding the da vinci surgical system or instrument. No malfunction of the da vinci system was alleged. | The instrument was returned and evaluated. Engineering found the main tube broken at the distal end. The tube fractured at the interface with the clevis. The grips could not open/close properly due to the tube breakage. Engineering concluded the damage was likely due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MEDIUMLARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3167084 | 3138856 | 3035391 | 2955842201301928 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-30 | No | Yes | Health Professional | 42805312 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-28 | Malfunction |  | A prograsp forceps instrument was returned. No complaint was communicated to intuitive surgical regarding the da vinci surgical system or instrument. No malfunction of the da vinci system was alleged. | The instrument was returned and evaluated. Engineering found the instrument had deep scratches on the main tube and a derailed grip cable. The distal end of the main tube had several scratch marks exhibiting light material removal and a rough surface finish. The scratches were short in length ranging between. 05 to. 09. Engineering concluded the damage may be due to mishandling. One grip cable was derailed off the clamping pulley at the proximal end. The grips can still open and close, but movement may not be precise. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The engineering findings does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3167132 | 3138866 | 2978146 | 2955842201301926 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-30 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-03 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the monopolar curved scissors instrument couldn't come out of the cannula and got stuck. Then arm was undocked and the instrument was taken out. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The tube extension was broken and missing a. 280 x. 160 piece at the proximal clevis interface. The clevis was dislodged from the tube extension. Breakage/dislodgment likely contributed to the instrument getting stuck in the cannula. The tube extension fractured next to one of the keys that mates with the proximal clevis. Engineering concluded the tube extension likely broke due to excessive side loading or other mishandling/misuse. Additional observation not reported by the site was tube damage. The distal end of the main tube had a 1 long section with material removed on one side of the tube. The damaged area was parallel to the tube axis and had a rough surface finish. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3167147 | 3138884 | 3035419 | 2955842201301929 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-06 | 2013-05-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-03 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure a prograsp forceps instrument stopped working. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch down cable was broken at the distal clevis hub. The cable segment that contains the crimp was missing. The clevis did not exhibit any damage or wear marks. Other cables at the instruments wrist were not damaged. Additional observation not reported by the site was tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 060 -. 110 in length and not aligned with the tube axis. Engineering concluded the damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3167175 | 3138890 | 3035425 | 2955842201301930 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-06 | 2013-05-30 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-05-06 | Malfunction |  | It was reported that after reprocessing a fenestrated bipolar forceps instrument that one of the pitch cables was broken. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instruments wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3167195 | 3138925 | 3035460 | 2955842201301931 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-30 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-11 | Malfunction |  | It was reported that during a da vinci si heller myotomy procedure, the thread came out of the wrist of a prograsp forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The pitch down cable was frayed at the distal clevis hub. An additional observation not reported was tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 225 in length and not aligned with the tube axis. Engineering concluded the damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3161033 | 3131863 | 3028418 | 2955842201301871 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-29 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, a wire broke on a large needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely. The cable segment stuck out at the instrument's wrist. The other cables at the wrist were not damaged. An additional observation not reported by the site was a damaged pulley. The distal idler pulley on which the broken cable was seated had two sections that were bent inwards. Engineering concluded the pulley damage was likely due to mishandling and that damage likely contributed to the cable breakage. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3161082 | 3131915 | 3028450 | 2955842201301872 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-05-03 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken shaft on the mega needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed that the main tube (shaft) of the instrument was broken at the distal end, roughly 2. 5 above the proximal clevis. The main tube was bent as a result of the main tube breakage and will not fit smoothly through a cannula. Engineering concluded that the main tube damage may be due to mishandling. No other damage was found. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3161812 | 3132654 | 3029209 | 2955842201301859 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-29 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-05-07 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken cable that was sticking out of the pk dissecting forceps instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a broken pitch up cable at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3161939 | 3132791 | 3029326 | 2955842201301860 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-05-29 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-05-08 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified the orange tape peeling back by the tip of the monopolar curved scissors instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed that the tube extension (orange tape) was broken and missing a piece at the distal end of the instrument. The clevis was not dislodged from tube extension. The tube extension likely broke due to excessive side loading or other type of mishandling. Electrical continuity testing was performed and passed. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3162007 | 3132866 | 3029401 | 2955842201301858 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-05-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-05-09 | Malfunction |  | It was reported that the user facility identified a broken cable on the instrument after it was cleaned and sterilized. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a broken conductor wire. The drive cables were all intact at the instrument's wrist but the conductor wire was broken at the yaw pulley exit. The wire was sticking out from the yaw pulley and conductor cap. No char marks were observed on the distal end components. Electrical continuity testing was performed and failed. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3162091 | 3132965 | 3029504 | 2955842201301861 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-05-29 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the large needle driver instrument tips were noted to not rotate. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument grip tips were bent, causing an offset and limited movement. The one grip was bent, causing side-to-side misalignment of the grips. There was a. 022 offset at the tips. The evidence was inconclusive, but damage was likely due to mishandling. The grip cable crimp was dislodged from the distal clevis and was hanging on the cable. Engineering also found the distal idler pulley damaged, a deep scratch, and a crack on the housing. The distal idler pulley had a straight edge on the one side due to an indentation. The evidence was inconclusive, but damage was likely due to mishandling. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. A small crack was found on the housing near release levers. The evidence was inconclusive, but damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3162250 | 3133140 | 3029675 | 2955842201301862 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-07 | Malfunction |  | On (b)(4) 2013, isi received (b)(4) with the following event description: pk dissecting forceps not working properly per surgeon. Instrument replaced. Upon inspection of instrument, there was a frayed segment on the instrument. No harm to patient. No x-ray required, as item was intact. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3162503 | 3133396 | 3029931 | 2955842201301857 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2013-01-24 | 2013-05-29 | No | Yes | Health Professional | 42022703 | No | No | 2013-01-24 | Device Not Retu | No | Unkown | 2013
42030901 | 2013-05-08 | Malfunction |  | During a da vinci s hysterectomy procedure, the cable on the mega suturecut needle driver instrument reportedly broke. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a broken grip close cable at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment was sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. Additional observations that were not initially reported by the customer were scratches on the surface of the distal pulley. Engineering was unable to conclude how the pulley damage occurred. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3162506 | 3133400 | 3029935 | 2955842201301863 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-09 | 2013-05-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-09 | Yes | No | Reuse | 2013
42017910 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si surgical procedure, a monopolar curved scissors instrument was not recognized by the system. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument, showing it was expired. The pogo pins were not stuck and were not contaminated. Additional damage found were deep scratches and a kinked flush tube. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage may be due to mishandling. The instrument was also found with the flush tube kinked inside the housing. Engineering concluded the damage was likely due to mishandling during the cleaning procedure. Electrical continuity test passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3162510 | 3133425 | 3029960 | 2955842201301864 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-22 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the robot would not recognize a maryland bipolar forceps instrument. The instrument was not used in the procedure. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument. The pogo pins did not stick and were not contaminated. Dcp verification of instrument passed. An electrical continuity test passed. The instrument was driven and moved intuitively and the grips opened and closed. An additional finding not reported was a scratch on the main tube. The distal end of the main tube had a scratch mark with light material removal. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3162732 | 3133683 | 3013917 | 2955842201301865 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2013-04-04 | Injury | Required InterventionHospitalization | On (b)(6) 2013, isi received medwatch report number hca (b)(4). The event details are provided below: robotic-assisted total hysterectomy, cystoscopy and ventral hernia repair on (b)(6) 2013. During this surgery an (b)(6) uterus with multiple fibroids were removed. The uterus was noted to be markedly enlarged and irregularly shaped due to the multiple fibroids. Patient developed multiple back bleeders and these were cauterized with bipolar cautery. Due to the "ginormous" size of the uterus, it did need to be morcellated. Indigo carmaine was given iv and the cystoscopy was performed; 70 degree scope was used and both ureteral jets were easily visualized and released indigo carmine. At the completion of this procedure, there was no free flow from vagina. Five days post op ((b)(6) 2013) the patient developed severe pain and clear discharge from vagina. Patient returned to the operating room and it was discovered the patient had sustained a thermal injury to her r ureter. A stent was attempted but unable to pass through area of injury. A percutaneous nephrostomy tube drain was then used. On (b)(6) 2013 successful ureteral re-implantation occurred. | On (b)(4) 2013 intuitive surgical, inc. (isi) contacted the initial reporter regarding the reported event. The initial reporter indicated that patient underwent a da vinci ureteral re-implantation procedure to repair the damage to the patient's ureter. Per the initial reporter, no video recording of the procedure is available. The initial reporter indicated that there was no report by the surgical staff that a malfunction of the fenestrated bipolar forceps instrument occurred during the surgical procedure and no malfunction of the da vinci surgical system or accessories occurred. The hospital's investigation into the reported event did not determine the root cause of the injury to the patient and it was unknown if the injury sustained by the patient occurred as a result of a user error or as a result of a malfunction of the fenestrated bipolar forceps instrument that was not noted during the surgical procedure. The surgeon and surgical staff inspected the instrument for damage and since no damage to the instrument was found the instrument was placed back into circulation. The initial reporter indicated that the instrument will not be returned to isi for failure analysis evaluation. Review of the site's system logs for the reported procedure found no system errors were generated during the surgical procedure that would have caused or contributed to injury sustained by the patient. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3162760 | 3133714 | 3030249 | 2955842201301866 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | RISK MANAGER | Initial | 2013-04-29 | 2013-05-29 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2013-04-29 | No | No | Reuse | 2013
IS3000 A60P3 | 2012-02-02 | Injury | Life Threatening | On (b)(6) 2013, intuitive surgical, inc. (isi) received a legal complaint. The legal complaint alleges that during a da vinci si mitral valve repair procedure, the patient's coronary artery was inadvertently obstructed, causing severe permanent heart damage. The legal complaint alleges that the patient experienced cardiac arrest three times. | On (b)(6) 2013, intuitive surgical, inc. (isi) contacted the hospital's risk management department and claims department. The risk manager and the claims department indicated that they are unable to provide any additional information concerning the reported event since the reported event is in litigation. Due to hippa regulations, the hospital is unable to release the patient's medical records without a signed release form from the patient. Based on the limited information provided regarding the reported event, at this time the root cause of the injuries sustained by the patient has not been determined. A follow-up medwatch report will be submitted to the fda if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3162897 | 3133820 | 3030419 | 2955842201301869 | 1 | NAY | Manufacturer | Other | Other | InitialFollowup | 2013-05-02 | 2013-05-29 | Yes | No | Health Professional | IS3000 A60P3 | No | No | 2013-07-23 | No | No | Reuse | 2013
IS3000 A60P4 | 2013-05-02 | Malfunction |  | It was reported that during a da vinci si partial nephrectomy procedure, the surgeon experienced non-intuitive motion on patient side manipulator (psm) arm 1 while manipulating the right master tool manipulator (mtm) on the surgeon side cart (ssc). The site reseated the sterile adapter and instrument; however, the issue persisted. The site contacted intuitive surgical, inc. Technical support engineering (tse) to report the issue. The tse attempted to help the site troubleshoot the issue; however, at the time of the call, the surgeon declined to troubleshoot the issue as the patient's vessels were cross clamped and the cross clamp time was extended as a result of the reported failure mode. The surgical staff indicated to the tse that the surgeon stated that there was a possibility that the patient may have sustained kidney damage due to the extended cross clamp time. The tse asked the site if the surgeon had considered completing the surgical procedure using open surgical techniques; however, the surgical staff indicated that conversion of the procedure to open had not been considered. Review of the site's system logs by the tse found that an invalid tool error occurred with the installation of three different instruments onto psm 1. The surgeon completed the planned surgical procedure utilizing the left master tool manipulator. No patient harm, adverse outcome or injury was reported to have occurred. | The investigation conducted by the field service engineer (fse) was unable to replicate the failure mode experienced by the customer. The fse tested the patient side manipulator (psm) arm 1 using all of the instruments that were used during the surgical procedure. The sterile adapter used during the surgical procedure was not tested, as the site had already discarded the sterile adapter. All of the instruments passed recognition testing and were ready for follow on matching grip mode. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instrument. Psm 1 was replaced as a precaution. On (b)(4) 2013 intuitive surgical, inc. (isi) contacted the initial reporter regarding the reported event. The initial reporter indicated that the surgeon experienced non-intuitive movement on psm 1 while using a prograsp instrument and needle driver instrument and that the surgical staff reseated the sterile adapter and replaced the instruments, however, the issue persisted. The initial reporter indicated that the patient's vessels were cross clamped approximately 42 minutes. The initial reporter indicated that the patient tolerated the procedure well and to the best of her knowledge the patient has not yet presented any symptoms of kidney damage. On (b)(4) 2013, the isi clinical sales representative who was present during the surgical procedure reported that she spoke to the surgeon after completion of the surgical procedure and the surgeon indicated to her that he felt that the procedure went well. Psm 1 was returned to isi for failure analysis evaluation. Engineering was unable to replicate the failure mode experienced by the customer as functional testing of the psm on an in-house is3000 test system using isi in-house test instruments found that the psm functioned within specifications. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3162854 | 3133834 | 3030369 | 2955842201301867 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-02 | 2013-05-29 | No | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the scissors would not open on a monopolar curved scissors instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to replicate or confirm the reported complaint. The instrument was placed on a system and driven. Recognition and engagement tests passed. The instrument moved intuitively with a full range of motion in all directions. The blades opened and closed properly and a cut test was successful. The instrument was fully functional. Additional damage found was pad printing removal. The instrument tube extension exhibited pad printing removal. Engineering concluded the damage was likely due to mishandling. An electrical continuity test passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; pad printing removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3162826 | 3133847 | 3030382 | 2955842201301868 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-04-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the permanent cautery spatula (pcs) instrument sparked. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the base of the instrument's yaw pulley exhibited burn marks and melted material, indicating that an arcing event occurred. The conductor wire insulation did not exhibit damage and the instrument passed electrical continuity testing. The customer reported complaint does not itself does not constitute a mdr reportable event; however; the arcing damage to the instrument's yaw pulley found during failure analysis may likely cause or contribute to an adverse event if the malfunction were to recur. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3162901 | 3133897 | 3030452 | 2955842201301870 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-29 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-04-25 | Malfunction |  | It was reported that during a davinci s surgical procedure, the customer noted a broken cable on the mega needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment stuck out at wrist. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were 0. 085 - 0. 190 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3163499 | 3134914 | 3031449 | 2955842201301873 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-08 | Malfunction |  | During a da vinci si prostatectomy procedure, the user facility reportedly identified a monopolar curved scissors instrument that was broken on its distal end between the isolator part and the jaw. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. . | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3163556 | 3134960 | 3031495 | 2955842201301874 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-08-13 | 2013-05-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-08-13 | Yes | No | Reuse | 2013
42017910 | 2013-05-06 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a piece of the monopolar curved scissors instrument's tip was missing when it was opened from the peel pack. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. . | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3163583 | 3134971 | 3031510 | 2955842201301875 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-06 | 2013-05-29 | No | Yes | Health Professional | 42017910 | No | Yes | 2013-05-06 | No | No | Initial | 2013
42017207 | 2013-04-25 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted a broken cable on the maryland bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Instrument passed electrical continuity test. Additional observations not reported by site were scratches on the surface of the distal pulley and main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the broken cable if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3163632 | 3135013 | 3031552 | 2955842201301876 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-15 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure the fenestrated bipolar forceps instrument was noted not to be cauterizing. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of the instrument not cauterizing. The instrument was placed on an in-house system for a cautery test and performed with no issues. The instrument passed electrical continuity testing. Engineering found a missing piece of the yaw pulley below the base of one of the grips. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3163682 | 3135060 | 3031599 | 2955842201301877 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-04-30 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that insulation was chipped on the thoracic grasper instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the black tube insulation was found to be damaged at the distal end. The insulation was gouged in various places and it exhibited some material missing. One of the insulation missing materials measured 0. 177 x 0. 081 and another measured roughly 0. 124 x 0. 227, in both instances the metal shaft was exposed as a result. Engineering concluded that damage was likely due to mishandling. 80 - the instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the chipped insulation if to reoccur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3163744 | 3135168 | 3031703 | 2955842201301878 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-22 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the large needle driver instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of no recognition. The instrument was checked for recognition on an in-house system and it was successfully recognized. The pogo pins were not stuck or contaminated. The instrument was driven, moved intuitively and the grips opened and closed. The instrument was fully functional. Engineering also found a deep scratch. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. The evidence was inconclusive, but damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC ACCESSORY | 3163867 | 3135321 | 3031880 | 2955842201301879 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-14 | Malfunction |  | It was reported that during a da vinci si partial nephrectomy procedure, the customer noted a thread on the wrist of the large needle driver instrument, and the endowrist movement was restricted. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cables were loose at the distal clevis due to a broken cable found at the back idler pulleys of the proximal end. Additional observations not reported by site were main tube damage, corroded clamping pulleys and back idler pulleys. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were approximately 0. 042 x 0. 3140 in length and were not aligned with the tube axis. All three clamping pulleys exhibited corrosion/contamination around the groves and the screws. The back idler pulleys exhibited yellowish material on the surface. Engineering concluded that damages found were likely due to improper cleaning and/or mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the scratches found on the main tube if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164024 | 3135457 | 3031996 | 2955842201301880 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-07 | 2013-05-29 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2013-05-07 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-12 | Malfunction |  | It was reported that during a da vinci si surgical procedure the fenestrated bipolar forceps instrument thread was noted to be out at the wrist. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3164141 | 3135533 | 3032068 | 2955842201301881 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-22 | Malfunction |  | During a da vinci si pyeloplasty procedure, the surgeon reportedly complained about dull scissors on the monopolar curved scissors instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed the alleged dull scissors issue. The returned instrument was placed on an in house da vinci system for a latex cut test. The scissors did not cut cleanly through (b)(4) latex and the latex got snagged at the scissor tips. Blade edges were not damaged but the edges exhibited wear at the tips, negatively affecting cut performance. Additional observation that was not reported by the customer were pad printing removal and deep scratches on main tube. Several locations on the tube extension exhibited pad printing removal. The distal end of main tube had scratch marks exhibiting light material removal and a rough surface finish. Evidence not conclusive, but damage may be due to mishandling. Electrical continuity testing was performed and passed. No other damage found. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a mdr reportable event; however,the pad print removal and main tube scratches found during failure analysis evaluation may cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3164180 | 3135543 | 3032078 | 2955842201301883 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-09 | 2013-05-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-09 | Device Not Retu | No | Reuse | 2013
420006 | 2013-01-17 | Malfunction |  | It was reported that during a da vinci si vaso epididymal anastomosis procedure, the customer noted a thread on the wrist of the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164155 | 3135565 | 3032104 | 2955842201301882 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 420006 | No | Yes | 2013-04-30 | Device Not Retu | No | Unkown | 2013
42017207 | 2013-04-22 | Malfunction |  | It was reported that a da vinci si the maryland bipolar forceps instrument was noted to have a problem with the wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of problem in wire. No damages were found to the instrument cables or the conductor wire. The instrument passed electrical continuity testing. The instrument was placed on in-house system and driven. The recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The instrument was fully functional. Engineering also found deep scratches and a char mark. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. The evidence was inconclusive, but the damage was likely due to mishandling. The yaw pulley exhibited a. 0505 char mark near the grip cables grooves. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3164202 | 3135590 | 3032129 | 2955842201301884 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-13 | Malfunction |  | During a da vinci si surgical procedure, the thread reportedly came out of the large suturecut needle driver instrument's wrist. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed the broken grip close cable at the distal idlers and the cable segment was sticking out at wrist. The idler pulley spun freely and was not damaged. Other cables at the instrument's wrist were not damaged. Additional observation that was not reported by the customer were multiple clamping pulleys exhibiting corrosion around the screw head clamping pulley and bearings. Some cables were partially buried under this residue showing one loose wire. In addition, the cover of the housing showed some stain on the outside cover. Evidence not conclusive, but failure most likely due to improper cleaning. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164185 | 3135602 | 3032161 | 2955842201301885 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-10 | 2013-05-29 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-10-02 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure the surgeon believed the monopolar curved scissors instrument tips may be blunt. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was placed on a system for a latex cut test. The scissors cut cleanly through. 006 of latex. The blades were not damaged but did exhibit some wear at the instrument tips. Additional damage found were deep scratchs and pad printing was removal. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage may be due to mishandling. The tube extension also exhibited pad printing removal at the midpoint. Electrical continuity test passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal and or pad printing removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3164255 | 3135614 | 3032149 | 2955842201301888 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-29 | 2013-05-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-29 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-17 | Malfunction |  | During a da vinci si surgical procedure, the user facility reportedly identified a thread that came out of the large needle driver instrument's wrist. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed the a loose grip cable at the distal idlers. The idler pulley spun freely and was not damaged. Upon opening up the instrument's housing, a broken grip cable was found causing the alleged loose wire by the instrument's tip area. Engineering also noted corrosion on the clamping pulleys and orange color residue around clamping pulley around the bearings. The screws were found to be secured. Evidence not conclusive, but damage most likely due to improper cleaning. Additional observation that was not reported by the customer were pulley and main tube damages. Engineering noted an indentation at the edge of the distal pulley. There were also visible scratches on the surface of the pulley. The distal end of main tube had various scratch marks showing light material removal. Scratches were short in length and not axially aligned with the tube. Engineering concluded that pulley and main tube damages may likely be due to mishandling. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164191 | 3135654 | 3032104 | 2955842201301887 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-10 | 2013-05-29 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-20 | Malfunction |  | It was reported that during a da vinci si gynecology procedure, the customer noted a thread on the wrist of the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was found to be broken at the distal clevis hub. The cable crimp was missing from the clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. Additional observations not reported by site were indentations on the edge of the distal pulley and main tube damage. Engineering observed scratches on the surface of the distal pulley. The distal end of the main tube also exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3164253 | 3135659 | 3032198 | 2955842201301886 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-18 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the prograsp forceps instrument jaws were noted not to open properly. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip cable was derailed at the pulley in between the distal clevis. The grips still opened and closed, but not fully. Engineering also found deep scratches on the main tube. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. The evidence was inconclusive, but damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164273 | 3135671 | 3032210 | 2955842201301889 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-01 | 2013-05-29 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-29 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the customer noted a frayed cable on the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be frayed at the distal clevis hub. The frayed strands stuck out at the wrist of the instrument. - additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were 0. 095 - 0. 165 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164294 | 3135675 | 3032214 | 2955842201301890 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-01 | 2013-05-29 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-05-09 | Malfunction |  | After a da vinci si hysterectomy procedure, the user facility reportedly identified a separated wire on the mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed two frayed grip cables at the distal clevis: one cable close near the proximal clevis and the other cable close to the grips. One frayed strand was found sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164281 | 3135704 | 3032239 | 2955842201301891 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-10 | 2013-05-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-10 | Device Not Retu | No | Reuse | 2013
40027202 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the bottom tip of the harmonic ace curved shears insert accessory installed on the harmonic ace curved shears instrument broke off and fell into the patient's cavity. The broken accessory piece was retrieved. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found that the insert accessory was returned with a. 660 long segment of the curved blade detached from the distal end. The curved blade fractured within the shaft, adjacent to the clamp arm pin. Engineering concluded that the location of the fracture may suggest that the curved blade made contact with the clamp arm during energy activation, likely creating a crack that propagated through the blade. No other damage found. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC INSTRUMENT ACCESSOR | 3164287 | 3135730 | 3032269 | 2955842201301892 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-30 | 2013-05-29 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Unkown | 2013
42009308 | 2013-04-29 | Malfunction |  | During a da vinci si prostatectomy procedure, the user facility reportedly identified a prograsp forceps instrument that was not functioning properly. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Additional observations found that were not reported by the customer were pulley and main tube damages. Engineering found scratches on the surface of the distal pulley. The distal end of main tube had various scratch marks showing light material removal. Scratches were short in length and not axially aligned with the tube. Evidence not conclusive, but damages to the pulley and main tube may be due to mishandling. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported and observed malfunctions may cause or contribute to an adverse event, if they were to reoccur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164394 | 3135782 | 3032321 | 2955842201301893 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2013-05-13 | 2013-05-29 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-17 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy, a prograsp forceps instrument was not recognized by the robot. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument. The pogo pins did not stick and were not contaminated. Dcp verification of the instrument passed. An additional observation found the instrument grip cable was derailed at the distal clevis, resulting in limited motion of the grip jaws. The pitch cable was frayed at the distal clevis hub, approximately 0. 4 in length. No excessive damage was found on the pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the frayed cable, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164399 | 3135805 | 3032324 | 2955842201301894 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-01 | 2013-05-29 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-05-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted that the mega suture cut needle driver instrument was no longer responding; upon inspection, it was noticed that a cable was damage. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip cable was found to be broken at the distal idlers. The idler pulley was able to spin freely, but it exhibited some wear. The broken cable segment stuck out at the wrist approximately 0. 4270 in length. Other cables at wrist were not damaged. Additional observations not reported by site were distal pulley indentation and main tube damage. The distal idler pulley located opposite side of the broken cable exhibited an indentation at the edge. The distal end of the main tube also exhibited a scratch mark with light material removal. The scratch mark was approximately 0. 0240 in length and it was not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 3164401 | 3135807 | 3032326 | 2955842201301895 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-03 | 2013-05-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
42011006 | 2013-04-30 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the precise bipolar forceps instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of not recognized by system. The instrument was checked for recognition on an in-house is3000 system and it was successfully recognized. The pogo pins exhibited a dark discoloration, but did not stick when actuated. The dcp (dallas chip programmer) verification of the instrument passed. Engineering was unable to replicate recognition issue. Engineering also found the bipolar pins were dislodged and main tube damage. The top bipolar pin was pushed into the back end and the bottom pin was bent. The chassis feature that acted as a hard stop for the pins was broken. The evidence was inconclusive, but the instrument was likely mishandled, breaking the chassis feature. The distal end of the main tube has various scratch marks with light material removal. The scratches were. 125 -. 800 in length and some were aligned with the tube axis. The evidence was inconclusive, but the tube damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3164427 | 3135838 | 3032377 | 2955842201301896 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-01 | 2013-05-29 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-05-01 | Malfunction |  | It was reported that during reprocessing a bowel grasper instrument, a knick in the insulation was found. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. Black tube insulation was damaged at the distal end. The insulation was gouged on one side and had a section lifted up roughly 5. 5 above the snake wrist. When the gouged insulation segment is pushed down, it covers the exposed section of main tube, indicating no material was removed from the insulation. Additional observation not reported by site was a bent tube. The tube was bent at the proximal end, just below the flush hole in the shaft. The distal tip of the instrument wobbled when shaft was rotated due to axial misalignment caused by bending. Engineering concluded the damage was likely due to mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3164430 | 3135842 | 3036923 | 2955842201301897 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-01 | 2013-05-29 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No | Reuse | 2013
42805912 | 2013-03-15 | Malfunction |  | It was reported that during a da vinci si surgical procedure the crocodile grasper instrument shaft was noted to be wrinkled up. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the main tube had various scratch marks with lifted material but no missing pieces. The scratches were short in length and not aligned with the tube axis. The evidence was inconclusive however, was likely caused by instrument collisions or rough handling. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | CROCODILE GRASPER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3164549 | 3135945 | 3032484 | 2955842201301898 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2013-05-01 | 2013-05-29 | No | Yes | Health Professional | 42805912 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No | Reuse | 2013
42011006 | 2013-04-05 | Malfunction |  | It was reported that during a da vinci si surgical procedure the precise bipolar forceps instrument was working at the beginning of the surgery. The surgeon used another instrument and then wanted to reuse the precise bipolar forceps, it was not recognized by the system. They tried some manipulations (remove, replace) but it was still not recognized. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument, showing 3 uses left. The pogo pins were not stuck and were not contaminated. Dcp verification of instrument passed. Electrical continuity test passed. The instrument was driven and moved intuitively and grips opened and closed. Additional findings not reported were scratches on the main tube. The distal end of main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3164571 | 3135965 | 3032504 | 2955842201301899 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-02 | 2013-05-29 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
| 2013-05-28 | Malfunction |  | During the or procedure, the davinci scissor tip was burned during use. Burn went through scissor tip to instrument and burned a coating on end of scissor. The packaging was not retained thus cannot submit all the numbers on package, just the device itself. What was the original intended procedure?the patient was having robotic surgery for removal of uterine fibroids. There was no harm to the patient. Device #1is this a laboratory device or laboratory test?no. |  | DAVINCIS MONOPOLAR CURVED SCISSORS 8MM | SYSTEM SURGICAL COMPUTER CONTR | 3177846 | 3148690 | 3045225 | 3148690 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-05-29 | 2013-05-29 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
| 2010-08-31 | Injury | Other | Radical prostatectomy using da vinci robot, (b)(6) hospital. Developed post - surgery groin pain that required wheel chair use. (b)(6) hospital diagnosed pain as inguinal hernia associated with prostate surgery above. Corrective surgery on (b)(6) 2011, (b)(6) hospital implanted with johnson and johnson ethicon prolene hernia mesh. Complications have increased to include nerve damage, bowel problems, swollen testicle, leg and groin pain, kidney pain, flatulence, increased stress urinary incontinence and increased blood pressure. Va has ruled this is associated with hernia repair above, which as noted above was ruled by va to be associated with the da vinci surgery. |  | INTUITIVE SURGICAL | DA VINCI ROBOTIC SURGICAL SYST | 3173327 | 3140856 | 3037391 | MW5030353 | 1 | NAY | Voluntary |  | Patient | Initial | 2013-05-20 | 2013-05-28 | Yes | Yes | Service Personnel |  |  | No |  |  |  |  | 2013
42017910 | 2011-11-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that a flash arced from the monopolar curved scissors (mcs) tip cover accessory (installed on the mcs instrument) and burnt the tip cover and the patient's ovarian tube. The initial reporter indicated that this was not an issue since the patient's ovarian tubes and ovaries were being removed as part of the planned surgical procedure. No further issues were reported and the procedure was completed. No patient harm, adverse outcome or injury was reported. | This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. Engineering evaluation also found a char mark on the instrument proximal clevis, adjacent to one of the window features that exposes the clevis seal. The tip cover accessory used in conjunction with the mcs instrument during the procedure was returned. Per the customer reported complaint, engineering observed a hole burnt through the interface between the silicone and ultem sleeve. The silicone exhibits localized melting around the hole, indicating an arcing event occurred. The hole size is approximately. 045 in diameter. A small tear was also observed on the distal end of the silicone tip. The tip cover hole position aligns with the char mark on the mcs instrument's clevis, which leads engineering to believe that the char mark is likely where energy escaped through the tip cover. Additionally, this event was initially reported via medwatch report number 2955842-2011-00418. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159808 | 3130937 | 3027475 | 2955842201301856 | 1 | NAY | Manufacturer | Other | Nurse | Initial | 2011-11-30 | 2013-05-25 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-27 | Device Not Retu | No | Reuse | 2013
| 2013-04-29 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure that arm 3 did not recognize instruments. The customer verified the sterile adapter positions and reseated the drape. The system was power cycled and the issue was cleared shortly. No parts were replaced. The issue was due to a damaged cannula mount. The field engineer found the system was recognizing instruments but there was an issue with advancing the instrument through the cannula. The field engineer had the site ensure that the drape was not getting stuck in the arm and that the sterile adapter was seated properly. Upon the field engineer's arrival, it was noted that the cannula was faulty. The field engineer found that cannula was badly damaged. It was bent and had a large crack. The site stated the cannula was not tested prior to the procedure, and the field engineer advised the site to test each cannula prior to procedure to prevent another issue. No patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the cannula, which was found during troubleshooting by the field engineer could likely cause or contribute to an adverse event if the failure mode were to recur. | CANNULA ACCESSORY | ENDOSCOPIC INSTRUMENT ACCESSOR | 3157559 | 3128629 | 3025164 | 2955842201301816 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-29 | 2013-05-24 | No | Yes | Health Professional |  | No | No | 2013-04-29 | No | No | Unkown | 2013
42029602 | 2013-04-26 | Malfunction |  | It was reported that during a da vinci si sigmoid colectomy procedure the large suturecut needle driver instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the grip cable was frayed at the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at the wrist were undamaged. Engineering also found the same grip cable was derailed at the distal idler pulley. The grips could still open and close, but movement may not be precise. The cable derailment was likely due to cable losing contact with pulley during wrist articulation. The fleet angle between the proximal and the distal pulleys may have contributed to derailment. There was an indentation at the edge of the distal pulley. The evidence is not conclusive however; it was likely due to mishandling. Also visible scratches on the surface of the pulley. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157550 | 3128630 | 3025165 | 2955842201301820 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-29 | 2013-05-24 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-29 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the pk dissecting forceps instrument was noted to have delayed function. During cleaning there was noted to be visible disintegration of the cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received for evaluation on (b)(4) 2013. The instrument has not yet been evaluated; therefore, the root cause of the customer reported failure mode has not been determined. A follow-up mdr will be submitted when the instrument has been evaluated, or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. . | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3157573 | 3128632 | 3025167 | 2955842201301821 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-05-23 | 2013-05-24 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-05-23 | Yes | No | Reuse | 2013
42017207 | 2013-04-26 | Malfunction |  | It was reported that during a da vinci si sigmoid colectomy procedure the maryland bipolar forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch down cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. Engineering also found tube damage. The distal end of main tube had various scratch marks with light material removal. The scratches were. 090 -. 175 in length and not aligned with the tube axis. The evidence was not conclusive, but damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3157561 | 3128633 | 3025168 | 2955842201301822 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-29 | 2013-05-24 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-29 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si sigmoid colectomy procedure, several wires broke on a bowel grasper instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. Two drive cables were broken at the distal end of the snake wrist. The broken strands protruded out of the snake wrist. The instrument wrist could not straighten and motion was not intuitive due to cable breakage. Additional observation not reported by site was tube damage. The distal end of the main tube had light scratch marks in the insulation. Scratches were not deep and there was no material missing from the insulation. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157565 | 3128634 | 3025169 | 2955842201301827 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-04-18 | Malfunction |  | It was reported that during a da vinci si sigmoid colectomy procedure, several wires broke on a bowel grasper instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip cable was broken at the snake wrist. The cable segment stuck out at the wrist approximately. 1555 in length. The other cables at the wrist were not damaged and no additional damage was found at the snake wrist. Additional damage found were signs of abrasion/wear to the main tube insulation. No evidence of any missing material found. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157566 | 3128635 | 3025170 | 2955842201301828 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-25 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted a broken cable on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at wrist. Other cables at wrist were not damaged. Engineering also observed an indentation and scratches on the inner edge of the distal pulley. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157546 | 3128646 | 3025184 | 2955842201301814 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-01 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted that the tips of the mega suture cut needle driver instrument were not moving. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that intuitive motion was not being experienced due to a broken grip cable. The broken grip cable stuck out at the wrist. The idler pulley was able to spin freely and exhibited pulley rim and face damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157548 | 3128648 | 3025186 | 2955842201301815 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-04-26 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical thoracic grasper instrument black sheath was noted to be coming off. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the black tube insulation was damaged at the distal end. The insulation was gouged in various places near the distal end and had a. 0415 x. 0665 piece of insulation missing. The metal shaft was exposed at two different locations as a result. Engineering also found a bent main tube shaft. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157560 | 3128649 | 3025187 | 2955842201301817 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-26 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s prostatectomy procedure, the customer noted a cut cable on the prograsp forceps instrument. The instrument was not responding. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable at distal idler. Engineering also observed scratches on the surface of the distal pulley. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the main tube damage found during failure analysis if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS INSTRUMENT | Reuse | 3157549 | 3128650 | 3025188 | 2955842201301818 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-06-13 | 2013-05-24 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-06-13 | Yes | No |  | 2013
42034403 | 2013-04-26 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical curved bipolar dissector instrument black sheath was noted to be coming off. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. The evidence was not conclusive, but damage was likely due to mishandling. The instrument passed electrical continuity testing. Engineering also found the one grip was bent, causing side-to-side misalignment of grips. There was a. 050 offset at the tips, indicating overloading at the tip. The evidence was not conclusive; however, damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CURVED BIPOLAR DISSECTOR INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3157551 | 3128651 | 3025189 | 2955842201301819 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42034403 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-28 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the maryland bipolar forceps instrument was noted to have a torn bowden at the top of the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the broken conductor wire allowing the grip to move within the pulley and have a larger range of motion in the yaw. The yaw pulley was not broken. The instrument failed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. The evidence was not conclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3157574 | 3128652 | 3025190 | 2955842201301823 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-29 | 2013-05-24 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-29 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-18 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical prograsp forceps instrument tip spring was noted to be broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the grip cable was frayed at the distal idler pulley. No other cables were damaged but the distal idler pulley exhibited indentation along the edges. The evidence was not conclusive but the damage to the distal idler pulley was likely due to mishandling. Engineering also found deep scratches on the main tube. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. The evidence was not conclusive, but damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157575 | 3128653 | 3025191 | 2955842201301825 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-29 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the mega suturecut needle driver instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. There was an indentation on the edge of the distal pulley. The evidence is not conclusive; however, the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157576 | 3128654 | 3025192 | 2955842201301824 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-29 | 2013-05-24 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-29 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-01 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted the tips of the mega suture cut needle driver instrument were not moving. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that intuitive motion was not being experienced due to a broken pitch cable. The pitch cable was broken at the wrist. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURE CUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157577 | 3128656 | 3025186 | 2955842201301826 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42805312 | 2013-04-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, a broken wire at the distal end of the medium-large instrument was identified during set up. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument does not have any wires or drive cables. The grips are driven open/closed by a push pull rod. Additional observation not reported by site was a broken clevis. Both ears on the clevis were broken at their base. The ears were retained by grip axis pin. The grips were dislodged from the pushrod cross pin and the open/close of the grips was no longer functional due to clevis breakage. Engineering concluded the damage was likely due to mishandling. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEDIUMLARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157581 | 3128663 | 3025201 | 2955842201301829 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42805312 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-03-28 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure dr (b)(6) was about to stitch the vag cuff. The v-loc suture got caught on wire of a davinci mega needle driver instrument. As physician used cobra instrument to remove stitch, wire popped on needle holder. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157582 | 3128664 | 2984211 | 2955842201301830 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-04-25 | Malfunction |  | It was reported that during a da vinci s sacrocolpopexy procedure, upon suturing peniotneum, noticed cable on instrument busted during davinci. Removed instrument, no injury to patient, replaced instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. Additional finding was an indentation at the edge of the distal pulley. Engineering concluded it was most likely due to mishandling. Also visible were scratches on the surface of the pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157584 | 3128666 | 3025204 | 2955842201301831 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the pk dissecting forceps instrument was docked and was not burning and noticed small wire stuck out. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3157585 | 3128667 | 3021466 | 2955842201301832 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-29 | 2013-05-24 | No | Yes | Health Professional | 42022703 | No | No | 2013-04-29 | No | No | Reuse | 2013
42022703 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si tounge base resection, the pk dissecting instrument was docked and surgeon started dissecting and after a few minutes it stopped working not burning. Undocked instrument and noticed small wire came out. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3157586 | 3128668 | 3021466 | 2955842201301833 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-29 | 2013-05-24 | No | Yes | Health Professional | 42022703 | No | No | 2013-04-29 | No | No | Reuse | 2013
42001105 | 2013-04-26 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure the instrument was noted not to slide through the cannula. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that instruments fit through the cannula, but with increased friction. The cannula tube diameter was checked with the obturator that was returned. The obturator fit through the cannula, but there was resistance felt towards the distal end of the cannula. A 5mm instrument was also installed and fit through the cannula with slightly increased resistance. The cannula was laid on bowl and rotated to check for wobble at the tip. The tip was observed to wobble as the cannula was rotated, suggesting the tube was bent. The cannula tube appeared to be slightly bent near the pn and lot laser marked area. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- | CANNULA 5MM | Reuse | 3158200 | 3129275 | 3025813 | 2955842201301811 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42001105 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No |  | 2013
42034301 | 2013-04-26 | Malfunction |  | During a da vinci si prostatectomy procedure, the articulation cable on the thoracic grasper instrument allegedly broke. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed that the pitch cable was broken and sticking out of instrument's snake wrist. Wrist motion was not intuitive as a result of cable breakage. Other cables at the instrument's snake wrist were not damaged. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3158382 | 3129539 | 3026057 | 2955842201301834 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-29 | 2013-05-24 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-04-29 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-17 | Malfunction |  | It was reported that during a da vinci si surgical procedure the maryland bipolar forceps instrument was noted to have poor articulation. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was broken at the proximal clevis to the main tube interface, which likely was the cause of poor articulation. The clevis was dislodged from the main tube as a result. The main tube had missing pieces. The breakage was likely due to overloading at the tip. Engineering also found corrosion. The back idler pulleys had corrosion along the edges. The evidence was not conclusive, but damage was likely due to improper cleaning. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159211 | 3130426 | 3026961 | 2955842201301835 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-30 | 2013-05-24 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-06-12 | Malfunction |  | It was reported that during a da vinci si cystectomy procedure, the tip was broken on a prograsp forceps instrument due to user error. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3159281 | 3130475 | 3026993 | 2955842201301836 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42009308 | No | No | 2013-04-26 | No | No | Reuse | 2013
42017910 | 2013-03-14 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the customer noted that the monopolar curved scissors (mcs) instrument was unable to cut. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the instrument was placed on the in-house da vinci robot system for latex cut test. The mcs instrument could not cut cleanly through 0. 006 latex. The latex got snagged at instrument's scissor tips. The blade edges were not damaged, but the edges exhibited wear at the tips, negatively affecting cut performance. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a mdr reportable event; however; main tube crack found during the failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159517 | 3130663 | 3027198 | 2955842201301837 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-27 | 2013-05-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-01 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the customer noted that the monopolar curved scissors (mcs) instrument was dull. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the instrument was placed on the in-house da vinci robot system for latex cut test. The mcs instrument could not cut cleanly through (b)(4) latex. The latex got snagged at instrument's scissor tips. The blade edges were not damaged, but the edges exhibited wear at the tips, negatively affecting cut performance. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a mdr reportable event; however; main tube crack found during the failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159530 | 3130683 | 3014314 | 2955842201301838 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2013-04-27 | 2013-05-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-10 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted that orange casing was broken on the monopolar curved scissors (mcs) instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension was found to be cracked with no material removed at the proximal clevis interface. The clevis was dislodged from the tube extension. Engineering concluded that the instrument broke due to excessive side loading or other mishandling. Instrument passed electrical continuity test. Additional observation not reported by site was a frayed grip cable located at the distal idler pulley, there was no damage found on the pulley. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a mdr reportable event; however; main tube crack found during the failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159581 | 3130734 | 3027269 | 2955842201301839 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2013-04-27 | 2013-05-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-05 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted that orange part on the monopolar curved scissors (mcs) instrument was loose. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension was found to be broken with a missing piece located at the proximal clevis interface. The clevis was dislodged from the tube extension. Engineering concluded that the instrument broke due to excessive side loading or other mishandling. Instrument passed electrical continuity test. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a mdr reportable event; however; main tube crack found during the failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159613 | 3130752 | 3011442 | 2955842201301840 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2013-04-27 | 2013-05-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted arcing while using the monopolar curved scissors (mcs) instrument and mcs tip cover. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on in-house da vinci robot system and driven with no problem. Instrument passed recognition and engagement tests. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly, and performed cautery functions with no trouble. The tip cover accessory used in conjunction with the mcs instrument during the procedure was discarded by the hospital and will not be returned. Additionally, this event was initially reported on (b)(4) 2013, via medwatch report number 2955842-2013-01752. The customer reported complaint does not itself constitute a mdr reportable event; however; main tube crack found during the failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. 48 - this instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159609 | 3130759 | 2985127 | 2955842201301841 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-27 | 2013-05-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-27 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-26 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure, the surgical staff alleged that the tip of the mega suturecut needle driver instrument broke off and fell into the patient. The surgical staff indicated that the fragment was retrieved and the surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, one grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely. However, engineering evaluation found damage on the surface of the idler pulley. The cable segment was found to be sticking out at the instrument's wrist. The other cables of the instrument's wrist did not exhibit any damage. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | Reuse | 3159698 | 3130849 | 3027384 | 2955842201301844 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-04-30 | 2013-05-24 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No |  | 2013
IS2000 A51P7 | 2013-02-26 | Injury | Required InterventionLife Threatening | It was reported that the patient underwent a da vinci s hysterectomy procedure on (b)(6) 2013, and on (b)(6) 2013, the patient allegedly experienced septic shock. The surgeon indicated that the hysterectomy procedure went well and he did not observe any errors during the case. According to the surgeon, the patient went to the emergency room (er) in septic shock on (b)(6) 2013, and a second doctor performed a diagnostic laparotomy. During the diagnostic laparotomy, the surgeon who performed the hysterectomy procedure was present. He claimed that no factors were identified for the cause of the sepsis during the diagnostic laparotomy. The second doctor then performed an exploratory laparotomy per the request of the surgeon since the patient was in sepsis. During the exploratory laparotomy, the patient's appendix was found to be inflamed and a small mark was observed on the patient's bowel. The surgeon denied that feces was observed in the patient's abdominal cavity. The second doctor reportedly applied several sutures over the small mark on the patient's bowel and a temporary diverting ileostomy was created. | On (b)(6) 2013, intuitive surgical inc. (isi) contacted the isi clinical sales representative (csr) who was present during the da vinci s hysterectomy procedure. The csr indicated that he did not observe any malfunction of the da vinci s surgical system, an instrument, or an accessory during the surgical procedure. The csr reportedly spoke to the surgeon, who performed the hysterectomy procedure, regarding the patient's return to the hospital on (b)(6) 2013. The surgeon reportedly informed the csr that the mark on the patient's bowel appeared to be cylindrical in shape and was smaller than the size of a dime. The surgeon also reportedly informed the csr that the mark appeared to be a burn mark that was more superficial. According to the csr, the surgeon concluded that the burn mark was probably due to an instrument used during the hysterectomy procedure. The csr indicated that there was no evidence that the patient's bowel was perforated. During the exploratory laparotomy procedure, the second doctor removed the patient's appendix due to being inflamed. The csr did not know the current status of the patient. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. Based on the information provided, there is no clear evidence as to what caused or contributed to the patient's septic shock, the alleged burn mark, or the inflammation of the patient's appendix. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3159716 | 3130866 | 3027404 | 2955842201301842 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2013-04-26 | 2013-05-24 | Yes | No | Health Professional | IS2000 A51P7 | No | Yes | 2013-04-26 | No | No | Reuse | 2013
IS1200 A43P9 | 2013-04-26 | Other | Other | It was reported that during a da vinci prostatectomy procedure, the surgical staff indicated that an unspecified scissor instrument was not fully articulating. However, the surgical staff claimed other instruments were working normally on the same arm on the patient side cart (psc). In addition, the surgical staff indicated that the same scissor instrument was working normally on a different arm on the psc. Due to the alleged instrument articulation issue, the initial reporter of this complaint indicated that the doctor made the decision to abort the surgical procedure. | On (b)(4) 2013, intuitive surgical inc. (isi) contacted the initial reporter of this complaint to obtain additional information regarding the reported event. The reporter indicated that the alleged articulation issue started occurring 2-3 hours into the da vinci prostatectomy procedure. According to the reporter, the surgical staff attempted to troubleshoot the alleged issue for about 1 hour before the surgeon made the decision to convert the surgical procedure to open surgical techniques. The reporter denied that the patient experienced any intra-operative or post-operative complications because of the alleged issue. The reporter also indicated that the patient was discharged as planned. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3159697 | 3130868 | 3027406 | 2955842201301843 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-26 | 2013-05-24 | Yes | Yes | Health Professional | IS1200 A43P9 | No | Yes | 2013-04-26 | No | No | Reuse | 2013
42030901 | 2013-04-25 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff indicated that mega suturecut needle driver instrument had frayed wires and an unspecified fragment broke off and fell into the patient. The surgical staff indicated that the fragment was retrieved and the surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation confirmed that the instrument had frayed wires. The instrument was found to have one grip cable broken at the distal idlers. The broken cable segment was found to be. 3330 in length and was sticking out at the instrument's wrist. The other cables of the instrument's wrist did not exhibit any damage. Engineering evaluation also observed damage along the edges of the idler pulleys. Each idler pulley was found to have an indentation on the edge. Engineering evaluation concluded that the damages on the pulleys were likely due to mishandling. The idler pulleys were able to spin freely. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3159755 | 3130907 | 3021561 | 2955842201301845 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-24 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-05-01 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff alleged that arcing was observed with the permanent cautery spatula instrument. The planned surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. A visual inspection of the instrument showed a char mark on the yaw pulley in between the conductor wire and drive cable. Also, light charring was found on the inner surface of one distal clevis ear. Engineering evaluation indicated that the char marks were evidence of an arcing event. Engineering evaluation also found a damaged conductor wire. The conductor wire insulation had a section missing adjacent to the char mark on the yaw pulley. Bare wire on the conductor wire was observed to be exposed. Electrical continuity testing of the instrument passed. On (b)(4) 2013, intuitive surgical inc. (isi) obtained additional information regarding the reported event from a nursing supervisor at the site. The nursing supervisor provided the following information: the surgical procedure was not recorded. The permanent cautery spatula instrument was inspected before the surgical procedure was performed and no issues were observed. She also indicated that the instrument was in use for about 10 minutes before the arcing event was first noticed. Due to the alleged arcing issue, the surgical staff removed and replaced the permanent cautery spatula instrument with another spatula instrument. At the time the arcing event occurred, a fenestrated bipolar forceps instrument was also in use. The nursing supervisor denied that there were any instrument collisions during the surgical procedure. She indicated that the electric energy arced from the permanent cautery spatula instrument to other instruments. At the time the arcing event occurred, the site was using a valley lab electrosurgical generator unit (esu) with a setting of 35 for cutting and coagulation. The nursing supervisor indicated that the grounding pad wash positioned correctly. She denied that the patient experienced any intra-operative or post-operative complications. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159737 | 3130908 | 3027426 | 2955842201301846 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-24 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-27 | Malfunction |  | During the da vinci s prostatectomy procedure, the surgeon could not cut the tissue smoothly. According to the customer, the surgical staff attempted to solve this problem by resetting the monopolar scissors instrument; however, the problem could not be solved. The surgeon felt an abnormal cutting feeling comparing to before. The staff exchanged it into a backup instrument, and the planned procedure was completed without problem. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the instrument was placed on the in-house da vinci robot system for latex cut test. The mcs instrument could not cut cleanly through (b)(4) latex. The latex got snagged at instrument's scissor tips. The blade edges were not damaged, but the edges exhibited wear at the tips, negatively affecting cut performance. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a mdr reportable event; however; main tube crack found during the failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159799 | 3130927 | 3042085 | 2955842201301849 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-27 | 2013-05-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-07-26 | Malfunction |  | It was da vinci si myomectomy procedure, the wrist on the monopolar curved scissors (mcs) instrument broke apart. There were no missing or fallen pieces reported. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a broken tube extension and missing a piece at the proximal clevis interface. The clevis was dislodged from tube extension. Engineering concluded that the tube extension likely broke due to excessive side loading. The instrument also passed electrical continuity testing. The customer reported complaint does not itself constitute a mdr reportable event; however; the cracks in the instrument's main tube, found during secondary failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159801 | 3130929 | 3027467 | 2955842201301851 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-07-26 | 2013-05-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2012-07-27 | Device Not Retu | No | Reuse | 2013
40017910 | 2013-01-28 | Malfunction |  | It was reported that during a da vinci urological procedure, during preparation of the patient's kidney tumor, the monopolar curved scissors instrument moved with difficulty. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode with the following clarification that the grip cable on the instrument was broken, making it incapable of proper movement. One grip cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment was sticking out at the wrist. The other cables at the wrist were not damaged. The instrument also passed electrical continuity testing. No other damage was found. This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. The customer reported complaint does not itself constitute a mdr reportable event; however; the cracks in the instrument's main tube, found during secondary failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159802 | 3130930 | 3027472 | 2955842201301853 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-27 | 2013-05-24 | No | Yes | Health Professional | 40017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
40017910 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci surgical procedure, during preparation at the patient's ureter, it was impossible to move the monopolar curved scissors (mcs) instrument. The bowden was observed to be torn. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode with the following clarification that the grip cable on the instrument was broken, making it incapable of proper movement. The idler pulley spun freely and was not damaged. The cable segment was sticking out at the wrist. The other cables at the wrist were not damaged. The instrument also passed electrical continuity testing. No other damage was found. This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. The customer reported complaint does not itself constitute a mdr reportable event; however; the cracks in the instrument's main tube, found during secondary failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159805 | 3130934 | 3027472 | 2955842201301854 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-27 | 2013-05-24 | No | Yes | Health Professional | 40017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
40017910 | 2013-02-21 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, the scissors on the monopolar curved scissors (mcs) instrument did not cut. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode with the following clarification that the grip cable on the instrument was broken, making the instrument incapable of cutting. One grip cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment was sticking out at the wrist. The other cables at the wrist were not damaged. The instrument passed electrical continuity testing. No other damage was found. This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. The customer reported complaint does not itself constitute a mdr reportable event; however; the cracks in the instrument's main tube, found during secondary failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159806 | 3130935 | 3027473 | 2955842201301855 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-27 | 2013-05-24 | No | Yes | Health Professional | 40017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
40017910 | 2013-03-07 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, it was not possible to move the monopolar curved scissors (mcs) instrument. Reportedly, the bowden was torn. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode with the following clarification that the instrument had a broken grip cable making it incapable of proper movement. One grip cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment was sticking out at the wrist. The other cables at the wrist were not damaged. The instrument also passed electrical continuity testing. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; main tube crack found during the failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159807 | 3130936 | 3027474 | 2955842201301852 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-27 | 2013-05-24 | No | Yes | Health Professional | 40017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
42017910 | 2011-09-12 | Malfunction |  | It was reported that post a successful da vinci si prostatectomy procedure, the patient returned to the hospital experiencing complications. It was discovered that the patient had a perforation in his right colon, approximately 6 cm from the caecum. The monopolar curved scissors (mcs) instrument was inspected prior to use. Additionally, no complications were observed by the surgeon or operating room staff with the mcs instrument or the mcs tip cover accessory during the procedure. | This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. Engineering evaluation was unable to find any evidence of arcing or melted materials. The instrument was placed on an in house test system for performance testing and no trouble was found. The tip cover accessory used in conjunction with the mcs instrument during the procedure was discarded by the hospital and will not be returned. Additionally, this event was initially reported via medwatch report number 2955842-2011-00330 on (b)(4) 2011. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159795 | 3130953 | 3027488 | 2955842201301847 | 1 | NAY | Manufacturer | Other | Other | Initial | 2011-09-21 | 2013-05-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-25 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted that the monopolar curved scissors (mcs) instrument was dull. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the instrument was placed on the in-house da vinci robot system for latex cut test. The mcs instrument could not cut cleanly through 0. 006 latex. The blade edges were damaged and circular shaped dents were found, negatively affecting cut performance. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a mdr reportable event; however; main tube crack found during the failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3159778 | 3130956 | 3027491 | 2955842201301848 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-27 | 2013-05-24 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-27 | Malfunction |  | During the da vinci s prostatectomy procedure, the surgeon could not cut the tissue smoothly. According to the customer, the surgical staff attempted to solve this problem by resetting the monopolar scissors instrument; however, the problem could not be solved. The surgeon felt an abnormal cutting feeling comparing to before. The staff exchanged it into a backup instrument, and the planned procedure was completed without problem. No pt harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, the instrument was placed on the in-house da vinci robot system for latex cut test. The mcs instrument could not cut cleanly through 0. 006" latex. The latex got snagged at instrument's scissor tips. The blade edges were not damaged, but the edges exhibited wear at the tips, negatively affecting cut performance. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a mdr reportable event; however, main tube crack found during the failure analysis eval could likely cause or contribute to an adverse event if the failure mode were to recur. This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of these micro-cracks is confirmed to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3188813 | 3145550 | 3042085 | 2955842201201849 | 1 | NAY | Manufacturer | DistributorOther | NOT APPLICABLE | Initial | 2013-04-27 | 2013-05-24 | No | Yes | Health Professional | 42017910 | Yes | No | 2013-04-27 | Yes | No | Reuse | 2013
42017910 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, a monopolar curved scissors instrument scissors were dull. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was placed on system for a latex cut test. The scissors did not cut cleanly through. 006 of latex. The latex got snagged at the scissor tips. There was an indentation on the instrument blade tip, which acted as a mechanical stop when trying to close the blades. This damage can contribute to poor cut performance. Engineering concluded the damage was likely due to mishandling. An additional finding not reported was pad printing removal and a dislodged flush tube. The flush tube was found sticking out of the flush port 1. Engineering concluded this damage was likely due to mishandling. An electrical continuity test passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; pad printing removal, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3154771 | 3125917 | 3022452 | 2955842201301793 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-25 | 2013-05-23 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure a monopolar curved scissors instrument scissors were dull. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was placed on a system for a latex cut test. The scissors did not cut cleanly through. 006 of latex. The latex got snagged at the scissor tips. An indentation on the instrument blade tip acted as a mechanical stop when trying to close blades. This damage can contribute to poor cut performance. Engineering concluded the damage was likely due to mishandling. An additional finding not reported was pad printing removal and deep scratches on main tube. The tube extension had a small area where material was removed. The distal end of the main tube had various scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded this damage was likely due to mishandling. An electrical continuity test passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; pad printing removal and or tube abrasions with light material removal, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3154772 | 3125918 | 3022452 | 2955842201301794 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-25 | 2013-05-23 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-13 | Malfunction |  | It was reported that prior to starting da vinci si surgical procedure the tips were not opening on a pk dissecting forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on a is3000 system and driven. Recognition and engagement passed. The instrument moved intuitively with full a range of motion in all directions. The grips opened and closed properly. Functional performance testing did not find any issues. Additional observations not reported by site were frayed cables and tube damage. The pitch down cable was frayed at the distal clevis hub. The frayed strands stuck out at the instruments wrist. Other cables at the wrist were not damaged. The distal end of main tube had various scratch marks with light material removal. The scratches were. 060 -. 090 in length and not aligned with the tube axis. Engineering concluded the damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; a frayed cable and tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3156386 | 3127543 | 3024078 | 2955842201301795 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-25 | 2013-05-23 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-02 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the customer reported that the fenestrated bipolar forceps instrument could not coagulate and it also exhibited a broken wire. No patient harm, adverse outcome or injury was reported. No missing or fallen pieces were reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was able to cauterize, but one drive cable was found to be broken. The conductor wires were not visibly damaged. Instrument passed electrical continuity test, which is an indicator of cautery function. The pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were scratches are 0. 070 - 0. 180 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3156511 | 3127664 | 3024199 | 2955842201301796 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-24 | 2013-05-23 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-17 | Malfunction |  | It was reported that prior to starting a da vinci si total benign hysterectomy procedure the tips were out of alignment and did not meet on a fenestrated bipolar forceps instrument. The instrument was replaced with new forceps. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip was bent, causing side to side misalignment of the instrument grips. There was an offset at the tips. The bent grip had separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the instrument tip. Engineering concluded the damage was likely due to mishandling. An electrical continuity test passed. Additional observation found the pitch cable frayed at the distal clevis hub. No damage was found at the clevis. Frayed segment was approximately 0. 05 in length. No other cable damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; a frayed cable, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3156546 | 3127672 | 3024207 | 2955842201301797 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-25 | 2013-05-23 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-20 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure a wire was sticking out of the wrist on a pk dissecting forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal clevis hub. Cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The surface of the distal pulley had scratches. An electrical continuity test passed. Additional finding was the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3156613 | 3127738 | 3024273 | 2955842201301798 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-25 | 2013-05-23 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
420093 | 2013-04-24 | Malfunction |  | It was reported that during a da vinci si cystectomy procedure, instruments were being removed for undocking at the end of the case, the distal tip of the prograsp instrument deformed/broke in arm 2. As a result, the instrument could not be removed from the cannula. (b)(6), the first assistant, undocked arm 2 from the cannula. She then reportedly angled the cannula under direct vision in order to remove the instrument and cannula from the patient port site. The instrument was inspected and it did not appear to have any missing pieces from the shaft. Dr. (b)(6) determined there to be no harm to patient. Due to the fact that the prograsp instrument had no additional uses left, (b)(4), the board runner, made the decision to discard it. The customer was advised to check the cannula with a gauge pin to ensure that it was not damaged as part of the instrument removal. Upon follow up, i was told that it was not damaged. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3156684 | 3127805 | 3024340 | 2955842201301799 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-04-24 | 2013-05-23 | No | Yes | Health Professional | 420093 | No | No | 2013-04-24 | No | No | Reuse | 2013
42019408 | 2013-04-15 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the customer noted a broken cable near the tip of the mega needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a loose grip cable located at the distal idler. The grip cable was found to be broken at clamming pulley area inside housing. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3156695 | 3127812 | 3024347 | 2955842201301800 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-24 | 2013-05-23 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-10-09 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure the maryland bipolar forceps instrument was noted to have a broken steelwire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that both pitch cables were broken at the distal clevis hub. The cable segments that contained the crimp were still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. Engineering also found the distal end of the main tube had various scratch marks with light material removal. The scratches were. 070 -. 185 in length and not aligned with the tube axis. The evidence was inconclusive, but the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | Reuse | 3156804 | 3127887 | 2819860 | 2955842201301801 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-05-29 | 2013-05-23 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-05-29 | No | No |  | 2013
42000606 | 2012-05-08 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure the large needle driver instrument was noted to have a broken steel wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the grip cable was broken and stuck out at the wrist. The idler pulley exhibited rim and face damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | Reuse | 3156809 | 3127893 | 2819637 | 2955842201301802 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-05-29 | 2013-05-23 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-05-29 | No | No |  | 2013
42009308 | 2012-09-04 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si cystectomy procedure the prograsp forceps instrument was noted to have a broken steel wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the wrist. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | Reuse | 3156811 | 3127894 | 2819850 | 2955842201301803 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-05-29 | 2013-05-23 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-29 | Yes | No |  | 2013
42034301 | 2013-04-26 | Malfunction |  | It was reported that during a da vinci si lobectomy procedure pieces of coating on the thoracic grasper instrument were noted to be chipped off. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the main tube insulation appeared to have scratches and gouge marks, with materials removed at the distal end. The evidence not conclusive, but damage was likely caused by excess force contact on the main tube surface. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3156820 | 3127905 | 3025015 | 2955842201301804 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-26 | 2013-05-23 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-20 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted a cable had come out from track on the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was missing from the clevis. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3156931 | 3128023 | 3024578 | 2955842201301805 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-25 | 2013-05-23 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-03-14 | Malfunction |  | It was reported that during a da vinci si nissen fundoplication procedure, the customer noted that the endowrist permanent cautery hook instrument did not work. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3157018 | 3128096 | 3024631 | 2955842201301806 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-25 | 2013-05-23 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-20 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, a wire was sticking out of the wrist of a pk dissecting forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, one grip close cable was derailed at the distal idler pulley. The grips can still open and close, but movement may not be precise. Additional damage found was deep scratches on the main tube. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage may be due to mishandling. Electrical continuity passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3157112 | 3128177 | 3024273 | 2955842201301807 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-25 | 2013-05-23 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42000604 | 2013-04-01 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the articulating motion was stuck on a large needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The grips could open/close when input discs were manually rotated, however, the cable tension was lower than usual due to a failure of one of the distal end components. Additional observation not reported by site was a broken proximal clevis. The instrument was broken at the proximal clevis to main tube interface. The clevis was dislodged from the main tube as a result. Both the proximal clevis and the main tube were missing pieces. Engineering concluded the breakage was likely due to overloading at the tip. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; a dislodged clevis, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157250 | 3128341 | 3024876 | 2955842201301808 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-23 | No | Yes | Health Professional | 42000604 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-04-25 | Malfunction |  | It was reported that during reprocessing that the insulation on the permanent cautery spatula instrument was cracked. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the ceramic sleeve was broken at the distal end. The small diameter section at the distal end of the sleeve was missing a triangular shaped piece; the footprint measured roughly. 085 x. 085. The larger diameter section of the sleeve exhibited various scratch marks below the broken section. The evidence was not conclusive, but damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3157334 | 3128393 | 3024931 | 2955842201301809 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-26 | 2013-05-23 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-17 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the prograsp forceps instrument grinds when it rotates. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on a is3000 system and driven. Recognition and engagement test passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Functional performance testing did not find any issues. Additional observation not reported by site was tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 055 -. 095 in length and not aligned with the tube axis. Engineering concluded the damage may be caused by mishandling/misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157353 | 3128424 | 3024959 | 2955842201301810 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-23 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-04-26 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure pieces of coating were chipping off a thoracic grasper instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The main tube insulation appears to have scratches and gouge marks, with materials removed at the distal end. Engineering concluded the damage was likely caused by excess force contact on the main tube surface. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157418 | 3128477 | 3025015 | 2955842201301812 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-26 | 2013-05-23 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-04-18 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy, a cable was sticking out from the distal end of a large suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. Engineering also had additional findings. There was wearing on the pin, indicating the cable likely became derailed and was rubbing against the pin to wearing out the inner edge, eventually leading to the cable breakage. There were also visible scratches on the surface of the pulley. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3157421 | 3128480 | 3025018 | 2955842201301813 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-23 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-18 | Malfunction |  | It was reported that after a da vinci si total benign hysterectomy procedure a broken cable was identified on a maryland bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, the instrument was found with a frayed pitch cable at distal clevis. No damage was found on the pulley and clevis. Additional damage found was a corroded back idler pulley. On the surface of each pulley exhibited a yellowish material. Engineering concluded the damage was likely due to improper cleaning. Electrical continuity passed. No other damage was found. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3152608 | 3123474 | 3020029 | 2955842201301770 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-23 | 2013-05-22 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-23 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-15 | Malfunction |  | It was reported that during reprocessing a cable on the mega suturecut needle driver instrument was found loose from the wheel. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, the instrument grip cable was derailed off the distal idler pulley. The pulley exhibited some wear but was not damaged. Additional damage found were deep scratch marks exhibiting light material removal and a rough surface finish on the distal end of the main tube. Two indentations were found on instrument blade, one at the midpoint and the other close to the tip of the blade. Engineering concluded the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal ,found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3152609 | 3123475 | 3020030 | 2955842201301769 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-23 | 2013-05-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-23 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-18 | Malfunction |  | It was reported that after a da vinci si total benign hysterectomy procedure, a broken cable was identified on a large needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, the instrument was found with a frayed cable at the grip pulley on the distal clevis. The frayed cable section was approximately 0. 05 in length. No damage was found on pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3152611 | 3123477 | 3020032 | 2955842201301771 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-23 | 2013-05-22 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-23 | Device Not Retu | No | Reuse | 2013
A51P8 | 2013-05-01 | Other | Other | Event Description: --/-- The site contacted intuitive surgical inc. (isi) technical service engineer (tse) on (b)(6) 2013 alleging a loss of both video signals. The reporter at the site mentioned that the video signal was ok during the system start. During troubleshooting, it was discovered that there was a defective digitizer. The power led at the backside of the digitizer was off. Tse had the reporter reseat all the power cords and related cables; however, issue persisted. Due to the alleged issue the surgeon made a decision to convert the procedure to laparoscopic surgery. Reporter contacted isi again later on the same day and mentioned that the surgeon decided to convert the procedure to open surgery instead of laparoscopic surgery. Reporter also mentioned that they replaced the camera head cable and now the vision works well; however, mentioned that the digitizer power led is still not working. Based on the alleged issue the field service engineer (fse) replaced the digitizer. | Manufacturer Narrative: --/-- The video digitizer was returned to isi for investigation on (b)(4) 2013. The digitizer was evaluated with the following findings: failure analysis (fa) was unable to replicate the initial issue that was reported by the site. The digitizer powered up with good video on both eyes. The subject digitizer passed the following tests: video checked on three stations (vsc, ssc, psc), focus in and out, composite and sdi, tile pro, and power cycle. The led on the back of the unit was pushed in and fa repaired the led in-house. --/-- Additional Manufacturer Narrative: --/-- The device has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Intuitive surgical inc. Sent follow up questions regarding the alleged issue and obtained the following additional information: a male patient who was (b)(6) was going in for a da vinci s thymectomy procedure for thymom with myasthenia gravis on (b)(6) 2013. Reporter mentioned that the da vinci s was tested and set up the evening before the procedure and no error message occurred during set up. On (b)(6) 2013 prior to the procedure, reporter mentioned that they started the da vinci s and tested the system. Approximately 45 minutes, the video screen turned black; however, the screen of the console bar showed the color bar. The da vinci s was not docked to the patient at any time and no port incisions were placed on the patient. Due to the alleged issue, the physician decided to convert the procedure to open. The patient was under anesthesia for about 30-45 minutes prior to the physician's decision to convert the procedure to open. The patient did not experience any complications during the open procedure. The patient is still admitted in the hospital and his status is normal. The video digitizer unit is being returned back to isi for investigation. | IS2000 DAVINCI S SYSTEM | Reuse | 3153291 | 3124379 | 3020934 | 2955842201301772 | 1 | NAY | Manufacturer | ForeignHealth Professional | Other | InitialFollowup | 2013-05-01 | 2013-05-22 | Yes | Yes | Health Professional | A51P8 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No |  | 2013
42017910 | 2013-04-25 | Malfunction |  | It was reported that the monopolar curved scissors instrument allegedly broke (specifics were not provided) during a da vinci si surgical procedure. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed observed a broken tube extension and a missing 0. 290 x 0. 160 piece at the proximal clevis interface. The clevis was dislodged from tube extension. The tube extension fractured next to one of the keys that mates with the proximal clevis. Engineering also observed the distal end of the main tube had various scratch marks with light material removal. Scratches were 0. 180 - 0. 205 in length and were not aligned with the tube axis. The tube extension had a couple of sections with the orange pad printing removed. Intact pad printing did not flake off when scraped with fingernail. Evidence not conclusive, but tube extension likely broke due to excessive side loading or other mishandling/misuse. Evidence not conclusive, but main tube damage may have been caused by mishandling/misuse. No other damage found. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3153468 | 3124485 | 3021040 | 2955842201301773 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-14 | 2013-05-22 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-14 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-05-03 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken pulling cable on the prograsp forceps instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. Additional observation that were not initially reported by the customer was main tube damage. The distal end of the main tube had various scratch marks with light material removal. Scratches were 0. 075 - 0. 175 in length and were not aligned with the tube axis. Evidence not conclusive, but damage may have been caused by mishandling/misuse. No other damage was found. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3153504 | 3124494 | 3021049 | 2955842201301774 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-22 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure, the pk dissecting forceps instrument stopped working and not burning a few minutes after the surgeon started dissecting. The instrument was undocked and it was noted that a small wire came out. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3153760 | 3124881 | 3021466 | 2955842201301775 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-23 | 2013-05-22 | No | Yes | Health Professional | 42022703 | No | No | 2013-04-23 | No | No | Unkown | 2013
42022703 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure the pk dissecting forceps instrument was not burning and it was noted to have a small wire sticking out. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3153776 | 3124911 | 3021466 | 2955842201301767 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-23 | 2013-05-22 | No | Yes | Health Professional | 42022703 | No | No | 2013-04-23 | No | No | Unkown | 2013
42017910 | 2013-04-19 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the monopolar curved scissors (mcs) instrument accessory tip was crumbled. The instrument was removed from the patient and it was noted that the orange cover was cracked. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that a broken tube extension. The tube extension was broken and missing a piece at the distal end. The clevis was dislodged from the tube extension. The tube extension likely broke due to excessive side loading or other type of mishandling. The instrument passed electrical continuity testing. Engineering also found damage on the distal end of the main tube. Material was missing. The surface of the main tube appeared to be scraped off. The damages were found in various areas approximately 0. 75 away from the distal end of the main tube. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event, in a fragment-causing failure (such as a cable failure, a crimp separating from a cable etc. ) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument; however, the damage to the main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3153830 | 3124942 | 3021497 | 2955842201301776 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-23 | 2013-05-22 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-23 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-04-24 | Malfunction |  | It was reported that during a da vinci s surgical procedure the monopolar cautery instrument cover of the tip was noted to be burnt. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of the cover of the tip of the instrument was burnt. No disposable hook/spatula tip was returned with the instrument. Visual inspection showed no char marks or localized melting on the monopolar adapter. The instrument passed electrical continuity testing. Engineering also found was tube damage. The black tube insulation was damaged at the distal end. The insulation was gouged on one side and had a. 375 x. 110 piece missing roughly. 700 above the snake wrist. No other damage was found. Evidence is not conclusive; however, the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3153843 | 3124950 | 3021505 | 2955842201301777 | 1 | GCJ | Manufacturer | DistributorHealth Professional | Other | Initial | 2013-04-24 | 2013-05-22 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42018907 | 2013-02-12 | Malfunction |  | It was reported that during a da vinci s distal pancreatectomy procedure the double fenestrated grasper instrument was noted to have a broken wire; one of the jaws did not move. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of broken wire and one of the jaws did not move. The drive cables were not broken at the wrist, but pitch cables were loose. The grips opened and closed when inputs were manually rotated. Engineering also found a derailed backend cable and main tube damage. The housing was removed to find one pitch cable derailed from the back idler pulleys. The cable was wedged between two pulleys and had lost tension. The clamping pulley screws were still tight. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 300 in length and not aligned with the tube axis. Engineering concluded that the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3153862 | 3124968 | 3021523 | 2955842201301778 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-24 | 2013-05-22 | No | Yes | Health Professional | 42018907 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-23 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, a mega suturecut needle driver instrument cables frayed at the distal end. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken cable. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. There were also scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3153914 | 3125006 | 3021561 | 2955842201301779 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-24 | 2013-05-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
41032204 | 2013-04-24 | Malfunction |  | It was reported that during a da vinci si low anterior resection procedure the endo wrist one vessel sealer instrument was noted to have broken wires. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the cable was broken at the snake wrist, between the 2nd and 3rd wrist disc. Some physical damage was found on the wrist disc. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | ENDO WRIST ONE VESSEL SEALER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3153926 | 3125030 | 3021585 | 2955842201301780 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-24 | 2013-05-22 | No | Yes | Health Professional | 41032204 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-25 | Malfunction |  | It was reported that during a da vinci s hysterectomy bso bilateral salpingo-oophorectomy procedure the mega suturecut needle driver instrument was noted to have a frayed cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3153972 | 3125054 | 3021609 | 2955842201301782 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-05-06 | Malfunction |  | During a da vinci si hysterectomy procedure, the wires on the mega suture cut needle driver instrument reportedly came out. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a broken grip close cable at the distal idlers and found the cable segment sticking out at the instrument's wrist. The idler pulley spun freely. Other cables at the instrument's wrist were not damaged. An additional observation that was not initially reported by the customer was a damaged pulley. The distal idler pulley, which has the broken cable seated on, had a section that was bent inwards. Evidence not conclusive, but the pulley damage was likely due to mishandling and likely contributed to cable breakage. No other damage was found. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3153940 | 3125076 | 3021631 | 2955842201301781 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-06 | 2013-05-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-05-07 | Malfunction |  | During central processing, the user facility identified a frayed cable from the large needle driver instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a frayed pitch cable at the instrument's wrist. The frayed segments were various in lengths. No damage was found at the clevis. No other cable damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3153975 | 3125106 | 3021661 | 2955842201301783 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-07 | 2013-05-22 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-05-07 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-16 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, there is a wire which sticks out the jaws of the fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with a frayed pitch cable at the distal end. No damage was found on pulley and clevis. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3154044 | 3125192 | 3021747 | 2955842201301784 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-24 | 2013-05-22 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-04-24 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, wires broke on a small grasping retractor instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed wires broken on the instrument. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit excessive damage. An additional observation found the instrument main tube with deep scratches at the distal end of the main tube, approximately 1. 05 from the base of the proximal clevis. Engineering concluded the damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3154309 | 3125417 | 3021972 | 2955842201301785 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-24 | 2013-05-22 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-05-07 | Malfunction |  | During a da vinci si hysterectomy procedure, the wires at the end of the pk dissecting forceps allegedly broke. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a broken conductor wire at the yaw pulley exit. The conductor wire was detached from its connection at the grip. Electrical continuity testing was performed and failed. The yaw pulley showed no signs of arcing. Additional observations that were not initially reported by the customer were pulley damages. The yaw pulley was missing a 0. 161 x 0. 053 piece opposite to the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. There were indentations at the edge of the distal pulley. Also visible scratches on the surface of the pulley. Engineering noted that the pulley damage most likely occurred due to mishandling. The instruments and accessories instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3154397 | 3125505 | 3022060 | 2955842201301786 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-07 | 2013-05-22 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-05-07 | Device Not Retu | No | Reuse | 2013
42805912 | 2013-05-07 | Malfunction |  | During a da vinci si cholecystectomy procedure, the crocodile grasper instrument allegedly bent halfway up the shaft and part of the sheathing was scraped. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering did not confirm the alleged bent shaft complaint. The subject instrument is a single-site instrument and the shaft is intended to be flexible. However, engineering confirmed that the distal end of main tube had various scratch marks with lifted materials but no missing piece. The scratches were short in length and not aligned with the tube axis, suggesting that the scratches may have been caused by instrument collisions or rough handling. No other damage was found. The instruments & accessories instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | CROCODILE GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3154515 | 3125596 | 3022151 | 2955842201301787 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-07 | 2013-05-22 | No | Yes | Health Professional | 42805912 | Device Returned To Manufacturer | Yes | 2013-05-07 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-15 | Malfunction |  | It was reported that during reprocessing a monopolar curved scissors instrument, the scissors were not opening or closing properly. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with the tube extension cracked with no material removed at the proximal clevis interface. The clevis was dislodged from the tube extension. Engineering concluded the tube extension likely broke due to excessive side loading or other mishandling. Electrical continuity passed. An additional observation found the instrument with multiple hairline cracks at the distal end of the main tube. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the hairline cracks, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3154566 | 3125665 | 3022220 | 2955842201301788 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-24 | 2013-05-22 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-21 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure the large needle driver instrument's movements were difficult. It was noted that the instrument cable was damaged. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3154568 | 3125667 | 3022222 | 2955842201301768 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-23 | 2013-05-22 | No | Yes | Health Professional | 42000606 | No | No | 2013-04-23 | Device Not Retu | No | Unkown | 2013
42017910 | 2013-04-23 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the housing came apart on a monopolar curved scissors instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. For clarification, the tube extension was cracked near the axial key at the proximal clevis. Based on this clarification, the complaint is confirmed. The clevis was not dislodged from tube extension. Engineering concluded the tube extension likely broke due to excessive side loading or other mishandling. Additional damage found was deep scratches on the main tube near the housing. The proximal end of the main tube had two scratch marks exhibiting light material removal and a rough surface finish. The scratches were short in length approximately. 0545 and. 0355. Engineering concluded the damage may be due to mishandling. Electrical continuity passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3154577 | 3125675 | 3007763 | 2955842201301789 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-24 | 2013-05-22 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-03-18 | Malfunction |  | It was reported that during a da vinci s surgical procedure a grasping retractor instrument had a broken separated cable. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. There were also scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3154594 | 3125691 | 3022246 | 2955842201301790 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-24 | 2013-05-22 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42004906 | 2013-04-25 | Malfunction |  | It was reported that during a da vinci si surgical procedure the cadiere forceps instrument was noted to have cable damage. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the wrist. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering also found main tube damage on the distal end. Material was missing. The surface of the main tube appeared to be scraped off. Damages can be found in various areas approximately 2. 5 away from the distal end of the main tube. The evidence was not conclusive, but damage was likely due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3154610 | 3125724 | 3022279 | 2955842201301791 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-25 | 2013-05-22 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-19 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure, an exposed wire was found on a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the base of the grip jaws and scissors. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3154743 | 3125845 | 3022400 | 2955842201301792 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-24 | 2013-05-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
| 2013-03-28 | Malfunction |  | Before using the maryland bipolar forceps (ref 420172) for the robotic cholecystectomy with the davinci, the scrub tech noticed that the strings were broken. There are uses left on the instrument to receive reimbursement. Device #1is this a laboratory device or laboratory test?no. |  | MARYLAND BIPOLAR FORCEPS | SYSTEM SURGICAL COMPUTER CONTR | 3193295 | 3164067 | 3060604 | 3164067 | 1 | NAY | User Facility |  |  | Initial | 2013-05-22 | 2013-05-22 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
| 2013-05-03 | Malfunction |  | none | none | FORCEPS BIPOLAR PRECISE | SYSTEM SURGICAL COMPUTER CONTR | 3193410 | 3164173 | 3060710 | 3164173 | 1 | NAY | User Facility |  |  | Initial | 2013-05-22 | 2013-05-22 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
| 2013-05-07 | Malfunction |  | During robotic case, the mega suturecut needle driver was not functioning properly. The instrument was immediately removed from the surgical field and replaced. Device #1is this a laboratory device or laboratory test?no. |  | MEGA SUTURECUT | SYSTEM SURGICAL COMPUTER CONTR | 3193441 | 3164183 | 3060740 | 3164183 | 1 | NAY | User Facility |  |  | Initial | 2013-05-22 | 2013-05-22 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
42018306 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure the console surgeon complained of difficulty in movement of the permanent cautery hook instrument at the start of the procedure. The scrub nurse removed the instrument for checking and found that the string of the instrument was cut so didn't use the instrument further there were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3150251 | 3121489 | 3018044 | 2955842201301762 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-23 | 2013-05-21 | No | Yes | Health Professional | 42018306 | No | No | 2013-04-23 | No | No | Reuse | 2013
42030901 | 2013-04-09 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure a cable broke on a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. There were also scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3150267 | 3121506 | 3018061 | 2955842201301763 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-23 | 2013-05-21 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-23 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-03-15 | Injury | Required Intervention | Event Description: --/-- It was reported that the tip of the monopolar curved scissors broke off during a prostatectomy procedure and fell into the patient. The broken tip was retrieved and there was no allegation of patient injury due to the alleged issue. The surgery was completed as planned. | Manufacturer Narrative: --/-- Product was replaced and is being requested back for investigation. The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Intuitive surgical inc. (isi) attempted to contact the reporter several times via email and has been unsuccessful in getting a hold of the reporter to answer any additional questions. No further clinical information was provided. The complaint is being reported based on the initial information that was provided to intuitive surgical inc. --/-- Additional Manufacturer Narrative: --/-- (b)(4). On (b)(6) 2013, the reporter contacted isi and mentioned that fragments did not fall into the patient as initially indicated in the initial report that was submitted to the fda. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2013, the tube extension of the mcs instrument was returned and evaluated. It was noted that the tube extension was broken and missing a. 350 x. 275 piece at the proximal clevis interface. The clevis was dislodged from the tube extension. The tube extension was found to be fractured next to one of the keys that mates with the proximal clevis. Engineering evaluation concluded that the damage was likely due to excessive side loading or other mishandling/misuse. Electrical continuity testing of the instrument passed. No other damage was found on the device. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HOT SHEARS TM MONOPOLAR CUR | Reuse | 3151083 | 3122201 | 3018756 | 2955842201301764 | 1 | NAY | Manufacturer | ForeignUser facility | Other | FollowupFollowup | 2013-04-26 | 2013-05-21 | Yes | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No |  | 2013
42019408 | 2013-04-29 | Malfunction |  | Event Description: --/-- Prior to starting a da vinci si surgical procedure, the user facility identified a broken cable on the mega needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. . --/-- Additional Manufacturer Narrative: --/-- Device evaluation: intuitive surgical, inc. Received the instrument involved with the complaint on (b)(4) 2013. Engineering confirmed there was a frayed grip close cable at the distal idler pulley. The frayed strands were sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | Reuse | 3151983 | 3122842 | 3019397 | 2955842201301765 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-04-29 | 2013-05-21 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-06-03 | Device Not Retu | No |  | 2013
42000707 | 2013-04-26 | Malfunction |  | Event Description: --/-- During a da vinci si surgical procedure, the user facility reportedly identified broken wires on the round tip scissors instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- Follow-up 1-device evaluation: intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed there was a broken pitch cable at the idler block and the clamping pulley. Additional observations noted that were not reported by the customer were corrosion damage. Both the cable and the clamping pulley were found with corrosion and residue from the cleaning agent. Evidence not conclusive, but corrosion damage likely due to mishandling during cleaning. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | ROUND TIP SCISSORS INSTRUMENT | Reuse | 3152033 | 3122891 | 3019446 | 2955842201301766 | 1 | NAY | Manufacturer | OtherUser facility | Other | InitialFollowup | 2013-04-29 | 2013-05-21 | No | Yes | Health Professional | 42000707 | Device Returned To Manufacturer | No | 2013-05-30 | Device Not Retu | No |  | 2013
420179 | 2013-05-10 | Injury | Disability | Robotic arm #1 with monopolar curved scissors jumped forward and nicked a vessel during surgical intervention. Clips were successfully applied to control the bleeding vessel. An add'l monopolar curved scissors jammed during this procedure. Reason for use: laparoscopic robotic supracervical hysterectomy. |  | DA VINCI S MONOPOLAR CURVED SCISSORS | DA VINCI S MONOPOLAR CURVED SC | 3166244 | 3135353 | 3031892 | MW5030313 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-05-21 | 2013-05-21 | No | Yes | Health Professional | 420179 | Yes | Yes |  |  |  |  | 2013
| 2013-04-01 | Injury | Hospitalization | Pt initially had robot assisted laparoscopic prostatectomy (b)(6) 2013. Readmitted with cellulitis and free air in abdomen (b)(6) 2013. After treatment discharged. Second readmission (b)(6) 2013 with ct positive for fistulous connection between sigmoid colon and deep pelvic abscess. During exploratory laparotomy, rectal perforation identified requiring colostomy creation. Vendor has notified facility of failure of monopolar curved scissors used during prostatectomy. Surgeon believes instrument contributed to rectal perforation. Vendor notification, urgent medical device notification - (b)(4). Intuitive surgical endowrist instrument hot shearsz monopolar curved scissors affected product: part numbers description (b)(4), 8mm monopolar curved scissors - a. K. A. Hot shears - 8 mm monopolar curved scissors - a. K. A. Hot shears - dear da vinci customer, this notification is to inform you that intuitive surgical has identified a potential issue with certain versions of its hot shearsz monopolar curved scissors mcs instrument and to inform you of precautionary steps you should take to ensure the safety of pts, while using your current inventory of mcs instruments. Certain (b)(4) versions of the mcs instruments may develop micro-cracks near the distal scissor end of the shaft following reprocessing. This may create a pathway for electrosurgical energy to leak to tissue during use and potentially cause thermal injury. The affected area is shown and is confined to an approx 1 cm section of the shaft, as (b)(4). |  | MONOPOLAR CURVED SCISSOR | SCISSOR | 3165857 | 3135356 | 3031895 | MW5030317 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-05-21 | 2013-05-21 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2013
VS3000 | 2013-04-17 | Death | Death | Intraoperative death secondary to cardiovascular arrest secondary to surgical bleeding during radical nephrectomy. |  | DA VINCI SURGICAL SYSTEM | DA VINCI | 3172927 | 3142082 | 3038617 | 3142082 | 1 | NAY | User Facility | Unknown | RISK MANAGER | Initial | 2013-05-17 | 2013-05-21 | Yes | No | Health Professional | VS3000 | Yes | Yes |  |  |  |  | 2013
IS2000 A51P8 | 2013-05-13 | Other | Other | Event Description: --/-- On (b)(6) 2013, the surgeon contacted the field service engineer (fse), who then contacted isi technical support for a non-recoverable error message 30253 on their da vinci s (is2000) system during a surgical procedure. An error 30253 means that the amps didn't turn on due to the failure reaction logic (frl) being triggered. The logs on the system were reviewed onsite and the field service engineer (fse) noticed a long list of error 256 (which means the emergency stop switch was pressed, this is where the supervisor recovered the system from a recoverable fault), error 10025 (means that amps don't enable due to estop)and error 283 (means estop is pressed). The surgeon confirmed that he only pressed the e-stop once by mistake at the beginning. The site rebooted the system; however, rebooting led to the following errors on the manufacturer safety data sheet (msds), 30253 and 213 (means the motorized arm did not complete its startup sequence , the system faults and transitions to a safe state). Isi technical support recommended the field service engineer (fse) to have the site emergency power off (epo) the surgeon side cart (ssc). Field service engineer (fse) did that a few times and the issue was not resolved. It was recommended that the site move the master tool manipulator (mtms) and restart the system. This was done and the system came up ok. The surgeon called back later the same day with the same issue and had a few 10025 errors on the screen that he was able to recover; however, then it became hard to recover from the error message. Isi technical support advised the surgeon to restart the system and doing that triggered error 30253 again. Surgeon was advised to emergency power off (epo) the console twice and restart normally a third time, but that was impossible due to the error message. Isi technical support then advised the site to cycle the e-stop button several times and moved the master tool manipulator (mtms) around. They tried one last time to emergency power off (epo) but the system came back with the error 30253. Due to the alleged error messages, the surgeon decided to convert the da vinci surgical procedure to open procedure and would postpone his da vinci procedure he had later that afternoon. Intuitive surgical inc. (isi) is ordering a gusp (gemini user switch panel), iop (i/o processor) and system control processor board (scpz) for the site. | Manufacturer Narrative: --/-- Isi regulatory complaint analyst attempted to get in contact with the site to obtain additional information; however, has been unsuccessful to get in contact with the reporter at this time. --/-- Additional Manufacturer Narrative: --/-- On (b)(6) 2013, the gusp (gemini user switch panel)was returned to isi for investigation and was evaluated by engineering with the following findings: engineering was unable to reproduce the alleged failure report by installing this board into pca system and run x50 power cycles. The unit was working fine without any issues. Engineering observed the power button was sticky and the power cover was broken. | DA VINCI S SURGICAL SYSTEM | Reuse | 3149461 | 3120747 | 3017302 | 2955842201301759 | 1 | NAY | Manufacturer | OtherUser facility | BIOMEDICAL ENGINEER | InitialFollowup | 2013-05-13 | 2013-05-20 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2013-05-13 | Device Not Retu | No |  | 2013
42017910 | 2013-04-08 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the orange insulator portion on the monopolar curved scissors (mcs) instrument had a piece missing. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension was found to be broken exhibiting a piece removed at the proximal clevis interface. The clevis was dislodged from tube extension. Engineering concluded that the tube extension possible broke due to excessive side loading. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3149595 | 3120820 | 3017375 | 2955842201301760 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-22 | 2013-05-20 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-08 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the tips did not align on a pk dissecting forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip was bent, causing a side to side misalignment of the grips. There was a. 015 offset at the instruments tip. The bent grip had a small separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. Additional observation not reported by site was tube damage. The distal end of main tube had various scratch marks with light material removal. The scratches were. 090 -. 200 in length and not aligned with the tube axis. Engineering concluded the damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3149980 | 3121205 | 2993422 | 2955842201301761 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-23 | 2013-05-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-23 | Device Not Retu | No | Reuse | 2013
410322 | 2013-05-20 | Malfunction |  | During robotic procedure, vessel sealer used for one bite then faulted stating blade exposed when it was not. Unable to fix error thus vessel sealer removed from the pt and the field. No harm to pt. |  | ENDOWRIST ONE VESSEL SEALER | VESSEL SEALER | 3166344 | 3135365 | 3031900 | MW5030310 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-05-20 | 2013-05-20 | No | Yes | Health Professional | 410322 | Yes | Yes |  |  |  |  | 2013
IS3000 A60P5 | 2013-05-01 | Other | Other | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure, the surgeon experienced a loss of image through the left eye of the high resolution stereo viewer (hrsv) on the surgeon side cart (ssc). Initially, the surgical staff indicated that there were no icons visible on the hrsv screen, and the image was good on the touchscreen left input. With instructions from an intuitive surgical inc. (isi) technical support engineer (tse), the surgical staff emergency powered off the surgeon side console (ssc) and reseated the blue optic cable. The alleged vision issue was not resolved. The surgeon then reported that the image was lost on the touchscreen left input but the icons were there. The tse advised the surgical staff to try substituting with a spare vision cable; however, the surgical staff indicated that they did not have a replacement cable. The tse reportedly informed the surgical staff of the ability to set up the da vinci si system for 2d. However, the surgeon made the decision to abort the da vinci si surgical procedure after anesthesia had already been administered and port incisions were created. | Manufacturer Narrative: --/-- On (b)(4) 2013, an intuitive surgical inc. (isi) field service engineer (fse) performed a field evaluation of the system. The fse resolved the reported vision issue by replacing the camera cable. The fse also swapped patient side manipulator (psm) 2 and 3. In addition, the fse found a battery issue from the patient side cart (psc). The battery reportedly failed a battery test. The fse indicated that the system would go down after 22 minutes. The battery was replaced. The system was then tested and found to be within specifications. The battery was returned and evaluated. Engineering evaluation was unable to replicate the reported battery issue found by the fse. The battery was installed on an in-house printed circuit assembly (pca) system. The system powered up at normal mode. The battery passed a system and power cycle test. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2013, isi obtained additional information from the initial reporter of this complaint regarding the reported event. The reporter indicated that the patient was under anesthesia for approximately 1 hour before the reported surgical procedure was aborted. During the time the patient was under anesthesia, the patient side cart (psc) was docked to the patient with the camera and cannulas inserted. The reporter denied that any instruments were inserted into the patient at any time before the surgical procedure was aborted. The reporter indicated that the patient did not experience any complications during the time he was under anesthesia. The reporter indicated that the patient did not develop any complications after the surgical procedure was aborted. He also indicated that the da vinci si prostatectomy procedure was rescheduled and the reported vision issue was resolved after the blue optic cable was replaced. | DA VINCI SI SURGICAL SYSTEM | Reuse | 3146851 | 3118783 | 3015338 | 2955842201301743 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-05-01 | 2013-05-19 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No |  | 2013
42017206 | 2013-04-10 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, it was noted that the arm connector broke off on the bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the bipolar insert located on the blue housing area was found to be detached from the back end. The connector pins and conductor wires had separated from each other from blue housing end. The insert was unable to fully sit in the chassis, as it could be pulled out if the bipolar cord was removed. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 3146846 | 3118799 | 3015354 | 2955842201301742 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-19 | 2013-05-19 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure the pk dissecting forceps instrument began acting independently and would not follow commands from the console. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument pitch cable was broken at the distal end, which was likely causing non-intuitive motion. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. Engineering also found a char mark on the top corner of the yaw pulley. The instrument passed electrical continuity testing. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3146907 | 3118859 | 3015414 | 2955842201301744 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-23 | 2013-05-19 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-23 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-19 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure a wire broke at the tip of a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported broken wire; however, for clarification the damaged instrument component was a frayed grip cable. The grip cable was frayed at the distal idler pulley. The frayed strands stuck out at the instruments wrist. Other cables at the wrist are not damaged. There were also scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3146974 | 3118895 | 3015450 | 2955842201301745 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-22 | 2013-05-19 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure a wire was broken on a fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a wire was broken. One conductor wire was broken at the yaw pulley exit. The wire was detached from its connection at the grip. An electrical continuity test failed. The yaw pulley showed no signs of arcing. Additional findings were the yaw pulley was missing a. 465 x. 059 piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. There was also an indentation at the edge of the distal pulley and some scratches on the surface. Engineering concluded it was most likely mishandling. The instrument was also found with a frayed pitch cable at distal idler. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the frayed cable and tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3147018 | 3118946 | 3005839 | 2955842201301747 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-22 | 2013-05-19 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-22 | Malfunction |  | It was reported that during reprocessing a cable broke on a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the broken cable. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. There were also scratches on the surface of the distal pulley. Additional findings were the distal end of main tube had a scratch mark showing light material removal. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3147074 | 3118979 | 3015534 | 2955842201301748 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-22 | 2013-05-19 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2013-04-17 | Death | Death | It was reported that during a da vinci si nephrectomy procedure, the patient's renal artery, aorta and vena cava were punctured and the patient subsequently expired. Reportedly, there was no malfunction of the da vinci surgical system. | On (b)(4) 2013, intuitive surgical, inc. (isi) received additional information from the initial reporter of this event. The initial reporter indicated that the patient's injuries were discovered during the surgical procedure and that repair of the affected vessels were not repaired as the patient expired. Intuitive surgical has contacted the site to obtain additional information concerning the reported event. As of the date of this report, no additional information has been provided by the site. On (b)(4) 2013, isi contacted the isi clinical sales representative (csr). The csr indicated that he spoke to the initial reporter of this report. The csr indicated that he was told by the initial reporter that the patient's renal artery was damaged during the application of staples. The csr was also told by the initial reporter that the surgeon was using an endowrist grasper instrument to grasp the patient's aorta and the camera had been removed. Reportedly, when the surgeon released the grasper instrument from the surgeon side console, and without direct visualization of the surgical site, damage to the patient's aorta occurred. The csr indicated that the initial reporter was unable to provide any other details concerning the events as they occurred. A review of the site's system logs found no system errors occurred that would have caused or contributed to the reported event. Based on the limited information provided, it is unknown what caused the patient's demise. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3147088 | 3119000 | 3015555 | 2955842201301749 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-18 | 2013-05-19 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-04-18 | No | No | Reuse | 2013
| 2012-02-14 | Injury | Required InterventionLife Threatening | Event Description: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) received voluntary report (b)(4) with the following event description: on (b)(6) 2012, i had a da vinci assisted lavh, being released the following day on (b)(6) 2012. I felt as though i was recovering well for the first few days. Began having abdominal pain on (b)(6) 2012. Returned to surgeon on (b)(6) 2012 because i was not feeling well. Put on antibiotic and sent home. On (b)(6), went to er and was admitted because it was discovered that my right ureter had suffered a thermal injury during my hysterectomy and my continued decline the days prior were due to the injured ureter leaking urine into my abdominal cavity. After two failed attempts at treating the injury conservatively, i had to be opened up, reoperated on, and my bladder reconstructed to reach my now shorter ureter. | Manufacturer Narrative: --/-- Intuitive surgical has made several attempts to contact the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. --/-- Additional Manufacturer Narrative: --/-- Corrected data can be found in event date = (b)(6) 2012. | DA VINCI SURGICAL SYSTEM | Unkown | 3147166 | 3119090 | 3015645 | 2955842201301750 | 1 | NAY | Manufacturer | Other | Other | InitialFollowup | 2013-04-22 | 2013-05-19 | Yes | No | Health Professional |  | No | No |  | Device Not Retu | No |  | 2013
IS3000 A60P5 | 2013-04-24 | Other | Other | It was reported that during a da vinci si distal pancreatectomy procedure, the patient side manipulator (psm) 2 arm insertion axis appeared to move downward. Reportedly, the system was docked to the patient and a harmonic shears instrument was installed on the psm when the issue occurred. No patient harm or injury was reported. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. | The investigation conducted by the field service engineer, determined that the issue experienced by the site was associated with the brake within the patient side manipulator (psm) arm 2. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering was unable to replicate the failure mode experienced by the site. Functional testing of the psm on an in-house is3000 test system found that the psm passed friction and brake testing. Engineering restarted the test system in normal mode and performed multiple power cycles while moving the psm in different positions. No brake failure occurred and there were no system errors generated. The psm also passed an advanced brake test. As a precaution, axis 1 and axis 2 of the psm was replaced. Engineering evaluation also found that the axis 2 produced a low reading. The axis 2 harmonic drive was replaced to repair the affected axis. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3147198 | 3119140 | 3015695 | 2955842201301751 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-24 | 2013-05-19 | Yes | Yes | Health Professional | IS3000 A60P5 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-08 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument orange sleeve was noted to be cracked. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was cracked, broken and missing a piece at the proximal clevis interface. The missing piece measured about 0. 237 x 0. 153. The clevis was dislodged from the tube extension. The tube extension likely broke due to excessive side loading. The instrument passed electrical continuity testing. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage was likely due to mishandling. No other damage was found. There was a crack on the main tube; it measured about. 302. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event, in a fragment-causing failure (such as a cable failure, a crimp separating from a cable etc. ) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3147201 | 3119148 | 3015703 | 2955842201301753 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-25 | 2013-05-19 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-02 | Malfunction |  | It was reported that during a da vinci si partial nephrectomy procedure the monopolar curved scissors (mcs) instrument were noted to be dull. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the scissors were dull. Engineering placed the instrument on an in-house system for a latex cut test. The scissors did not cut cleanly through. 006 latex. The latex was snagged at the scissor tips. The blade edges were undamaged, but the edges exhibited wear at the tips, negatively affecting cut performance. Additional damage found was pitting on the instrument blade. On the one blade tip, there was evidence of pitting (small holes) on the blade surface. Engineering concluded the damage was likely due to improper cleaning. Engineering also found pad printing removal on the tube extension and a bent banana plug. The tube extension exhibited material removal on the distal end. The banana plug was bent at the back of the chassis. The instrument passed electrical continuity testing. No other damage was found. The cleaning and sterilization user manual specifically states: o when scrubbing the tip of cautery instruments, take care not to damage the insulation. O do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. O prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event, in a fragment-causing failure (such as a cable failure, a crimp separating from a cable etc. ) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument; however, the pad printing was removed, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3147231 | 3119151 | 3015706 | 2955842201301756 | 1 | NAY | Manufacturer | DistributorHealth Professional | Nurse | Initial | 2013-04-25 | 2013-05-19 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42017909 | 2011-02-28 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the monopolar curved scissors (mcs) instrument stopped responding. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering placed the instrument on an in-house is3000 system and it was driven. Recognition and engagement testing passed. The instrument moved intuitively with full range of motion in all directions. The blades opened and closed properly. Engineering found poor cut performance and main tube damage. The scissors did not cut cleanly through. 006 latex. The latex was snagged at the scissor tips. The blade edges were undamaged, but the edges exhibited wear at the tips, negatively affecting cut performance. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 120 -. 330 in length and not aligned with the tube axis. Engineering concluded that the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3147203 | 3119152 | 3015707 | 2955842201301754 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-04-25 | 2013-05-19 | No | Yes | Health Professional | 42017909 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
40018012 | 2013-04-18 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure, the surgeon staff alleged that arcing was observed with the monopolar curved scissors (mcs) tip cover accessory used in conjunction with the mcs instrument. The mcs tip cover accessory was reportedly discarded. There were no missing or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The mcs tip cover accessory was not returned to isi for evaluation, as the initial reporter indicated that the tip cover was discarded, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, intuitive surgical inc. (isi) contacted the initial reporter of this complaint and obtained the following additional information: the surgical procedure was not recorded. The mcs instrument, mcs tip cover accessory, and cannulas were inspected prior to the start of the surgical procedure and nothing unusual was observed. The reporter indicated that the mcs tip cover accessory was installed correctly with the installation tool and electro lube was applied to the instrument prior to the start of the procedure. The reporter claimed that the scissors sparked (or arced) to vag cuff tissue, anterior and posterior. The reporter also indicated that no injury was observed during the procedure. After the reported arcing was observed, the mcs instrument was removed, wiped off, and inspected by the surgical staff. The surgical staff reportedly did not observe anything unusual and the mcs instrument was reinserted. According to the reporter, the grounding pad was correctly positioned during the procedure. According to the reporter, the site uses erbe for their electrosurgical generator unit (esu) with 5-65 settings. It is unknown what exact esu settings were used during the reported event. The surgical procedure was reportedly completed as planned and the mcs tip cover accessory was discarded by the surgical staff. The surgical staff indicated that the mcs tip cover accessory appeared to be intact and normal prior to removal from the mcs instrument. The reporter indicated that the tip was not saved as they didn't think about it being damaged. The reporter indicated that the patient did not experience any post-operative complications. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3147233 | 3119155 | 3015710 | 2955842201301752 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-19 | 2013-05-19 | No | Yes | Health Professional | 40018012 | No | No | 2013-04-19 | No | Yes | Unkown | 2013
42020505 | 2013-04-25 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical fenestrated bipolar forceps instrument was noted to have frayed wires. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was frayed at the distal clevis. No damage found on pulley and the clevis. Engineering also found deep scratches on the main tube. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded that the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3147267 | 3119205 | 2987170 | 2955842201301757 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-25 | 2013-05-19 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument was noted to be dull. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering placed the instrument on an in-house is3000 system for a latex cut test. The scissors cut cleanly through. 006 latex. The blade edges were undamaged. Engineering also found a crack in the main tube. The distal end of the main tube exhibited a hairline crack in the axial direction. The tube outer diameter had a slight bulge near where the crack was located, roughly. 300 above the metal reinforcement ring. The instrument was disassembled to check the inner surface of the main tube. No cracking was found on the inner surface; it appeared the crack initiated on the outer surface. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3147300 | 3119215 | 3015770 | 2955842201301755 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-25 | 2013-05-19 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-21 | Malfunction |  | On (b)(6) 2013 intuitive surgical recieved medwatch report uf/importer report (b)(4) with the following event description: during operation mega suturecut needle driver wire broke at tip of instrument. Needle driver removed from sterile field. What was the original intended procedure? robotic assisted hysterectomy, bso, sacrocolpopexy and tvt. Device usage problem: device failed (e. G. Broke ,couldn't get it to work or stopped working). | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3147398 | 3119325 | 3015880 | 2955842201301758 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-19 | 2013-05-19 | No | Yes | Health Professional | 42030901 | No | No | 2013-03-19 | No | No | Unkown | 2013
42017910 | 2013-04-24 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the monopolar curved scissors instrument (mcs) was noted to be cracked just below the orange piece. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was broken at the proximal clevis interface. No missing pieces were found but broken a piece was dislodged from the clevis. Engineering concluded that the tube extension was likely due to excessive side loading or other mishandling. Engineering also found cracks on the main tube. Two hairline cracks parallel to each other were found on the distal end near the tube extension. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event, in a fragment-causing failure (such as a cable failure, a crimp separating from a cable etc. ) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3145251 | 3117339 | 3013894 | 2955842201301724 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-24 | 2013-05-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-24 | Device Not Retu | No | Reuse | 2013
40009308 | 2013-04-03 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the prograsp forceps instrument didn't work. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering placed the instrument on an in-house is1200 system and it was driven. Recognition and engagement testing passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The functional performance testing did not find any issues. Engineering also found tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 - 1 in length and some were aligned with the main tube axis. Engineering concluded that the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3145263 | 3117340 | 3013895 | 2955842201301726 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-23 | 2013-05-17 | No | Yes | Health Professional | 40009308 | Device Returned To Manufacturer | No | 2013-04-23 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-24 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical maryland bipolar forceps instrument black insulation was noted to be worn in a spot. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the main tube exhibited a large crack that might pose as an insulation risk at the proximal end. Engineering concluded that the damage was likely due to mishandling. Engineering also found deep scratches on the main tube. The distal end of main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded that the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3145252 | 3117361 | 3013916 | 2955842201301725 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-24 | 2013-05-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-22 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure the maryland bipolar forceps instrument was noted to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. Engineering also found tube damage and corroded clamping pulleys. The main tube exhibited wear with light material removal and a rough surface finish in various areas along the tube length. The housing was removed to find excessive white residue and light pitting corrosion on all clamping pulleys. The backend components exhibited a brown discoloration. Engineering concluded that the tube and clamping pulley damage was likely due to improper cleaning processes. No other damage was found. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3145264 | 3117362 | 3013917 | 2955842201301727 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-23 | 2013-05-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-23 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-01 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the customer noted that the monopolar curved scissors (mcs) instrument were dull. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on in-house da vinci robot system to perform a latex cut test. The mcs instrument cut cleanly through a 0. 006 latex. The blades were not damaged. Additional observation not reported by site was main tube damage. Engineering found a crack near the distal end of the main tube. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3145615 | 3117759 | 3014314 | 2955842201301728 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-17 | 2013-05-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-02 | Malfunction |  | The mega suture cut instrument was returned without an initial complaint from the customer. | The instrument was returned and evaluated. The mega suture cut needle driver instrument was returned without an initial complaint from the customer. The instrument was returned with a derailed grip cable, blade damage and deep scratches on the main tube. One grip cable slipped off distal pulley, but the cables were still in tension. Engineering also observed that the instrument blade exhibited two indentations along blade edge. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3145639 | 3117847 | 3014402 | 2955842201301730 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-17 | 2013-05-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-11 | Malfunction |  | It was reported that during reprocessing a mega suturecut needle driver instrument the tips open to wide and gets stuck. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, one blade edge had an indentation at tip that acted as a mechanical stop for the other blade when closing. Engineering concluded the damage to blade was likely due to mishandling. Additional damage found were deep scratches on the main tube. The distal end of the main tube had various scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal ,found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3145703 | 3117848 | 3014403 | 2955842201301729 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-22 | 2013-05-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-04 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a frayed cable on the pk dissecting forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be frayed at the distal end. Instrument passed electrical continuity test. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3145716 | 3117915 | 3014470 | 2955842201301731 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-17 | 2013-05-17 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-11 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was noted the monopolar curved scissors (mcs) instrument was not being registered by the da vinci robot system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The mcs instrument was placed on in-house da vinci robot system and driven with no problem. Instrument passed recognition, engagement and electrical continuity tests. The pogo pins did not stick and were not contaminated. Additional observations not reported by site were main tube damage, pad printing removed and kinked flushed tubing. Engineering observed a small crack near the distal end of the main tube. Several marks were also found on tube extension pad printing area showing some orange label material removed. A kink was also observed on the flush tube that is located inside the blue housing. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3145832 | 3118014 | 2952177 | 2955842201301732 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-17 | 2013-05-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-29 | Malfunction |  | It was reported that prior to starting a da vinci si prostatectomy procedure a monopolar curved scissors instrument had a problem about angulation as a wire was out. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed an angulation wire was out. For clarification, the pitch cable was broken at the distal end causing a lack of proper instrument wrist movement. The cable segment that contains the crimp was missing from the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the wrist were not damaged. Additional damage found were pad printing removal from the tube extension and a deep scratch on the main tube. The tube extension exhibited multiple spots missing pad printing. The distal end of the main tube had various scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage may be due to mishandling. Electrical continuity passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3145958 | 3118245 | 3014800 | 2955842201301733 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-22 | 2013-05-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-04-18 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted a a broken cable on the mega needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at the wrist. Other cables at wrist were not damaged. Engineering also observed scratches on the surface of the distal pulley. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3145959 | 3118249 | 3014804 | 2955842201301734 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-18 | 2013-05-17 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer reported that the fenestrated bipolar forceps instrument was not rotating properly. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Instrument passed electrical continuity test. Additional observations not reported by site were main tube damage and pin connector damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. The connector pin located on the blue housing was also found to be bent upward. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3146034 | 3118297 | 3014852 | 2955842201301735 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-19 | Malfunction |  | It was reported that during a da vinci surgical procedure, the customer reported that the fenestrated bipolar forceps instrument was broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated Per the customer reported complaint engineering observed that one grip was found to be bent causing side to side misalignment of grips There was a 0 155 offset at the tips indicating overloading at the tip Instrument passed electrical continuity test Additional observation not reported by site was main tube damage The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish The scratches were short in length and were not aligned with the tube axis Engineering concluded that damage was likely due to mishandling The customer reported complaint does not itself constitute a reportable event however the reported malfunction if to reoccur could cause or contribute to an adverse event The endowrist instruments instructions for use ifu specifically states general precautions and warnings handle instruments with care Avoid mechanical shock or stress that can cause damage to the instruments | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3146230 | 3118504 | 3015059 | 2955842201301739 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-19 | 2013-05-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-21 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a broken cable on the monopolar curved scissors (mcs) instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at wrist. Instrument passed electrical continuity test. Engineering also observed a frayed pitch cable. The pitch cable was found to be frayed at the distal clevis, but no excess damage was found. Additional observations not reported by site were main tube damage and pin connector damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. The connector pin located on the blue housing was also found to be bent upward. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3146493 | 3118598 | 3015183 | 2955842201301736 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-18 | 2013-05-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
40017910 | 2013-02-12 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure, the customer noted that the monopolar curved scissors (mcs) instrument were unable to cut anymore. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a broken grip cable making it incapable of cutting. One grip cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at wrist. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. Engineering found a crack near the distal end of the main tube. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3146511 | 3118618 | 3015173 | 2955842201301737 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-18 | 2013-05-17 | No | Yes | Health Professional | 40017910 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-04-19 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the black sheath was peeling off from the thoracic grasper instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the main tube insulation appeared to have scratches and gouge marks, with materials removed at the distal end. Engineering concluded that damage was likely caused by excess force contact on the main tube surface. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3146556 | 3118684 | 3015259 | 2955842201301738 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-19 | 2013-05-17 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-04-19 | Malfunction |  | It was reported that during a da vinci s surgical procedure a wire sprung on a tenaculum forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3146685 | 3118719 | 3015294 | 2955842201301741 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-19 | 2013-05-17 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2010-11-22 | Injury | Other | Event Description: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint with the following allegation related to a da vinci si hysterectomy procedure performed on (b)(6) 2010: the legal complaint alleged that on (b)(6) 2010, the da vinci surgical system malfunctioned, and as a result of the reported issue, the patient allegedly suffered grievous and permanent personal injuries, pain, suffering, and loss of enjoyment of life, aggravation, exacerbation and worsening of pre-existing conditions, and incurred pecuniary damages, medical, hospital, rehabilitative and other expenses, and loss of ability to work. | Manufacturer Narrative: --/-- Intuitive surgical inc. (isi) contacted the risk management department and director of insurance of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. Isi has reviewed the system logs with a date of (b)(6) 2010, the date the reported da vinci si hysterectomy procedure was performed. Based on the review of the system logs, there is no evidence that a system error occurred during the surgical procedure. --/-- Additional Manufacturer Narrative: --/-- Intuitive surgical inc. (isi) contacted the risk management department and director of insurance of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. Isi has reviewed the system logs with a date of (b)(6) 2010, the date the reported da vinci si hysterectomy procedure was performed. Based on the review of the system logs, there is no evidence that a system error occurred during the surgical procedure. | DA VINCI SI SURGICAL SYSTEM | Reuse | 3146651 | 3118728 | 3015253 | 2955842201301740 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-04-23 | 2013-05-17 | Yes | Yes | Health Professional | IS3000 A60P8 | No | No | 2013-04-23 | No | No |  | 2013
IS2000 A51P8 | 2013-04-22 | Other |  | It was reported that prior to starting a da vinci s prostatectomy procedure, after powering on the system the site experienced system error code # 802. Reportedly, the site shut down and restarted the system; however, the system error code # 23 appeared. The site cleaned their system fiber cables and restarted the system; however, system error code # 802 recurred. Unable to resolve the issue, the site contacted isi for technical support assistance. An isi technical support engineer determined that the system error code # 23 experienced by the site was associated with the system's red fiber cable. The tse instructed the site to clean the cable and restart the system; however, system error code # 802 re-appeared. The tse then instructed the site to try a different power outlet; however, the issue persisted. The pt was under anesthesia and the port incisions had not yet been created when the surgeon made the decision to abort the planned surgical procedure. | The ppd was returned and evaluated. Engineering eval was unable to replicate the alleged "err 802 pt side cart battery backup errors" after installing the board into an in-house pca system. The module battery box was returned and evaluated. Engineering eval was able to replicate the reported issue by installing the battery box into an in-house pca system. The module battery box failed at power on with a message displayed on screen. The site's logs were reviewed during the investigation conducted by the field service engineer (fse). The fse found that the site also experienced system error code #5, system error code #23 and system error code #802 experienced by the site were associated with the module battery box, the red fiber cable, the auxiliary power board (apb) and the pt side power distribution board (ppd). The module battery box provides 24 volt battery backup. The fiber cable connects the pt side cart to the surgeon console and passes the video, audio and data during system operation. The apb controls the battery backup system, cart safety interlock and converts the 24 volt battery to 48 volts during ac power failure. The pt side power distribution board is a printed circuit (pca) board located on the pt side cart that divides the electrical power feed into subsidiary circuits while providing a protective circuit breaker for each circuit. The system was repaired by replacing the affected module battery box, red fiber cable, the apb and the ppd. As of the date of this report, no add'l info concerning the reported incident has been provided to isi by the site. Intuitive surgical has made several attempts to contact the site to obtain add'l info concerning the reported event; however, no add'l info has been provided as of the date of this report. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3160289 | 3122675 | 3019230 | 2955842201301746 | 1 | NAY | Manufacturer | ForeignUser facility | NOT APPLICABLE | Initial | 2013-05-17 | 2013-05-17 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No |  | Yes | No | Reuse | 2013
42030901 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy a cable broke on a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a cable was broken. The grip cable was broken at the distal idlers. The cable segment stuck out at the instruments wrist. The idler pulley spun freely and was not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER IINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142790 | 3115075 | 3011630 | 2955842201301695 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-17 | Malfunction |  | It was reported that during reprocessing, the pk dissecting forceps instrument had a frayed wire in a loop. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the frayed wire, in a loop. The instrument was found with a frayed pitch cable at the distal clevis. No damage was found on the pulley or clevis. The frayed strand was approximately 0. 0455 in length. Additional damage found were deep scratches on the main tube. The distal end of the main tube had two scratch marks exhibiting light material removal and a rough surface finish. The scratches were short in length approximately. 1155 and. 0655. Engineering concluded the damage may be due to mishandling. Electrical continuity passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142803 | 3115085 | 3011640 | 2955842201301691 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-16 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-18 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, a broken cable was identified on a fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the broken cable. The pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instruments wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142805 | 3115086 | 3011641 | 2955842201301692 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-16 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-18 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure a broken cable was identified on a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken cable. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. Additional findings were scratches on the surface of the distal pulley. The distal end of the main tube also had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142806 | 3115087 | 3011642 | 2955842201301693 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-17 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure a broken cable wire was found on a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken cable wire. One grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142808 | 3115089 | 3011644 | 2955842201301694 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the cord snapped off a small grasping retractor instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the cord snapped. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. There were also scratches on the surface of the distal pulley. An additional finding was the distal end of the main tube had a scratch mark light with material removal. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3143957 | 3116088 | 3012643 | 2955842201301696 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-18 | 2013-05-16 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-15 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, while suturing, the customer noted a broken cable on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that both grip cables were found to be broken at the wrist. The cable segments stuck out at wrist. The idler pulleys were able to spin freely and exhibited pulley rim and face damage. Instrument has zero uses left. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3144555 | 3116701 | 3013256 | 2955842201301697 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-17 | 2013-05-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-02 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the monopolar curved scissors instrument was dull and would not cut tissue. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on a system for a latex cut test. The scissors did not cut cleanly through (b)(4) latex. The latex snagged at the scissor tips. The blade edges were damaged and circular shaped dents were found, negatively affecting cut performance. Electrical continuity passed. Additional observation found a hairline crack at the distal end of the main tube. The customer reported complaint does not itself constitute a mdr reportable event; however; cracks on the main tube, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3144664 | 3116750 | 3013305 | 2955842201301699 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-19 | 2013-05-16 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
400180 | 2013-04-16 | Malfunction |  | It was reported that during a da vinci s partial nephrectomy procedure, when the surgical staff attempted to remove the monopolar curved scissors (mcs) instrument with the mcs tip cover accessory installed from the patient, the mcs instrument became stuck inside of the cannula. Reportedly, the mcs instrument and cannula were removed simultaneously to allow for removal of the instrument from the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The tip cover accessory and monopolar curved scissors (mcs) instrument have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, intuitive surgical contacted the initial reporter regarding the reported event. The initial reporter indicated that inspection of the monopolar curved scissors (mcs) instrument by the surgical staff found that the mcs tip cover accessory had been over installed onto the mcs instrument, causing the mcs instrument to become stuck during removal of the instrument from the cannula. The initial reporter indicated that it is unknown if the tip cover accessory and mcs instrument will be returned to isi failure analysis evaluation as the site made the decision to retain the devices to use for training purposes at the site to prevent future occurrences over installation of the tip cover onto the mcs instrument. The initial reporter was unable to provide the version number of the affected tip cover accessory. The instruments and accessories user manual indicates that the tip cover accessory is not properly installed on the mcs instrument if the tip cover is installed beyond the orange surface and over the shaft of the instrument. This causes a bulge on the instrument shaft and may prevent it fitting through the cannula. The instruments and accessories user manual specifically states: warning: failure to install the tip cover accessory properly may result in: improper scissor opening; tip cover accessory falling off; electrical arcs and alternate site burns the customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 3144645 | 3116751 | 3013306 | 2955842201301698 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-16 | 2013-05-16 | No | Yes | Health Professional | 400180 | No | No | 2013-04-17 | No | No | Unkown | 2013
42018306 | 2013-04-17 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure, the surgeon staff indicated that the tip of the permanent cautery hook instrument broke off and fell into the patient. The broken fragment was retrieved and there was no patient harm, adverse outcome, or injury reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation confirmed the alleged issue that the hook of the instrument came out. The instrument was returned with the hook detached from the distal end. The hook shank was found to be fractured within the yaw pulley at the cross section of the hole feature. The instrument's ceramic sleeve was found to be still intact. Engineering evaluation also found the instrument with a broken proximal clevis. One proximal clevis ear was observed to be broken and missing a piece at the proximal pin. The other clevis ear was found to be cracked along the base of the ear, but no missing pieces were observed. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, intuitive surgical inc. (isi) contacted the initial reporter of this complaint and obtained the following information: the reporter indicated that the instrument was inspected prior to use and no disinfectants were used prior to cleaning of the device. She also denied that the site has recently changed any of their cleaning/sterilization processes. The reporter indicated that the procedure was not recorded. She explained that the hook of the permanent cautery hook instrument got caught on one of the tip openings of a maryland bipolar forceps instrument and then broke off. The fragment was reportedly retrieved through a laparoscopic accessory port using a laparoscopic grasper instrument. The reporter denied that an x-ray or ultrasound was used to help find the broken fragment. The reporter indicated that the patient did not experience any intra-operative complications because of the alleged issue. The reporter also denied that the patient underwent any post-operative tests or secondary procedures because of the alleged issue. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3144681 | 3116786 | 3013341 | 2955842201301701 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-17 | 2013-05-16 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-04 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the scissors were dull on the monopolar curved scissors instrument. The planned surgical procedure was completed with a replacement instrument of the same type and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on system for a latex cut test. The scissors cut cleanly through. 006 of latex. The blades were not damaged. An electrical continuity test passed. An additional observation found the main tube with hairline cracks at the distal end. The customer reported complaint does not itself constitute a mdr reportable event; however; hairline cracks on the main tube, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3144711 | 3116804 | 3013359 | 2955842201301700 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-19 | 2013-05-16 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-15 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the customer noted that the orange sleeve expanded and kinked. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the mcs instrument tube extension was cracked at the proximal clevis interface. The clevis was found to be dislodged from tube extension. Engineering concluded that the tube extension cracked due to excessive side loading. Instrument passed electrical continuity test. Engineering also found scratches on the surface of the distal pulley. Additional observation not reported by site was main tube damage. Engineering found a crack on the main tube close to the extension tube side. The crack measured roughly about 0. 508. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3144714 | 3116806 | 3013361 | 2955842201301702 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-17 | 2013-05-16 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-11 | Malfunction |  | It was reported that during a da vinci si surgical procedure the maryland bipolar forceps instrument was noted to have wires protruding out at the wrist area. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3144749 | 3116836 | 3013391 | 2955842201301705 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-16 | No | Yes | Health Professional | 42017207 | No | No | 2013-04-18 | No | No | Unkown | 2013
42017207 | 2013-04-12 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the maryland bipolar forceps instrument was noted to have a frayed separated cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. The distal end of the main tube had a scratch mark showing light material removal. The scratches were short in length and not axially aligned with the main tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3144777 | 3116837 | 3013392 | 2955842201301706 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-16 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-04-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the tennaculum forceps instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering placed the instrument on an in-house system. The system successfully recognized the instrument. The pogo pins were not stuck or contaminated. Engineering also found a broken pitch cable. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was missing in the clevis. The clevis did not exhibit any damage or wear marks. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3144755 | 3116838 | 3013393 | 2955842201301708 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-19 | 2013-05-16 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-03-25 | Malfunction |  | It was reported that prior to starting a da vinci si low anterior resection a permanent cautery hook instrument pole was bent for cautery connection. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The banana plug was bent to one side. A monopolar cord could still be installed, but would not seat flush against the chassis. Electrical continuity test passed. Additional observation not reported was the distal end of the main tube has various scratch marks with light material removal. The scratches were. 105 -. 165 in length and not aligned with the tube axis. Engineering concluded the damage to the banana plug and tube damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3144773 | 3116853 | 3013408 | 2955842201301703 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-19 | 2013-05-16 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-04-18 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the tenaculum forceps instrument cable popped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3144774 | 3116855 | 3013410 | 2955842201301704 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-18 | 2013-05-16 | No | Yes | Health Professional | 42020706 | No | No | 2013-04-18 | No | No | Unkown | 2013
42018907 | 2013-04-18 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure the double fenestrated grasper instrument cables broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the grip cable was frayed at the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at the wrist were undamaged. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3144752 | 3116860 | 3013415 | 2955842201301707 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-16 | No | Yes | Health Professional | 42018907 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-09 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical monopolar curved scissors (mcs) instrument was noted to have wires exposed by the tip of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found multiple hairline cracks near the distal end of the main tube. The customer reported complaint does not itself constitute a mdr reportable event, in a fragment-causing failure (such as a cable failure, a crimp separating from a cable etc. ) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument; however, the hairline cracks, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3144753 | 3116861 | 3004532 | 2955842201301709 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-19 | 2013-05-16 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-11 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical prograsp forceps instrument tip was noted to be split apart. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the damaged instrument component was a frayed pitch cable. The instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3144756 | 3116865 | 3013420 | 2955842201301710 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-19 | 2013-05-16 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-19 | Malfunction |  | It was reported that during reprocessing the da vinci si radical prostatectomy fenestrated bipolar forceps instrument was noted to have a torn cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was at the distal clevis. Engineering also found bent grips, deep scratches, and pogo pins sticking. The one grip was bent, causing side-to-side misalignment of grips. There was a. 0220 offset at the tips. Engineering concluded that damage was likely due to mishandling. The distal end of the main tube had multiple scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. The one pogo pin was found stuck at the base. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3144780 | 3116866 | 3013421 | 2955842201301711 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-22 | 2013-05-16 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-22 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical pk dissecting forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the cable was frayed. The pitch down cable was frayed at the distal clevis hub. The frayed strands stuck out at the wrist. The pitch up cable was loose, but not damaged. Engineering also found a derailed backend cable and main tube damage. The housing was removed to find one pitch cable was derailed from the back idler pulleys. The cable was wedged between two pulleys and had lost tension. The backend derailment likely contributed to the fraying at the wrist. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 075 -. 160 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3144782 | 3116870 | 3013425 | 2955842201301712 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-22 | 2013-05-16 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-26 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility reportedly identified broken wires at the tip of the mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering confirmed a frayed grip cable at the distal idler pulley. The frayed strands were sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. Additional observations that were not initially reported by the customer were derailed cable and pulley damage. One grip close cable was derailed at the distal idler pulley. The grips was still able to open and close but movement may not be precise. The cable derailment was likely due to cable losing contact with pulley during wrist articulation. The fleet angle between proximal and distal pulleys may contribute to derailment. There were also scratches on the surface of the distal pulley. Engineering concluded that the scratches on the pulley were most likely due to tge derailed grip cable. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3144787 | 3116878 | 3011348 | 2955842201301713 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-26 | 2013-05-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-04-16 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy a wire broke a mega needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken cable. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. There were also scratches on the surface of the distal pulley. An additional finding were various scratch marks showing light material at the distal end of the main tube. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3144949 | 3117025 | 3013600 | 2955842201301714 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-19 | 2013-05-16 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-19 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure a large needle driver instrument had wires unraveled at the tip. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. However for clarification, the nonconformance was a broken grip cable. The broken grip cable stuck out at the instruments wrist. The idler pulley spun freely but exhibited pulley rim and face damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3144986 | 3117063 | 3013638 | 2955842201301715 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-19 | 2013-05-16 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-26 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility reportedly identified broken wires at the tip of the mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a frayed grip cable at the distal idler pulley. The frayed strands were sticking out at the instrument's wrist. Other cables at instrument's wrist were not damaged. Additional observations that were not initially reported by the customer were a derailed cable, scratches on the pulley and scratches on the main tube. The same frayed grip cable was also derailed at the distal idler pulley. Grips was still able to open and close but movement may not be precise. The cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. The fleet angle between proximal and distal pulleys may have contributed to derailment. There were scratches on the surface of the distal pulley with the derailed grip cable on top of it. The distal end of main tube had a scratch mark showing light material removal. The scratch was short in length and was not axially aligned with the tube. Evidence not conclusive, but damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings -handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3144995 | 3117072 | 3011348 | 2955842201301716 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-26 | 2013-05-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-25 | Malfunction |  | During a da vinci si surgical procedure, a cable from the fenestrated bipolar instrument allegedly broke. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp was missing. The clevis does not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3145051 | 3117128 | 3013703 | 2955842201301718 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-30 | 2013-05-16 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-22 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical pk dissecting forceps instrument was noted to have a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUME | ENDOSCOPIC ELECTROSURGICAL INS | 3145053 | 3117130 | 3013705 | 2955842201301717 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-22 | 2013-05-16 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-22 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical maryland bipolar forceps instrument was noted to have a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was derailed and severely frayed at the distal idler pulley. Frayed strands stuck out at the wrist. The other cables at wrist are not damaged. Engineering also found a dislodged luer plate and bipolar insert, disconnected conductor wire, and damaged main tube. The luer plate was dislodged from the chassis and pushed into the backend. A section of the chassis inner wall feature that holds luer plate was broken off. The bipolar insert and pins were missing from the back end. The conductor wires were sticking out of the chassis. The insert and pins may have been detached when bipolar cord was removed. The main tube exhibited wear with material removal and a rough surface finish throughout its entire length. The epoxy tube coating has been removed to expose rough fibers below. Engineering concluded that the luer plate and tube damage were likely cleaning related issues. No other damage was found. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3145057 | 3117134 | 3013709 | 2955842201301720 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-22 | 2013-05-16 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-19 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the mega suturecut needle driver instrument wire was noted to be broken and sticking out. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. Engineering also found an indentation on the edge of the distal pulley. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3145059 | 3117136 | 3013711 | 2955842201301719 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-22 | 2013-05-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-04-22 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical tenaculum forceps instrument was noted to have a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was missing from the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3145060 | 3117137 | 3013712 | 2955842201301721 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-22 | 2013-05-16 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42018906 | 2013-03-29 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy the double fenestrated grasper instrument would only move to one side. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on a system and driven. Recognition and engagement test passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The instrument was fully functional. Additionally engineering found deep scratches on the distal end of the main tube. Distal end of main tube has scratch marks exhibiting light material removal and a rough surface finish. Evidence not conclusive, but damage may be due to mishandling. Instrument has 6 uses remaining. Pic attached. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3145092 | 3117171 | 3013746 | 2955842201301722 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-20 | 2013-05-16 | No | Yes | Health Professional | 42018906 | Yes | No | 2013-04-20 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-11 | Malfunction |  | It was reported that during reprocessing the tip of the pk dissecting forceps instrument would not open wide enough. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on a system and driven. Recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The instrument was fully functional. Electrical continuity passed. Additional damage found were deep scratches on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; tube abrasions with material removal , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PK DISSECTING FORCEPS INSTRUMENET | ENDOSCOPIC ELECTROSURGICAL INS | 3145116 | 3117195 | 3013770 | 2955842201301723 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-22 | 2013-05-16 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
40022702 | 2013-02-01 | Malfunction |  | It was reported that during a da vinci surgical procedure the pk dissecting forceps instrument can not open anymore. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument does not open/close. One grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. Additional observation not reported by site was a broken conductor wire. One conductor wire was broken at the yaw pulley exit. The wire stuck out from the yaw pulley. No char marks were observed on the yaw pulley. Electrical continuity test failed. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3140582 | 3112988 | 3009544 | 2955842201301667 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 40022702 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
40017910 | 2013-03-06 | Malfunction |  | It was reported that during a da vinci cystocele repair procedure, during preparing tissues, no moves were possible anymore on the monopolar curved scissors instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument would not move. For clarification, the instrument had a broken grip cable making it incapable of proper movement. One grip cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. Additional damage found was pad printing removal on the tube extension. Electrical continuity passed. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3141493 | 3113814 | 3010370 | 2955842201301668 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 40017910 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
40022702 | 2013-03-26 | Malfunction |  | It was reported that during a da vinci cystectomy procedure, during coagulation of tissue, the pk dissecting forceps instrument doesn't want to close. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument does not close. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. The other cables at the wrist were not damaged. Electrical continuity passed. Additional findings not reported were scratches on the surface of the main tube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may have been due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3141675 | 3114000 | 3010556 | 2955842201301669 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 40022702 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-08 | Malfunction |  | It was reported that during reprocessing a maryland bipolar forceps instrument the jaws were not closing together properly and evenly. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the jaws were not closing together properly and evenly. One grip was bent, causing side to side misalignment of the instrument grips. There was a. 021 offset at the instrument tips. Engineering concluded the damage was likely due to mishandling. Additional damage not reported were a bent bipolar pin and deep scratches on main tube. The bottom bipolar pin was bent at the back of the chassis. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the additional findings may be due to mishandling. Electrical continuity test passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the material removal from the main tube, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142046 | 3114369 | 3010925 | 2955842201301670 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-08 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure a pk dissecting forceps instrument, when inserted, read not recognized. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument. The pogo pins did not stick and were not contaminated. Dcp verification of the instrument passed. The instrument was driven, moved intuitively, and the grips opened and closed. Additional findings were scratches on the surface of the distal pulley. The distal end of the main tube also had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the material removal on the main tube , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142134 | 3114419 | 2989442 | 2955842201301671 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-18 | Malfunction |  | It was reported that the pk dissecting forceps instrument was damaged. The reporter was unable to indicate when the damaged instrument was identified and provide specifics of the damage. No fallen pieces were reported. There was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering confirmed a broken pitch cable at the distal end of the instrument. The cable segment that contains the crimp was missing from clevis. Clevis did not exhibit any damage or wear marks. Other cables at instrument's wrist were not damaged. Additional observations not reported by the customer was a char mark at the distal end of the instrument. The yaw pulley exhibited charring and localized melting at the grip base. Electrical continuity testing was performed and passed. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142150 | 3114447 | 3011003 | 2955842201301672 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-03 | 2013-05-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-05-03 | Device Not Retu | No | Reuse | 2013
| 2013-04-01 | Injury | Other | It was reported that during a da vinci si prostatectomy procedure, the surgical staff alleged that an unspecified instrument could not be removed from a trocar. The surgeon reportedly attempted to use another unspecified instrument to straighten the instrument that could not be removed from the trocar. During the attempt, one of the instruments allegedly broke and fragments fell into the patient. According to the initial reporter of this complaint, the surgical procedure was completed robotically. The initial reporter did not know the status of the patient. | Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. At this time, it is unknown what exact instruments were involved in this complaint, if any interventions were taken to locate and/or retrieve the broken fragment(s) that fell into the patient, if the broken fragment(s) were retrieved, if the patient developed any complications because of the reported issue, and what the patient's current status is. | UNKNOWN INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142164 | 3114456 | 3011012 | 2955842201301673 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-15 | 2013-05-15 | Yes | Yes | Health Professional |  | No | No | 2013-04-15 | No |  | Unkown | 2013
42018405 | 2013-03-08 | Malfunction |  | It was reported that after a da vinci si sacrolpopexy procedure the blue housing came off permanent cautery spatula instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument housing had no damage and remained on the housing. An additional observation not reported was a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the broken pitch cable, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142218 | 3114494 | 2935676 | 2955842201301674 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-26 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility reportedly identified broken wires at the tip of the mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering confirmed that one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit any damage. The cable segment was sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142389 | 3114509 | 3011348 | 2955842201301675 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-26 | 2013-05-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-20 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure the maryland bipolar forceps instrument was noted to have reduced movement. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering placed the instrument on an in-house system and the instrument was driven. The instrument passed recognition and engagement testing. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The instrument was fully functional. The instrument passed electrical continuity testing. Engineering found scratches on the main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the main tube. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142328 | 3114511 | 3011066 | 2955842201301677 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-04 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the maryland bipolar forceps instrument was not recognized by the system. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering placed the instrument on an in-house system and the system successfully recognized the instrument. The pogo pins were not stuck or contaminated. The dcp (dallas chip programmer) verification of the instrument passed. The instrument passed electrical continuity testing. The instrument was driven and moved intuitively and the grips opened and closed. Engineering also found scratches on the main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. There were scratches on the surface of the distal pulley. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142397 | 3114669 | 3011225 | 2955842201301676 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42019007 | 2013-04-05 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical cobra grasper instrument was noted to have bent tips. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of grips. There was a. 2805 offset at the tips. Engineering concluded that the damage was likely due to mishandling. Engineering also found a deep scratch on the main tube. The distal end of the main tube had a scratch mark exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. Instrument had expired. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142401 | 3114672 | 3011228 | 2955842201301678 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42019007 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-26 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility reportedly identified broken wires at the tip of the mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering confirmed that one grip close cable was broken near the proximal pulleys and proximal clevis cable hole. The idler pulley spun freely and did not exhibit any damage. The cable segment was sticking out at wrist. No major wear was found on the proximal clevis cable hole. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142403 | 3114674 | 3011348 | 2955842201301679 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-26 | 2013-05-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-04-15 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the grasping retractor instrument was not recognized. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was checked for recognition on an in-house system. The system successfully recognized instrument. The pogo pins did not stick and were not contaminated. Dcp verification of the instrument passed. Additional damage found were various deep scratch marks exhibiting light material removal and a rough surface finish on the distal end of the main tube. Engineering concluded the damage may have been due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. This the customer reported complaint does not itself constitute a mdr reportable event; however; the main tube abrasions , found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142408 | 3114676 | 3011232 | 2955842201301680 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-26 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility reportedly identified broken wires at the tip of the mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering confirmed that one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment was sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. Additional observations that were not initially reported by the customer were scratches on the surface of the distal pulley. Engineering was unable to conclude how the scratches occurred. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142423 | 3114695 | 3011348 | 2955842201301681 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-26 | 2013-05-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-26 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility reportedly identified broken wires at the tip of the mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering confirmed that one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment was sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. Additional observations that were not initially reported by the customer were scratches on the surface of the distal pulley. Engineering was unable to conclude how the scratches occurred. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142437 | 3114729 | 3011348 | 2955842201301682 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-26 | 2013-05-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
420227 | 2013-04-16 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure that the pk dissecting forceps instrument would not fully engage on arm 1. It was noted the instrument had a broken cable and the instrument was replaced to continue the case. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The customer has disposed of the instrument and it will not be returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142529 | 3114792 | 3011347 | 2955842201301683 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-16 | 2013-05-15 | No | Yes | Health Professional | 420227 | No | No | 2013-04-16 | No | No | Reuse | 2013
42030901 | 2013-04-26 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility reportedly identified broken wires at the tip of the mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering confirmed that one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment was sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142530 | 3114793 | 3011348 | 2955842201301684 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-26 | 2013-05-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-22 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have a broken connector. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal end. The cable segment that contained the crimp was missing from the clevis. Engineering also found indentations on the distal clevis. Engineering concluded the damage was likely due to mishandling and the cause of the broken cable. The other cables at the wrist were undamaged. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142539 | 3114818 | 3011374 | 2955842201301685 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
40020504 | 2013-04-04 | Malfunction |  | It was reported that during a da vinci s partial nephrectomy procedure, the fenestrated bipolar forceps instrument bipolar cable was noted to be difficult to connect and disconnect, just like if electrodes were too big. The surgical staff was used to this instrument and took care of it the same way each time. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the bipolar cable was difficult to connect and disconnect from the instrument. Both the pin and insert was missing; most likely due to difficult to remove and broke off when pulling off with force. Pins were broken off from the conductor wires, therefore failed electronic continuity testing. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the main tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142540 | 3114822 | 3011377 | 2955842201301686 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 40020504 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-04-09 | Malfunction |  | It was reported that during a da vinci si surgical procedure the cable snapped and separated on the large suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. Scratches were on the surface of the distal pulley. The cable segment stuck out at the instruments wrist. The other cables at the wrist were not damaged. Additional findings were the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142589 | 3114864 | 3011419 | 2955842201301688 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure that a broken wire was sticking out of the fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was sticking out from the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity test passed. Additional findings were the distal end of main tube has various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142575 | 3114870 | 3011425 | 2955842201301687 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42004806 | 2013-04-18 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the long tip forceps instrument was noted to have a crack at the very tip of the joint. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the grip tips were found bent at the distal end. The one grip was bent, causing side-to-side misalignment of the grips. There was a. 077 offset at the tips. Engineering concluded that the damage was likely due to mishandling. Engineering also found damaged distal idler pulleys and deep scratches on the main tube. Both distal idler pulleys exhibited indentations along the edges. Engineering concluded that the damage was likely due to mishandling. The distal end of the main tube had a scratch mark exhibiting light material removal and a rough surface finish. The scratch mark was approximately. 1190 long. Engineering concluded that the damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | LONG TIP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3142595 | 3114886 | 3011441 | 2955842201301689 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42004806 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-05 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the monopolar curved scissors (mcs)instrument orange part was noted to be loose. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was cracked with a piece sticking out at the proximal clevis interface. Engineering concluded that the tube extension likely broke due to excessive side loading or other mishandling. The instrument passed electrical continuity testing. Engineering also found a hairline cracked at the distal end of the main tube. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3142576 | 3114887 | 3011442 | 2955842201301690 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-18 | 2013-05-15 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
| 2013-05-02 | Malfunction |  | Robot arm alarmed and disarmed during case. Error was not able to be cleared when they tried to resolve fault the arm was disabled. Surgery completed with 3 arms. |  | DA VINCI | ROBOT | 3152941 | 3121838 | 3018393 | MW5030212 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-05-15 | 2013-05-15 | No | Yes | Health Professional |  | Device Returned To Manufacturer | Yes |  |  |  |  | 2013
42027804 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci s surgical procedure that the cable broke on grasping retractor instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, one grip cable was frayed and not broken at the distal clevis. The frayed strands stuck out at the instruments wrist. The other cables at the wrist were not damaged. Additional damage found was deep scratch exhibiting light material removal and a rough surface finish on the distal end of the main tube. Engineering concluded the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3138439 | 3110496 | 3007051 | 2955842201301653 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-16 | 2013-05-14 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-04-16 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-22 | Malfunction |  | During a da vinci si surgical procedure, the monopolar curved scissors instrument's tips were reportedly not working properly. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering was unable to confirm the alleged complaint. The instrument was placed on an in house da vinci system and driven: recognition and engagement tests passed. The instrument moved intuitively with full range of motion in all directions. Blades opened and closed properly and successfully performed cut test. Instrument was fully functional. Additional observations not reported by the customer was a small hairline crack above the tube extension. Electrical continuity testing was performed and passed. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3139058 | 3111208 | 3007763 | 2955842201301651 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2013-05-08 | 2013-05-14 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-26 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken tip on the large needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering confirmed a broken grip cable by the instrument's tip area. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment was sticking out by the grip area. Other cables at wrist were not damaged. Additional observations not reported by the customer were indentations on the edge of the distal pulley. Engineering was unable to conclude how the pulley damage occurred, but most likely due to mishandling. There were also some visible scratches on the surface of the pulley. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3139538 | 3111801 | 3008356 | 2955842201301654 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-30 | 2013-05-14 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42805206 | 2013-04-19 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a hook that would not attach to the tip of the monopolar cautery instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering was able to replicate and confirm the alleged complaint. Engineering was not able to screw on an in house hook adapter because both electrical contacts inside distal end were bent inside. Evidence not conclusive, but damage is like due to excessive force or other type of mishandling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3139586 | 3111863 | 3008418 | 2955842201301655 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-30 | 2013-05-14 | No | Yes | Health Professional | 42805206 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-06 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the maryland bipolar instrument could not articulate the surgeon's movement. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on in-house da vinci robot system and driven with no problem. Instrument passed recognition, engagement and electrical continuity tests. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The instrument was found to be fully functional. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3139744 | 3112092 | 3008647 | 2955842201301656 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-15 | 2013-05-14 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-14 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the customer experienced difficulty while advancing the monopolar curved scissors (mcs) instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The mcs instrument was placed on in-house da vinci robot system and it was able to advance pass the cannula opening with no trouble. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3139822 | 3112267 | 3008822 | 2955842201301657 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-15 | 2013-05-14 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42805206 | 2013-04-16 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the monopolar cautery instrument copper leads were noted to be bent inside the distal tip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the electrical prong was found bent inside the cautery adapter, resulting in the cautery tip not fitting properly. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3140148 | 3112538 | 3009093 | 2955842201301658 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2013-04-16 | 2013-05-14 | No | Yes | Health Professional | 42805206 | Device Returned To Manufacturer | No | 2013-04-16 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-11 | Malfunction |  | It was reported that prior to starting a da vinci si gynecological surgical procedure the maryland bipolar forceps instrument was not recognized. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument, showing the instrument was expired. The pogo pins did not stick and were not contaminated. Additional findings found were an indentation on the distal idler pulley, deep scratches, and a bent bipolar pin. The distal idler pulley had a mechanical indentation at the edge. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. The bottom bipolar pin was bent at the back of the chassis. Engineering concluded these additional findings were likely due to mishandling. Electrical continuity passed. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; scratch marks on the main tube exhibiting light material removal, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3140214 | 3112684 | 3009239 | 2955842201301659 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-17 | 2013-05-14 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-03 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the orange sleeve was cracked on the monopolar curved scissors instrument and was unable to be removed from the trocar. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The complaint the orange sleeve was cracked was confirmed. For clarification, the tube extension was cracked at the proximal clevis interface. The clevis was dislodged from the tube extension as a result. Engineering concluded the tube extension likely broke due to excessive side loading or other mishandling. Additional damage found was deep scratches exhibiting light material removal and a rough surface finish on the distal end of the main tube. Engineering concluded the damage may be due to mishandling. Electrical continuity passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the deep scratches exhibiting light material removal, found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3140276 | 3112689 | 3000173 | 2955842201301660 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-16 | 2013-05-14 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-16 | Device Not Retu | No | Reuse | 2013
42020705 | 2013-04-15 | Malfunction |  | The tennaculum forceps instrument was returned, no further information was provided. | The instrument was returned and evaluated. Engineering found a broken pitch cable, corrosion and cracks on the clamping pulleys inside the housing, and deep scratches on the main tube. The pitch cable was broken at the distal end. The cable segment that contained the crimp was remained in the clevis but was stuck out. The clevis did not exhibit any damage or wear marks. Multiple clamping pulleys exhibited cracking and corrosion around the screw head. Engineering concluded that the failure was likely cleaning related. The distal end of the main tube had various scratch marks exhibited light material removal and a rough surface finish. Engineering concluded the damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3140279 | 3112691 | 3009246 | 2955842201301661 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-17 | 2013-05-14 | No | Yes | Health Professional | 42020705 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-04-18 | Malfunction |  | It was reported that during a da vinci si surgical procedure the wires of the articulation have been broken on the large needle driver instrument. The procedure continued with another instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. There were also indentations on the edge and scratches on the surface of the distal pulley. Engineering concluded it was most likely due to mishandling. An additional finding not reported was a scratch mark showing light material removal on the distal end of the main tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3140284 | 3112695 | 3009250 | 2955842201301662 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-19 | 2013-05-14 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-08 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure the cable broke on the monopolar curved scissors instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, one grip close cable was broken at the distal idlers. The idler pulley spun freely and is not damaged. The cable segment stuck out at the instruments wrist. The other cables at the wrist were not damaged. Additional damage found was a deep scratch mark exhibiting light material removal and a rough surface finish on the distal end of the main tube. Engineering concluded the damage may be due to mishandling. Electrical continuity passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3140313 | 3112731 | 3009286 | 2955842201301663 | 1 | NAY | Manufacturer | OtherUser facility | Other | FollowupInitial | 2013-04-17 | 2013-05-14 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42029602 | 2013-04-17 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure the large suturecut needle driver instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. Engineering also found indentations on the edge of the distal pulley. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3140387 | 3112782 | 3009337 | 2955842201301664 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-17 | 2013-05-14 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-02 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument was noted to be dull. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. These cracks were reported on the initial mdr report. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of the instrument dull scissors. Engineering placed the instrument on an in-house system for a latex cut test and the scissors cut cleanly through (b)(4) latex. The blades were undamaged. The instrument passed electrical continuity testing. Engineering found a hairline crack at the distal end of the main tube. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3140358 | 3112800 | 3009355 | 2955842201301665 | 1 | NAY | Manufacturer | OtherUser facility | Other | FollowupInitial | 2013-04-18 | 2013-05-14 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42029602 | 2013-03-25 | Malfunction |  | It was reported that during a da vinci si sacrolpopexy procedure the large suturecut needle driver instrument was not cutting. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument was not cutting. The instrument grip cable was found broken and derailed at the distal idler pulley, resulting in the instrument not being able to cut. An additional observation found damage on the surface and the edge of the idler pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3140411 | 3112830 | 3002229 | 2955842201301666 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-17 | 2013-05-14 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-24 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a cracked orange sleeve on the monopolar curved scissors instrument's tube extension. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering confirmed the cracked orange sleeve portion of the tube extension at the proximal clevis interface. The tube extension cracked next to one of the keys that mates with the proximal clevis. Evidence not conclusive, but tube extension likely cracked due to excessive side loading or other mishandling/misuse. Additional observations not reported but the customer was a cracked main tube. The distal end of the main tube had a hairline crack in the axial direction. The crack location overlapped the section of the tube extension that was installed inside the main tube. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3137311 | 3109238 | 3005793 | 2955842201301642 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-24 | 2013-05-13 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-24 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken wire on the fenestrated bipolar forceps instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering found a broken drive cable. The pitch down cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. Additional observations not reported but the customer main tube damage. The distal end of main tube had various scratch marks with light material removal. Scratches were 0. 070 - 0. 145 in length and were not aligned with the tube axis. Evidence not conclusive, but damage may have been caused by mishandling/misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3137344 | 3109284 | 3005839 | 2955842201301643 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-24 | 2013-05-13 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-30 | Malfunction |  | During a da vinci si hysterectomy procedure, a cable on the fenestrated bipolar forceps instrument allegedly broke. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering confirmed a broken pitch cable at the distal end of the instrument. The cable segment that contains the crimp was missing from the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. Additional observations not reported but the customer main tube damage. The distal and proximal ends of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Scratch mark on the distal end was approximately 0. 0315 in length and scratch marks on the proximal end were approximately 0. 1385 and 0. 905. Evidence not conclusive, but damage may be due to mishandling. Electrical continuity testing was performed and passed. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3137404 | 3109341 | 3005896 | 2955842201301644 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-30 | 2013-05-13 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-23 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility reportedly identified a misaligned wire on the prograsp forceps instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering confirmed a derailed grip cable that was off the pulley in between the distal clevis. The instrument's grips can still open and close but not fully. Additional observations not reported by the customer was main tube damage. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Evidence not conclusive, but damage may be due to mishandling. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3137473 | 3109407 | 3005962 | 2955842201301645 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-05-02 | 2013-05-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-09 | Malfunction |  | During a da vinci si prostatectomy procedure, the grip on the fenestrated bipolar forceps instrument allegedly was not closing. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering confirmed that one of the grip was bent, causing misalignment of the grips. There was an 0. 101 offset at the tips, indicating overloading at the instrument's tip. Electrical continuity testing was performed and passed. Additional observations not reported by the customer were pulley and main tube damages. There were scratches on the surface of the distal pulley. Engineering was unable to determine how the scratches occurred. The distal end of main tube had various scratch marks showing light material removal. Scratches were short in length and not axially aligned with the tube. Evidence not conclusive, but damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3137598 | 3109666 | 3006221 | 2955842201301646 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-06 | 2013-05-13 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-11 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted a broken cable on the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed cable at distal idler pulley. The frayed cable section was approximately 0. 08 in length. The damages were found on the surface of the pulley. Engineering also found a derailed grip cable at the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3137666 | 3109712 | 3006267 | 2955842201301647 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-15 | 2013-05-13 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-10 | Malfunction |  | During central processing, the user facility reportedly identified bent wires on the prograsp forceps instrument. Nothing reportedly fell into a patient. The instrument was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering confirmed the alleged complaint. The pitch cables were loose but were not visibly broken at the instrument's wrist. The pitch up cable was slightly kinked and unwound due to decreased cable tension. The pitch cable crimp was not seated properly in the crimp pocket causing slight tension decrease due to shifted crimp position. Other cables at the instrument's wrist were not damaged. Additional observations not reported by the customer was main tube damage. The distal end of the main tube had various scratch marks with light material removal. Scratches were 0. 100 - 0. 130 in length and were not aligned with the tube axis. Evidence not conclusive, but damage may have been caused by mishandling/misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3137710 | 3109758 | 3006313 | 2955842201301648 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-07 | 2013-05-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-05-07 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-17 | Malfunction |  | During central processing, the user facility reportedly identified jaws that were not meeting on the fenestrated bipolar forceps instrument. Nothing reportedly fell into a patient. The instrument was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering was unable to confirm the alleged complaint. The instrument's grips were not bent and the teeth meshed properly when grips were in the closed position. The grips were able to open/close fine when input discs were manually rotated. Additional observations not reported by the customer was a broken cable. The pitch down cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Other cables at instrument's wrist were not damaged. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3137760 | 3109827 | 3006382 | 2955842201301650 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-05-07 | 2013-05-13 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-05-07 | Device Not Retu | No | Reuse | 2013
42019407 | 2013-04-11 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the customer noted a broken cable on the mega needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be frayed at the distal idler pulley. The frayed strands stuck out at the wrist. Other cables at wrist were not damaged. Engineering also observed that the same frayed cable was also found to be derailed from the distal idler pulley. The grips were able to open and close, but their movement could not be precise. Engineering concluded that the derailed cable likely contributed to the frayed cable. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3137796 | 3109859 | 3006414 | 2955842201301649 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-15 | 2013-05-13 | No | Yes | Health Professional | 42019407 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42011704 | 2013-03-21 | Malfunction |  | It was reported that prior to starting a da vinci s gastric bypass procedure, the scrub nurse noticed that the housing of the needle driver instrument was damaged. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that three of four housing snap tabs and all four adjacent housing pins were broken on the instrument's blue housing. The housing could be removed as a result of the damage. Engineering concluded that damage was likely due to mishandling. Additional observation not reported by site was main insulation tube damage. The distal end of the main tube exhibited deep scratches with material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3138114 | 3110135 | 3006690 | 2955842201301652 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-15 | 2013-05-13 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted that the fenestrated bipolar instrument was not moving and they also observed wires coming out of the tip of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Instrument passed electrical continuity test. Engineering also observed scratches on the surface of the distal pulley. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3136330 | 3108396 | 2997165 | 2955842201301625 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-11 | 2013-05-12 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-28 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy the prograsp forceps instrument tip was not articulating completely. The customer trouble shot by cleaning the tip, fixing the drape, and reloading the instrument. The instrument was replaced with another prograsp forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument was not articulating completely. The pitch cable was found broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. Clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3136327 | 3108402 | 3004957 | 2955842201301634 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-15 | 2013-05-12 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-14 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the console surgeon noted difficulty with the patient side manipulator (psm) arm movements of the fenestrated bipolar forceps instrument in the middle of the procedure. The scrub nurse removed the instrument for inspection and found that the string of the instrument was cut. The instrument was replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with damage insulation on the conductor wire by the yaw pulley. Instrument passed electrical continuity test. Engineering concluded that damage was likely due to mishandling. Additional observations not reported by site were grip bent and insulation tube damage. The grip bent was causing side to side misalignment of grips. There was a 0. 102 offset at the tips, indicating overloading at the tip. The distal end of the main tube also exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that these damages were likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131867 | 3103837 | 3000392 | 2955842201301598 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-10 | 2013-05-10 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-03-27 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that a bent tip on the thoracic grasper instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grips were not bent and grip teeth mesh properly when in closed position. Distal end of main tube exhibited a slight bend roughly 2 above the snake wrist. When tube was rotated a slight wobble in the tube could be observed due to a bending. Additional observation not reported by site was main insulation damage. The black tube insulation was damaged at the distal end. The insulation was gouged on one side and had a 0. 100 x 0. 100 piece missing roughly 0. 250 above the snake wrist. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3131863 | 3103849 | 3000404 | 2955842201301596 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-10 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-25 | Malfunction |  | It was reported that during the da vinci s prostatectomy procedure, the surgeon could not open and/or close the tip of the monopolar curved scissors instrument due to a broken wire on the instrument. The surgical staff exchanged it into a backup instrument, and the planned procedure was completed without problem. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at the wrist. Other cables at wrist were not damaged. Instrument passed electrical continuity test. Engineering also found scratches on the surface of the distal pulley. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131850 | 3103850 | 3000405 | 2955842201301597 | 1 | NAY | Manufacturer | DistributorHealth Professional | Physician | Initial | 2013-04-09 | 2013-05-10 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-10 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure that a wire hanging out separated on the fenestrated bipolar forceps. The planned surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 03 in length. No other cable damage was found. Additional observation found a conductor wire broken at the distal clevis hub. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131856 | 3103855 | 3000410 | 2955842201301599 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-10 | 2013-05-10 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-10 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure that the prograsp forceps instrument had frayed cables around the jaw. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with a frayed pitch cable at the distal idler. No damage was found on the pulley and clevis. 48- additional findings were scratches on the surface of the distal pulley. Various scratch marks showing light material on the distal end of the main tube were also found. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3131857 | 3103856 | 3000411 | 2955842201301600 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-10 | 2013-05-10 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | No | Reuse | 2013
42020705 | 2013-04-10 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure that cables were broken on the inside of the tenaculum forceps instrument and it was not recognized. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, the pitch cable was broken at the distal end. The cable segment that contains the crimp was still in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the instruments wrist were not damaged. Additional findings not reported were broken snaps and pins on the housing causing it to be loose. The housing had two snaps and four pins broken off. Engineering concluded the damage was likely due to mishandling. Also found were various scratches on the main tube showing light material removal and a rough surface finish. Engineering concluded the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3131861 | 3103859 | 3000414 | 2955842201301601 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-10 | 2013-05-10 | No | Yes | Health Professional | 42020705 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-09 | Malfunction |  | It was reported that after a da vinci si surgical procedure the prograsp forceps instrument was noted to have a wire that was sticking out of the jaws of the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was derailed at the idler block pulley, resulting in a loose pitch cable at the distal clevis hub. The pitch cable was also found to be frayed. No damage was found on the clevis. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3132656 | 3104834 | 3001389 | 2955842201301586 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-15 | 2013-05-10 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-25 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s surgical procedure the monopolar curved scissors (mcs) instrument was noted to have a cable sticking out. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damages on the edge and on the surface. Engineering concluded the damage on the pulley was likely due to excess force contact and mishandling. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. Engineering also found the tube extension pad printing was removed at the distal end. The customer reported complaint does not itself constitute a mdr reportable event, in a fragment-causing failure (such as a cable failure, a crimp separating from a cable etc. ) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument; however, the pad printing removed, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3132684 | 3104843 | 2975471 | 2955842201301588 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-04-15 | 2013-05-10 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42022703 | 2013-04-09 | Malfunction |  | It was reported that prior to starting a da vinci si hysterectomy procedure the pk dissecting forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal end. The cable segment that contained the crimp was missing from the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. Engineering also found deep scratches on the main tube and a char mark on the yaw pulley. The distal end and the proximal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage was likely due to mishandling. The yaw pulley exhibited a char mark near the grip tip. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3132694 | 3104853 | 3001408 | 2955842201301581 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-10 | 2013-05-10 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | No | Reuse | 2013
42018907 | 2012-11-28 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the console surgeon experienced difficulty while opening and closing the jaws of the double fenestrated bipolar forceps instrument during the middle of the procedure. The scrub nurse removed the instrument for inspection and found that the string of the instrument was cut. The instrument was replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was found to be broken at the distal clevis hub. The cable segments that contain the crimps have fallen out of the clevis. The clevis did not exhibit excessive damage. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3132858 | 3105092 | 3001646 | 2955842201301602 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-10 | 2013-05-10 | No | Yes | Health Professional | 42018907 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-11 | Malfunction |  | It was reported that during a da vinci si cystectomy procedure that the prograsp forceps instrument was not working properly. It was alleged the instrument was not responding to the surgeons control. The customer stated that she believed a wire may have broke during the case. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3133105 | 3105230 | 3001785 | 2955842201301603 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-11 | 2013-05-10 | No | Yes | Health Professional | 42009308 | No | No | 2013-04-11 | No | No | Reuse | 2013
IS2000 A51P8 | 2013-04-10 | Malfunction |  | It was reported that during a procedure a schoelly 0 degree endoscope outer lens came off. The lens was discarded. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The endoscope was received with mechanical damage to the distal window assembly hermetic seal resulting in fluid invasion and subsequent major damage to the optical components. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3133247 | 3105317 | 3001871 | 2955842201301605 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-10 | 2013-05-10 | No | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-26 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility reportedly identified dull scissors on the monopolar curved scissors instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering confirmed the dull scissors complaint. The instrument was placed on an in house da vinci system for a latex cut test and failed testing. The instrument's scissors did not cut cleanly through 0. 006 latex and the latex was snagging at scissor tips. The blade edges were not damaged but the edges exhibited wear at the tips, negatively affecting cut performance. Additional observations not initially reported by the customer were a cracked main tube and tube extension. The distal end of main tube exhibited a hairline crack in the axial direction. The tube extension was cracked at the distal end, next to one of the keys that mates with the proximal clevis. Engineering noted that tube extension may likely cracked due to excessive side loading or other mishandling/misuse. Electrical continuity testing was performed and passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3133244 | 3105332 | 3001887 | 2955842201301604 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-05-08 | 2013-05-10 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-02 | Malfunction |  | It was reported that during a da vinci si nephrectomy, a cable broke on the fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with a frayed cable at the distal idler pulley. The frayed cable section was approximately 0. 05 in length. No damage was found on the pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3133353 | 3105450 | 3002005 | 2955842201301606 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-10 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-27 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy, a cable broke on the large suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed one grip cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. The other cables at the wrist were not damaged. Additional damage found was a corroded bearing. Yellowish material was found on the small bearing located on the upper chassis. Engineering concluded the damage was likely due to improper cleaning. No other damage was found. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3133572 | 3105675 | 3002229 | 2955842201301607 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-10 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-04-11 | Malfunction |  | It was reported that during reprocessing that the tip broke off the small clip applier instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the tip broke off the instrument. One grip was bent at the midpoint and the other grip was broken off. The grip was not returned with the instrument. Engineering concluded the bending was likely due to applying force normal to the grip while trying to install a clip, and by over force broke the tip off. The broken off tip roughly measured 0. 138 x 0. 056. The damage was a result of misuse. An additional finding was an indentation at the edge of the distal pulley. Engineering concluded it was most likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135552 | 3107638 | 3004192 | 2955842201301608 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-10 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2013-04-11 | Other |  | It was reported that during a da vinci s prostatectomy procedure, the patient side cart (psc) suddenly transitioned to battery power. The doctors attempted to restart the system; however, the error message psc is not connected. Check red cable and psc power was displayed on the monitor and the system would not power on. The surgeon made the decision to convert to traditional open surgical techniques to complete the planned surgical procedure. No patient harm was reported. | The investigation performed by the field service engineer concluded that the system issue experienced by the site was associated was associated with the patient side cart medical grade power supply (mgps). The mgps converts the ac power from the wall socket (through the power cord) and converts it to dc power for use on the system. The mgps was returned to isi for evaluation. Engineering was unable to replicate the issue experienced by the site. The mgps was installed on an isi in-house patient side cart (psc) and the psc powered up in a normal mode. Video, voltage, and image testing using the affected mgps passed. The mgps was power cycled ten times and there were no errors generated. On (b)(4) 2013, isi contacted the initial reporter concerning this event. The initial reporter indicated that the patient was in stable condition and did not experience any intra-operative complications as a result of the reported event. As of (b)(6) 2013, there have been no reported recurrences of this failure mode at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135619 | 3107713 | 3004267 | 2955842201301609 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-04-11 | 2013-05-10 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-11 | Malfunction |  | It was reported that during reprocessing that frayed wires were identified on a prograsp forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed frayed wires on the instrument. The instrument was found with frays on both pitch cables at the distal clevis. No damage was found on the pulley and clevis. One fray strand was approximately 0. 028 in length and the other was. 007 in length. Additional damage found were deep scratches on the distal end of the main tube. The scratch marks exhibited light material removal and a rough surface finish. The scratches varied in size ranging approximately from. 057 to. 355 in length. Engineering concluded the damage may have been due to mishandling. A derailed grip cable was also found. The grip cable was derailed off the clamping pulley inside the housing. As a result, the grips did not open/close properly. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135649 | 3107730 | 3004285 | 2955842201301610 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-10 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2011-01-07 | Injury | Required InterventionLife Threatening | On (b)(6) 2012, isi received a legal complaint concerning a patient who underwent a da vinci si mitral valve repair procedure that was converted to another technique. The legal complaint alleged that during the surgery the robotic arms had too much pressure applied to them causing the machine to jolt. The legal complaint also alleged that the patient experienced 15 hours of internal bleeding before determining that exploratory surgery was needed, and the surgery discovered that the robotic arms spliced and punctured the patient's liver and spleen. Due to the reported injuries sustained by the patient during the da vinci robotic surgical procedure, the legal complaint alleged that the patient had to have multiple additional medical tests and physician consultations, and had suffered pain, loss of function, emotional distress, and permanent injury. The legal complaint alleged that despite additional treatment, the patient continued to suffer from abdominal pain, fatigue, and decreased energy and stamina. | Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. Isi has reviewed the operative/procedure report of the patient. Based on the review of the operative report, there is no evidence that a malfunction of the da vinci si system, instruments, or accessories occurred, or caused or contributed to the patient's alleged injuries. The operative report indicated that the patient's chest was approached with the da vinci system where the diaphragm was very elevated despite several traction sutures that were placed. The operative report indicated that the surgical staff did not have enough room for the arms to work so the surgeon decided to abort this procedure and then extend the working port to make a lateral approach for the operation. At the completion of the surgery, the operative report indicated that the patient returned to the coronary care unit in satisfactory condition. Isi has reviewed the system logs for the surgical procedure performed at the site on (b)(6) 2011. A review of the system logs showed that no system errors occurred during the surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135695 | 3107770 | 3004325 | 2955842201301612 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-19 | 2013-05-10 | Yes | No | Service Personnel | IS3000 A60P6 | No | No |  | No | No | Reuse | 2013
42030901 | 2013-04-11 | Malfunction |  | It was reported that during reprocessing that frayed wires were identified on a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed frayed cables. The instrument had one cable frayed at two different locations at the distal clevis. One strand stuck out at the proximal clevis and the shorter strand stuck out near the instrument grip tips. No other cables were damaged at distal clevis. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135707 | 3107771 | 3004326 | 2955842201301613 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-10 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2008-08-22 | Other | Required InterventionLife Threatening | On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint with the following allegation: it was reported that during a da vinci si nephrectomy procedure performed on (b)(6) 2008, the surgeon allegedly caused a hole to develop in the patient's renal artery at the junction of the aorta. The legal complaint indicated that the surgeon made the decision to convert the procedure to open surgical techniques for an unspecified reason. The surgeon was reportedly able to control the patient's bleeding and retrieve the patient's left kidney. During the procedure, the surgeon was also reportedly informed that the instrument count in the operating room was inaccurate. The legal complaint also indicated that radiological films were performed and a possible drain or sponge was found to be retained inside the patient. The patient was reportedly moved back to the surgical table, reopened, and the drain or sponge was removed. No further clinical information was provided. | On (b)(4) 2013, isi contacted the risk management department of the site to obtain additional information regarding the reported event. The risk manager indicated that she could not provide any information regarding the reported incident. Based on the information provided, it is unknown if the da vinci si system, an instrument, and/or an accessory caused or contributed to the patient's reported injury. There is also no allegation that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the reported surgical procedure. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135708 | 3107774 | 3004329 | 2955842201301614 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | Initial | 2013-04-11 | 2013-05-10 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-04-11 | No | No | Reuse | 2013
IS3000 A60P8 | 2009-03-18 | Injury | Life Threatening | On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint with the following allegation: it was reported that after a da vinci si nephrectomy procedure was performed on (b)(6) 2009, bleeding was noted from the patient's aorta. No further clinical information was provided. | On (b)(4) 2013, isi contacted the risk management department of the site to obtain additional information regarding the reported event. The risk manager indicated that she could not provide any information regarding the reported incident. Based on the information provided, it is unknown if the da vinci si system, an instrument, and/or an accessory caused or contributed to the patient's reported injury. In addition, at this time there is no allegation that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the reported surgical procedure. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135702 | 3107777 | 3004332 | 2955842201301616 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | Initial | 2013-04-11 | 2013-05-10 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-04-11 | No | No | Reuse | 2013
42018404 | 2013-04-15 | Malfunction |  | It was reported that during reprocessing of the permanent cautery hook instrument, it was noted that the tip of the instrument looked burnt. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the yaw pulley and distal clevis had char marks. Indicating that an arcing event occurred. The yaw pulley char marks were located on the side opposite of the conductor wire and the distal clevis char marks were on the same side as the conductor wire. Engineering evaluation also found that the instrument's conductor wire and the main tube exhibited damage. The conductor wire was broken near the distal idler pulley and the wire segment was sticking out at the wrist. The instrument failed electrical continuity testing. Engineering concluded that the wire breakage likely created a path for arcing. The distal end of the main tube exhibited scratches with light material removal. The scratches were. 90 -. 530 in length and some scratches were aligned with the tube axis. Engineering concluded that the damage to the main tube is likely due to mishandling and misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. On (b)(4) 2013, intuitive surgical, inc. (isi) contacted the initial reporter concerning this event. The initial reporter indicated that during set up of a da vinci surgical procedure, it was noted by the surgical team that the permanent cautery spatula (pcs) instrument appeared to be unclean. The surgical team sent the instrument back to central reprocessing to be reprocessed and it was during reprocessing of the instrument that it was noted that the instrument was not dirty, but was charred. The initial reporter indicated that there was no report by the surgical staff that arcing from the instrument was observed during a surgical procedure and there was no report by the surgical staff that any patient injury occurred as a result of arcing from the pcs instrument while in use during a surgical procedure. The initial reporter indicated that there was no report by the surgical staff that any piece(s) from the instrument fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3135706 | 3107791 | 3004345 | 2955842201301611 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-15 | 2013-05-10 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2008-09-26 | Other | Required InterventionLife Threatening | On (b)(6) 2013, intuitive surgical inc. (isi) received a legal complaint with the following allegation: it was reported that during a da vinci si nephrectomy procedure performed on (b)(6) 2008, one of the patient's lumbar veins began to bleed. The legal complaint claimed that the bleeding started while the surgeon was dissecting the patient's left ovarian vein. The legal complaint also indicated that while attempting to control the patient's bleeding, the surgeon allegedly injured the patient's distal aorta with a harmonic scalpel instrument. The surgeon allegedly created an 8 to 10 mm hole in the patient's distal aorta. During the event, the legal complaint claimed that the patient lost her blood pressure for several minutes and cardiopulmonary resuscitation was initiated. The surgeon reportedly aborted the nephrectomy procedure and converted to open surgical techniques to repair the patient's aorta. The patient's kidney was not harvested. No further clinical information was provided. | On (b)(4) 2013, isi contacted the risk management department of the site to obtain additional information regarding the reported event. The risk manager indicated that she could not provide any information regarding the reported incident. Based on the information provided, it is unknown if the da vinci si system, an instrument, and/or an accessory caused or contributed to the patient's reported lumbar vein injury. In addition, at this time there is no allegation that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the reported event. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135700 | 3107792 | 3004346 | 2955842201301615 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | Initial | 2013-04-11 | 2013-05-10 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-04-11 | No | No | Reuse | 2013
IS3000 A60P8 | 2010-01-27 | Other | Required InterventionLife Threatening | On (b)(6) 2013, intuitive surgical inc. (isi) received a legal complaint with the following allegation: it was reported that during a da vinci si nephrectomy procedure performed on (b)(6) 2010, the surgeon allegedly transected a renal artery branch to the lower pole of the patient's kidney with the robot's harmonic scalpel. Due to transecting the renal artery branch, the surgeon reportedly disconnected the robot and converted to an open surgical procedure to remove the patient's kidney. No further clinical information was provided. | On (b)(6) 2013, isi contacted the risk management department of the site to obtain additional information regarding the reported event. The risk manager indicated that she could not provide any information regarding the reported incident. Isi has reviewed the system error logs for the site with a date of (b)(6) 2010 and no surgical procedures were found. At this time there is no indication that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the reported surgical procedure. Based on the information provided, it is unknown if the da vinci si system, an instrument, and/or an accessory caused or contributed to the patient's reported injury. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135713 | 3107796 | 3004350 | 2955842201301617 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | Initial | 2013-04-11 | 2013-05-10 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-04-11 | No | No | Reuse | 2013
IS3000 A60P8 | 2009-12-02 | Other | Required InterventionLife Threatening | On (b)(6) 2013, intuitive surgical inc. (isi) received a legal complaint with the following allegation: it was reported that during a da vinci si adrenalectomy, distal pancreatectomy, and splenectomy procedure performed on (b)(6) 2009, the surgeon allegedly injured the patient's aorta. The legal complaint indicated that the injury occurred during dissection and the surgeon made the decision to convert the robotic surgery to open surgical techniques. The legal complaint alleged that during transition, the robot's arm was inappropriately moved, causing a tear in the patient's aorta. The surgeon reportedly repaired the aorta. The legal complaint indicated that post-operatively, the patient experienced renal failure and the family withdrew life support shortly after surgery. No further clinical information was provided. | On (b)(6) 2013, isi contacted the risk management department of the site to obtain additional information regarding the reported event. The risk manager indicated that no malfunction of the da vinci system, instruments, or accessories occurred during the reported event. The risk manager claimed that the patient injury was due to user-error and a communication issue between the surgical staff. The risk manager explained that during the conversion to open surgery, a surgical staff member moved the robot and was not aware that one of the robot's arms was still grasping a bleeder blood vessel which caused another tear. The risk manager denied that the robot arm moved on its own. She also indicated that the patient was not on life support, but was a hospice patient. The risk manager indicated that she could not provide any additional information regarding the reported incident. Based on the information provided from the risk manager, there is no indication that a malfunction of the da vinci system, an instrument, and/or an accessory caused or contributed to the patient's initial injury to the aorta during dissection. At this time there is no allegation that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the reported event. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135716 | 3107798 | 3004352 | 2955842201301619 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | Initial | 2013-04-11 | 2013-05-10 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-04-11 | No | No | Reuse | 2013
IS3000 A60P8 | 2011-03-21 | Injury | Other | On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint with the following allegation: it was reported that shortly after a da vinci si nissen fundoplication procedure performed on (b)(6) 2011, the patient reportedly developed gastroparesis. The patient consulted with a gastroenterologist who concluded that the patient's gastroparesis allegedly developed because her vagus nerve was damaged. No further clinical information was provided. | On (b)(4) 2013, isi contacted the risk management department of the site to obtain additional information regarding the reported event. The risk manager indicated that she could not provide any information regarding the reported incident. Isi has reviewed the system error logs for the site with a date of (b)(6) 2011 (the date of the reported surgical procedure). Based on a review of the system error logs, there is no evidence that a malfunction of the da vinci si system, an instrument, or an accessory occurred during the surgical procedure. Based on the information provided, it is unknown if the da vinci si system, an instrument, and/or an accessory caused or contributed to the patient's reported injury. At this time, it is unknown how the patient's vagus nerve was allegedly injured. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135720 | 3107800 | 3004354 | 2955842201301621 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | Initial | 2013-04-11 | 2013-05-10 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-04-11 | No | No | Reuse | 2013
41032204 | 2013-04-18 | Malfunction |  | During a da vinci si surgical procedure, the vessel sealer instrument reportedly stopped working. No further details were reported. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering placed the instrument an in house da vinci system for testing and confirmed that the cautery function was not working. After inspecting the instrument, engineering found a damaged insulation on the conductor wire at the instrument's snake wrist. The conductor wire was exposed at two different locations. Electrical testing was performed and passed. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | ENDOWRIST ONE VESSEL SEALER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3135717 | 3107804 | 3004359 | 2955842201301618 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-10 | No | Yes | Health Professional | 41032204 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-11 | Malfunction |  | It was reported that during reprocessing that a wire was separated on a mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, the instrument grip cable was found frayed and derailed at the distal idler pulley. The frayed cable section was approximately 0. 06 in length. Additionally, damage on the surface of the distal idler pulley was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135719 | 3107806 | 2997360 | 2955842201301620 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-10 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-19 | Malfunction |  | During a da vinci si gastrectomy surgical procedure, the thread allegedly came out of the fenestrated bipolar forceps instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering found a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Additional observations not reported but the customer were: broken conductor wire; bent grips; pulley damage; and main tube damage. One conductor wire was broken at the yaw pulley exit. The wire was detached from its connection at the grip. Electrical continuity testing on the instrument was performed and failed. The yaw pulley showed no signs of arcing. One grip was bent causing side to side misalignment of grips. There was a 0. 067 offset at the tips, indicating overloading at the tip. Evidence not conclusive, but damage most likely due to mishandling. There were scratches on the surface of the distal pulley: unknown how the scratches occurred. The distal end of main tube had various scratch marks showing light material removal at the proximal clevis side. The scratches were short in length and were not axially aligned with the tube. Evidence not conclusive, but damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3135741 | 3107818 | 3004373 | 2955842201301623 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-30 | 2013-05-10 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-30 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2010-05-21 | Other | Required InterventionLife Threatening | On (b)(4) 2013, intuitive surgical inc. (isi) received a legal complaint with the following allegation: it was reported that during a da vinci si nephrectomy procedure, the surgeon allegedly injured the patient's aorta when first inserting a 'trocar,' or pathway device allowing access to the abdomen. The patient reportedly experienced bleeding from her aorta. The surgeon made the decision to abort the robotic procedure and to convert to open surgical techniques to stop the reported bleeding and to repair the injury sustained by the patient's aorta. The patient's kidney was not harvested. During the surgical procedure, the surgeon was reportedly informed that there was an incorrect instrument count. However, the surgeon reportedly closed the patient's surgical incision. After the patient was closed up, an x-ray was performed on the patient in an attempt to find if any instruments had been retained. The surgeon reviewed the x-ray and reportedly did not see any retained instruments. The patient was then taken from the operating room and while recovering, she allegedly experienced respiratory distress. A second x-ray performed revealed that the patient had retained a sponge. The formal complaint indicates that the surgeon performed a secondary open surgical procedure using a wide trocar to retrieve the sponge from the patient. The legal filing also alleged that the patient sustained a peroneal nerve injury as a result of surgical padding pressure on her peroneal nerve during surgery. No further clinical information was provided. | On (b)(4) 2013, isi contacted the risk management department of the site to obtain additional information regarding the reported event. The risk manager indicated that she could not provide any information regarding the reported incident. Based on the information provided, it is unknown if the da vinci si system, an instrument, or an accessory caused or contributed to the patient's reported injury. At this time, there is no allegation that a malfunction of the da vinci si system, an instrument, or accessory occurred during the reported event. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135728 | 3107821 | 3004355 | 2955842201301622 | 1 | NAY | Manufacturer | OtherCompany Representative | Other | Initial | 2013-04-11 | 2013-05-10 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-04-11 | No | No | Reuse | 2013
42019408 | 2013-03-05 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted one of the cables of the mega needle driver instrument was cut during the procedure. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135783 | 3107837 | 3004391 | 2955842201301624 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-08-13 | 2013-05-10 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-08-13 | Yes | No | Reuse | 2013
42029602 | 2013-04-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted a broken cable on the large suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at wrist. Other cables at wrist were not damaged. Engineering also observed scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135846 | 3107877 | 3004451 | 2955842201301628 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-10 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-03 | Malfunction |  | It was reported that during reprocessing the flush port was stuck and would not come back up. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the flush port was stuck and would not come back up. The instrument was returned with a luer fitting attached to the luer plate at the primary flush port. The luer fitting was stuck and could not be removed by applying axial and twisting forces. The luer fitting was the incorrect size for luer plate. Engineering concluded the issue was likely due to improper cleaning. An additional observation not reported by the site was tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 070 -. 130 in length and not aligned with the tube axis. Engineering concluded the damage may have been caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135817 | 3107903 | 3004458 | 2955842201301626 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-12 | 2013-05-10 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-12 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted a broken cable on the large suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at wrist. Other cables at wrist were not damaged. Additional observations not reported by the site were derailed grip cable and pulley damage. The derailed grip cable was found at the distal idler pulley. The grips could be open and close, but their movement could not be precise. The cable derailment was likely due to cable losing contact with the pulley during the wrist articulation. Fleet angle between proximal and distal pulleys may contribute to derailment. The distal pulley also exhibited some scratch marks on the surface most likely due to the cable. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135762 | 3107906 | 3004461 | 2955842201301627 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-10 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-03 | Malfunction |  | Event Description: --/-- It was reported that during reprocessing a monopolar curved scissors instrument, the blades of the scissors were black and was unable to be cleaned. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. These cracks were reported on the initial mdr report. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering confirmed the reported complaint. Both scissor tips had a black discoloration that did not come off when wiped with a moist towel or scraped with a pick. Engineering concluded the blades likely turned black as a result of heavy cautery use. Additional observation not reported by site was a cracked tube. The distal end of the main tube exhibited a couple of hairline cracks parallel to each other, oriented in the axial direction. The cracks overlapped with the section of tube extension that gets installed inside the main tube. There was a small bulge in the tube diameter near the cracks. The instrument was disassembled and found the cracks did not align with axial key slot. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3135849 | 3107917 | 2980244 | 2955842201301629 | 1 | NAY | Manufacturer | OtherUser facility | Other | InitialFollowup | 2013-04-12 | 2013-05-10 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42009308 | 2013-03-14 | Malfunction |  | It was reported that during a da vinci si prostatectomy a cable broke on the prograsp forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135828 | 3107919 | 3004474 | 2955842201301630 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-12 | 2013-05-10 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-12 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-04-11 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that an insulation piece came off from the monopolar cautery instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the main insulation tube exhibited scratch and gouge marks with materials removed at the distal end. Engineering concluded that damage was likely caused by excess force contact on the main tube surface. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3135853 | 3107920 | 3004475 | 2955842201301631 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-11 | 2013-05-10 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | Yes | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted the wire on the tip of the mega suture cut needle driver instrument is broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip cable was found to be broken at the distal idlers. The idler pulley was able to spin freely, but it exhibited some wear. The cable segment stuck out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURE CUT NEEDLE DRIVERINSTRUMENT |  | 3135863 | 3107924 | 3004479 | 2955842201301632 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-12 | 2013-05-10 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-12 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-04-17 | Other | Other | It was reported that prior to starting a da vinci si prostatectomy procedure, when the site powered on the system, the surgeon side cart (ssc) lost power. The intuitive surgical, inc. (isi) clinical sales representative was on-site when the site experienced the power issue. The csr attempted to power on the ssc; however, the issue recurred. The surgeon aborted the planned surgical procedure and re-scheduled it for a later date. No patient harm was reported. | The investigation performed by the field service engineer concluded that the system issue experienced by the site was associated was associated with the patient side cart medical grade power supply (mgps). The mgps converts the ac power from the wall socket (through the power cord) and converts it to dc power for use on the system. The mgps was replaced to repair the system. The mgps was returned to isi for evaluation. Engineering replicated the failure mode experienced by the site. The mgps was installed on an intuitive surgical, inc. (isi) in-house surgeon side cart (ssc). The ssc did not power on with the mgps installed. A visual inspection of the mgps found that it was in good condition. As of (b)(6) 2013, there have been no reported recurrences of this failure mode at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135869 | 3107932 | 3004487 | 2955842201301635 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-04-17 | 2013-05-10 | Yes | Yes | Health Professional | IS3000 A60P8 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-04-16 | Other | Other | It was reported that during a da vinci si gastrectomy procedure, while the surgeon was performing dissection of the patient's hepatic artery, the patient side manipulator (psm) arm 1 behaved in a sluggish and resistant manner and intermittently would not move when the surgeon manipulated the psm using the right master tool manipulator (mtmr) on the surgeon side cart. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | During the field investigation, the intuitive surgical, inc. , (isi) field service engineer (fse) was able to reproduce the reported failure mode experienced by the site. The fse experienced resistance motion while controlling an instrument on psm 3 as described by the customer after an hour of testing. The fse also clutched psm 3 and moved it by hand, as a patient side assistant would do. The fse reported that the outer yaw axis felt more difficult to move than expected. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instrument. The system was repaired by replacing psm 1 and psm 3. Psm 2 was replaced as a precaution. On (b)(6) 2013, the initial reporter of this complaint provided isi with additional information regarding the reported event. The initial reporter indicated that the site did not experience any issues with the da vinci surgical system during set-up. The initial reporter indicated that the patient did not experience any intra-operative complications as a result of the reported event and that the patient was in stable condition after successful completion of the open surgical procedure. Per the information provided by the initial reporter the patient was discharged from the hospital on (b)(6) 2013 and the patient did not return to the hospital due to experiencing any post-surgical complications. The psm 1 psm 2 and psm 3 were returned to isi for failure analysis investigation. Engineering evaluation found that the root cause of the customer reported failure mode was due to mechanical interference between the gears in psm 1 and psm 3. The gears will be replaced to repair the psms. As of (b)(6) 2013, there have been no reported recurrences of this failure mode at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3135864 | 3107943 | 3004498 | 2955842201301633 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-04-16 | 2013-05-10 | Yes | Yes | Health Professional | IS3000 A60P8 | Yes | No | 2013-04-16 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-04-10 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, one of the wheels would not turn on the mega needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was placed on a system and driven. Recognition and engagement passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. The instrument was fully functional. An additional finding not reported were various scratches showing light material removal and a rough surface finish on the distal end of the main tube. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135902 | 3107954 | 3004509 | 2955842201301637 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-15 | 2013-05-10 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-05 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si total benign hysterectomy procedure the monopolar curved scissors instrument was not recognized by the system. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument. The pogo pins dif not stick and were not contaminated. Electrical continuity passed. An additional finding not reported was pad printing removed on the tube extension. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3135895 | 3107962 | 2999621 | 2955842201301636 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-04-15 | 2013-05-10 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42017207 | 2012-10-10 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the tungsten cable at the endowrist joint was fraying on the maryland bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3135875 | 3107973 | 3004528 | 2955842201301638 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-15 | 2013-05-10 | No | Yes | Health Professional | 42017207 | No | No | 2013-04-15 | No | No | Reuse | 2013
42017910 | 2013-04-10 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the monopolar curves scissors instrument orange sleeve cracked when it was removed from the trocar. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument orange sleeve was cracked. The tube extension was broken with a piece removed at the prox clevis interface. The clevis was dislodged from tube extension. Engineering concluded the tube extension likely broke due to excessive side loading or other mishandling. Electrical continuity test passed. An additional observation not reported was a hairline crack at the distal end of the main tube. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3135910 | 3107977 | 3004532 | 2955842201301639 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-15 | 2013-05-10 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-10-08 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the large needle driver instrument jaw was fractured after 5 uses. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135911 | 3107978 | 3004533 | 2955842201301640 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-15 | 2013-05-10 | No | Yes | Health Professional | 42000606 | No | No | 2013-04-15 | No | No | Reuse | 2013
42017704 | 2013-04-15 | Malfunction |  | It was reported that during reprocessing a bowl grasper instrument had cut insulation on the shaft. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the reported complaint. The instrument shaft had insulation coating peeled off near distal end. No pieces were missing. There were also scratches along the main tube. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3135912 | 3107979 | 3004534 | 2955842201301641 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-15 | 2013-05-10 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-03 | Malfunction |  | It was reported that during reprocessing that a wire was broken on the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the edge and surface. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. Engineering concluded the damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3131017 | 3103034 | 2999588 | 2955842201301569 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-09 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-05 | Malfunction |  | It was reported that during reprocessing that the wire was broken by the tip on the monopolar curved scissors instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely but it was damaged on the edge. Engineering concluded the damage was likely due to mishandling. The cable segment stuck out at the instruments wrist. The other cables at the wrist were not damaged. Additional damage found was pad printing removed on the tube extension. Several markings were found on the tube extension exhibiting pad printing removal. Electrical continuity passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131064 | 3103067 | 2999621 | 2955842201301571 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-09 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-08 | Malfunction |  | It was reported that during reprocessing the mega suturecut needle driver instrument that the wire broke. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the surface and edge. Engineering concluded the damage was likely due to mishandling. The cable segment stuck out at the instrument's wrist. Other cables at the wrist were not damaged. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3131034 | 3103081 | 2999636 | 2955842201301570 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-09 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-26 | Malfunction |  | It was reported that during a da vinci si total hysterectomy, the monopolar curved scissors were blocked. The customer attempted to move with emergency grip release without success. The surgery was stopped and removed with trocars. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint with the following clarification: the tube extension was broken. The instrument was returned with tip cover installed. The tip cover exhibited tearing that exposed the proximal clevis. The tube extension was broken and missing a. 210 x. 150 piece at the proximal clevis interface. The clevis was dislodged from tube extension. Dislodgment may make removal of instrument from cannula difficult. The tube extension fractured next to one of the keys that mates with the prox clevis. Engineering concluded the tube extension likely broke due to excessive side loading or other mishandling/misuse. Electrical continuity passed. No other damage was found. On (b)(4) 2013, isi attempted to contact the initial reported of this complaint to obtain additional information regarding the reported event. At the time of this report, no additional information has been provided. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131276 | 3103274 | 2999829 | 2955842201301572 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-09 | 2013-05-09 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-04-08 | Injury | Other | Event Description: --/-- It was reported that during a da vinci si lobectomy procedure, the surgical staff allegedly noticed an electric current leak from the resin portion of the permanent cautery spatula (pcs) instrument. It was also reported that the patient's tissue and the area around the resin portion of the instrument were burned. Reportedly there was no bleeding around the tissue area that was burned. The surgical staff reportedly replaced the pcs instrument with a backup instrument and the planned procedure was completed with no additional problems. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering evaluation was unable to confirm the customer reported failure mode. A visual inspection of the instrument showed no damage to the conductor wire and there was no presence of char marks on any of the distal end components. The instrument passed electrical continuity testing. Engineering also performed a functional cautery test with the instrument using an in-house generator and monopolar cord. Engineering found that the instrument was able to transfer energy to a wet towel in both cut and coag cautery modes. The instrument was driven on an in-house is3000 system and the instrument was found to move intuitively with full range of motion in all directions. No physical damage was found. Intuitive surgical has made several attempts to contact the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) obtained the following additional information regarding the reported event: the permanent cautery spatula instrument and cannula were inspected prior to the surgical procedure and no issues were observed. The cannula was inspected with a pin gage. Arcing was observed by the surgical staff during a left upper lobectomy procedure. According to the surgical staff, the electrical energy arced to the upper leaf parenchyma. The details of the burned tissue were not provided. The patient did not receive any treatment or require intervention due to sustaining the tissue burn. The arcing event was observed approximately 10 seconds after the operation started and at that time, a fenestrated bipolar forceps instrument was in use. A review of the surgical procedure on a recorded video did not show any evidence of instrument collisions. The permanent cautery spatula instrument was not removed at any time during the surgical procedure. The grounding pad was properly positioned by the patient's thigh. During the surgical procedure, a vio300d electrosurgical unit (esu) was used with a swift coag effect mode of 4 and 30w setting. The patient did not experience any post-operative complications and was in stable condition. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 3131300 | 3103295 | 2999849 | 2955842201301574 | 1 | NAY | Manufacturer | User facilityDistributor | Physician | InitialFollowup | 2013-04-09 | 2013-05-09 | Yes | No | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2013-04-09 | Device Not Retu | No |  | 2013
40018010 | 2013-04-01 | Malfunction |  | It was reported that during setup of a da vinci s prostatectomy procedure, the surgical staff found an unsealed package of a monopolar curved scissors (mcs) tip cover accessory. The mcs tip cover accessory was not used in the surgical procedure. The planned surgical procedure was completed as planned using an mcs tip cover accessory from a new package. No patient harm, adverse outcome, or injury was reported. | The mcs tip cover accessory was returned and evaluated. Engineering evaluation found no damage with mcs tip cover accessory. Per the customer reported complaint, engineering evaluation was unable to confirm the alleged issue of an unsealed mcs tip cover accessory package. The mcs tip cover accessory was returned to isi without the package that was reportedly unsealed. Engineering was unable to confirm if the mcs tip cover accessory package was sealed or not although a photograph of an unspecified package was provided. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3131328 | 3103318 | 2999873 | 2955842201301573 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2013-04-09 | 2013-05-09 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | Yes | 2013-04-09 | Device Not Retu | Yes | Initial | 2013
42018406 | 2013-04-10 | Malfunction |  | During central processing, the user facility identified cracked insulation on the permanent cautery spatula instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering confirmed a broken ceramic sleeve at the distal end of the instrument. Small diameter section of sleeve had a 0. 085 x 0. 085 piece broken off. Larger diameter section of sleeve exhibited various scratch marks below the broken section. Additional damage found not reported by the customer was a bent spatula. The spatula was bent to one side at the junction between the round shank and flattened spatula. Evidence not conclusive, but damage to sleeve and spatula may be due to mishandling/misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131437 | 3103416 | 2999970 | 2955842201301576 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2013-04-26 | 2013-05-09 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2013-04-26 | Device Not Retu | No | Reuse | 2013
41032204 | 2013-02-05 | Malfunction |  | During a da vinci si surgical procedure, the user facility reportedly identified a blade on the vessel sealer instrument was not returning to the housing. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering performed visual inspection of the instrument and confirmed an exposed blade between the instrument's grips. There was also some bio debris within the blade tracks, possibly affecting the blade travel. Additional damage found not reported by the customer were broken cables. Two drive cables were sticking out at the distal end of the snake wrist. Wrist motion will not be intuitive due to the cable breakage. Snake wrist discs did not exhibit any obvious damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | ENDOWRIST ONE VESSEL SEALER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131415 | 3103429 | 2999984 | 2955842201301575 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-22 | 2013-05-09 | No | Yes | Health Professional | 41032204 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-04-09 | Malfunction |  | During central processing, the user facility identified a hole in the insulation of the thoracic grasper instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering confirmed that the black main tube insulation was damaged at the distal end. The insulation was gouged on one side and had a 0. 050 x 0. 050 piece missing roughly 0. 800 above the instrument's snake wrist. The distal end also exhibited various scratch marks without material removal. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3131464 | 3103466 | 3000020 | 2955842201301577 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-26 | 2013-05-09 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | Yes | 2013-04-26 | Device Not Retu | No | Reuse | 2013
40018012 | 2013-03-26 | Malfunction |  | It was reported that during a da vinci si total hysterectomy, the monopolar curved scissors were blocked. The customer attempted to move with emergency grip release without success. The surgery was stopped and removed with trocars. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the mcs tip cover exhibited three different tears. The first tear was roughly 0. 280 long, perpendicular to the longitudinal axis of the tip cover, located 0. 860 below the distal tip. The first tear lined up with the tube extension breakage found on returned monopolar curved scissors instrument mdr number 2955842-2013-01572. The second tear on the combination pellethane/silicone section was located 0. 465 below the distal tip. The third tear was located at the distal end opening that was axially oriented. No arcing related damage was observed. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3131541 | 3103552 | 3000107 | 2955842201301578 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-09 | 2013-05-09 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Initial | 2013
41032204 | 2013-03-21 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted broken and exposed wires on the vessel sealer instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that two adjacent drive cables were sticking out at the distal end of the snake wrist. The instrument's motion could no longer be intuitive due to the broken cables. No obvious damage observed on snake wrist discs. The grips and the blade track exhibited dried bio debris, but blade was not exposed. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | VESSEL SEALING INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131592 | 3103606 | 3000161 | 2955842201301579 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-09 | No | Yes | Health Professional | 41032204 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-28 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, it was noted that the prograsp forceps instrument did not open, instrument was not used in case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on the in-house da vinci robot system and driven without any problem. The instrument passed recognition and engagement test. Instrument moved intuitively with full range of motion in all directions. Grips opened and closed properly. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were 0. 060 - 0. 120 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3131636 | 3103610 | 3000165 | 2955842201301580 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-09 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-03-29 | Malfunction |  | It was reported that prior to starting a da vinci si lobectomy procedure the thoracic grasper instrument was noted to have an issue with angulation. The staff inspected the instrument and the angulation wires were noted to broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was found with grip cables broken at the snake wrist, between the first and second wrist disc. Engineering also found multiple main tube scratches at the distal end with material removed. The main tube was bent at the distal end. Engineering concluded that the damages were likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3131655 | 3103617 | 3000172 | 2955842201301589 | 1 | NAY | Manufacturer | ForeignUser facility | Nurse | Initial | 2013-04-15 | 2013-05-09 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | Yes | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-03 | Malfunction |  | It was reported that during a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument smoke and bubbling was noted under the tip cover at the proximal end. The tip cover was removed and replaced and the procedure was continued without further occurrence. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering placed the instrument on an in-house system with the tip cover, and no smoke and bubbling was observed. The tip cover was returned with the instrument and no fitting issue was found. Engineering also found pad printing removed at multiple locations on the tube extension. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131656 | 3103618 | 3000173 | 2955842201301590 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-04-15 | 2013-05-09 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-04-12 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the bowel grasper instrument was noted to have a tear in the black rubber sheath. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the main tube insulation appeared to have scratches and gouge marks, with materials removed near the midpoint of the main tube. Engineering concluded that the damage was likely caused by excess force contact on the main tube surface. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | BOWEL GRASPER 5MM INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3131658 | 3103619 | 3000174 | 2955842201301593 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-12 | 2013-05-09 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-04-12 | Device Not Retu | No | Reuse | 2013
40016907 | 2012-10-11 | Malfunction |  | It was reported that during a da vinci si gastric bypass procedure the harmonic curved shears instrument accessory was noted to have a broken tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC INSTRUMENT ACCESSOR | 3131643 | 3103634 | 3000189 | 2955842201301582 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-04-11 | 2013-05-09 | No | Yes | Health Professional | 40016907 | No | No | 2013-04-11 | No | No | Unkown | 2013
40016907 | 2012-10-24 | Malfunction |  | It was reported that during a da vinci si gastric bypass procedure the harmonic curved shears insert instrument accessory was noted to have a broken tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC INSTRUMENT ACCESSOR | 3131645 | 3103636 | 3000191 | 2955842201301583 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-04-11 | 2013-05-09 | No | Yes | Health Professional | 40016907 | No | No | 2013-04-11 | No | No | Unkown | 2013
40016907 | 2012-09-19 | Malfunction |  | It was reported that during a da vinci si gastric bypass procedure the harmonic curved shears insert instrument accessory was noted to have a broken tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | HARMONIC CURVED SHEARS INSERTACCESSORY | ENDOSCOPIC INSTRUMENT ACCESSOR | 3131649 | 3103639 | 3000194 | 2955842201301584 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-04-11 | 2013-05-09 | No | Yes | Health Professional | 40016907 | No | No | 2013-04-11 | No | No | Unkown | 2013
42018304 | 2013-04-08 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the permanent cautery hook instrument cautery connection was noted to be broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the banana plug was missing from the back end. The nut that retained the plug was loose in the chassis. The plug became unscrewed. The instrument was built in 2006, prior to implementation of a lock washer and loctite to help with retention of the plug. The instrument passed electrical continuity (using metal contact at conductor end) testing. Engineering also found a damaged ceramic sleeve and proximal clevis. The entire ceramic sleeve was missing from the yaw pulley. The sleeve likely broke in more than one piece. The proximal clevis was cracked in the axial direction halfway between the ears. The crack orientation was consistent with overloading the tip with the wrist pitched. Engineering concluded that both failures were likely mishandling related. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131652 | 3103641 | 3000196 | 2955842201301585 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-09 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-15 | Malfunction |  | The pk dissecting forceps instrument was returned, no further information was provided. | The instrument was returned and evaluated. Engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp remained in the clevis. The clevis did not exhibit excessive damage. Engineering also found deep scratches on the instrument main tube, with some materials removed at the distal end. Engineering concluded the damage was likely caused by excess force contact on the main tube surface. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131595 | 3103643 | 3000198 | 2955842201301587 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-15 | 2013-05-09 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-20 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci si prostatectomy procedure the monopolar curved scissors (mcs)instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering placed the instrument on an in-house system and it was checked for recognition; the instrument was successfully recognized. The pogo pins did not stick and were not contaminated. The dcp (dallas chip programmer) verification of the instrument passed. The instrument was driven and moved intuitively and the grips opened and closed. Engineering found hairline cracks at the distal end of the main tube. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. These cracks were reported on the initial mdr report. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3131657 | 3103644 | 3000199 | 2955842201301591 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-04-15 | 2013-05-09 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42020505 | 2013-04-12 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the fenestrated bipolar forceps instrument wire snapped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. The distal end of the main tube had various scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131596 | 3103645 | 3000200 | 2955842201301592 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-12 | 2013-05-09 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-04-12 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-03-22 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the mega needle driver instrument was noted to have a broken connector. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but has damages on the edge. Engineering concluded the damage was likely due to mishandling. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3131659 | 3103646 | 3000201 | 2955842201301595 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-18 | 2013-05-09 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-18 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-04-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the maryland bipolar forceps instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument. The pogo pins were not stuck or contaminated. Engineering also found bent grips, deep scratches on main tube. One grip was bent, causing side-to-side misalignment of grips. There is a. 0545 offset at the tips. Engineering concluded that the damage was likely due to mishandling. The distal end of the main tube had various scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded that the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3131597 | 3103647 | 3000202 | 2955842201301594 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-17 | 2013-05-09 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-17 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-04 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si total benign hysterectomy procedure, the mega suturecut needle driver instrument was noted to have wires sticking out and broken. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. The pulley was scratched and the distal clevis post was worn. The distal idler pulley on which broken cable was seated exhibited scratch marks on the outer surface. The top and side of the clevis post on the same side as the broken cable exhibited wear. The evidence was inconclusive, but the pulley and post damage suggests the cable derailed and was rubbing against these components, which likely contributed to cable breakage. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | Reuse | 3128528 | 3100451 | 2997025 | 2955842201301561 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-05-15 | 2013-05-08 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-15 | Yes | No |  | 2013
42806503 | 2013-04-04 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the single-site port tubing pulled away from the port. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the port was returned with tubing detached. Visual inspection of tubing showed no damage to tubing components. Visual inspection of the port showed a couple of tears. The sidewall of the port had a tear in it adjacent to the hole where the tubing would be installed. The damage made it possible for the tubing to exit through the tear instead of the bottom of the hole. There was also a. 510 long tear along the underside of the top rim. Engineering concluded that damage was likely due to mishandling/misuse. No other damage was found. General precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Always have a backup instrument available to complete the surgical procedure in case of instrument failure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SINGLESITE PORT FOR 85 MM ENDOSCOPE | ENDOSCOPIC INSTRUMENT ACCESSOR | 3128540 | 3100494 | 2997049 | 2955842201301562 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-10 | 2013-05-08 | No | Yes | Health Professional | 42806503 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | Yes | Initial | 2013
42022703 | 2013-04-08 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical pk dissecting forceps instrument cable was noted to be frayed in a loop. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument had a frayed a pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3128548 | 3100502 | 2989442 | 2955842201301563 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-10 | 2013-05-08 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-10 | Malfunction |  | Event Description: --/-- It was reported that during reprocessing the da vinci si surgical fenestrated bipolar instrument was noted to have frayed cables at the tip of the instrument. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was received for evaluation on (b)(4) 2013. The instrument has not yet been evaluated; therefore, the root cause of the customer reported failure mode has not been determined. A follow-up mdr will be submitted when the instrument has been evaluated, or if additional information is received. the customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. . --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was frayed at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. The instrument passed electrical continuity testing. There were scratches on the surface of the distal pulley. Engineering also observed the distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. The evidence was inconclusive, but the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | Reuse | 3128651 | 3100604 | 2997159 | 2955842201301564 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-05-06 | 2013-05-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-05-06 | Yes | No |  | 2013
42020505 | 2013-04-11 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was frayed at the distal clevis. No damage was found on the pulley or the clevis. Engineering also found corroded back idler pulleys, a bent bipolar pin, and a deep scratch on the main tube. Yellowish residue material was found on the back idler pulleys inside the housing. Engineering concluded the damage was likely due to improper cleaning. The bottom bipolar pin was bent at the back of the chassis. Engineering concluded the damage was likely due to mishandling. The distal end of the main tube had a scratch mark exhibiting light material removal and a rough surface finish. Engineering concluded that the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3128657 | 3100610 | 2997165 | 2955842201301565 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the mega suturecut needle driver instrument was noted to have a bent tip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering performed visual inspection, and found the tips were perfectly aligned. Engineering also found a broken grip cable. The one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. No other damage was found. Engineering also found an indentation on the edge of the distal pulley. Engineering concluded the damage was likely due to mishandling. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3128812 | 3100766 | 2997321 | 2955842201301566 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-08 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-17 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical fenestrated bipolar forceps instrument was noted to be frayed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was frayed at the distal idler. No damage found on the pulley or the clevis. The instrument passed electrical continuity testing. Engineering also found there were scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3128817 | 3100773 | 2997328 | 2955842201301568 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-11 | 2013-05-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-11 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-11 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure the mega suturecut needle driver instrument was noted to have a frayed cable on the distal end. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument had a broken grip cable. The one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at wrist were undamaged. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3128849 | 3100805 | 2997360 | 2955842201301567 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-12 | 2013-05-08 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-12 | Device Not Retu | No | Reuse | 2013
8MM | 2013-05-06 | Malfunction |  | Davinci mega suturecut needle driver broke during use. The wires at the tip of the instrument broke during the suturing. What was the original intended procedure?inguinal hernia repair. |  | MEGA SUTURECUT | SYSTEM SURGICAL | 3138475 | 3110507 | 3007062 | 3110507 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-05-08 | 2013-05-08 | No | Yes | Invalid Data | 8MM | Yes | No Answer Provided |  |  |  |  | 2013
400272 | 2013-05-06 | Malfunction |  | While utilizing harmonic ace curved shears for robot, bottom piece of scissors broke off of shears. Broken piece obtained and entire instrument removed from field. No harm to pt. |  | HARMONIC ACE CURVED SHEARS INSERT | ROBOTIC HARMONIC | 3140943 | 3111492 | 3008047 | MW5030114 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-05-08 | 2013-05-08 | No | Yes | Health Professional | 400272 | Yes | Yes |  |  |  |  | 2013
| 2013-04-26 | Malfunction |  | Davinci vessel sealer malfunctioned during surgery. Blade would not retract back. No harm was done to patient. Device was intact upon inspection. Manufacturer requests device(s)for evaluation. This is the 4th defective device. Local rep is supposed to pick up products for evaluation. What was the original intended procedure?gyn procedure. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST ONE | INSTRUMENT COMPUTER CONTROLLED | 3140853 | 3113354 | 3009910 | 3113354 | 1 | NAY | User Facility |  |  | Initial | 2013-05-08 | 2013-05-08 | No | Yes | Physician |  |  | No Answer Provided |  |  |  |  | 2013
42017207 | 2013-03-29 | Malfunction |  | It was reported that during reprocessing that a cable fell off the pulley on the maryland bipolar forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. Both grip cables on one side of distal clevis were derailed from the distal idler pulleys. The grip open cable was seated on the outermost pulley and the grip close cable was exposed on the outside of the pulley. Yaw motion was non-intuitive as a result. Other cables at the instruments wrist were not damaged. An additional observation not reported by site was tube damage. Distal end of the main tube had various scratch marks with light material removal. The scratches were. 095 -. 130 in length and not aligned with the tube axis. Engineering concluded the damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3124607 | 3096026 | 2992581 | 2955842201301554 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-07 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-04-09 | Death | Death | Event Description: --/-- It was reported that during a da vinci si pulmonary wedge resection procedure performed on (b)(6) 2013, the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. At the time of the conversion, the da vinci surgical system was undocked from the patient and the patient was reportedly in stable condition and was doing fine. At an unspecified time during the open surgical procedure, the patient subsequently expired. No additional clinical information was provided. | Manufacturer Narrative: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) was able to obtain the operative report concerning the reported event. Based on the operative report and in reference to a tumor, the surgeon noted, it also became apparent to me that since it invaded the phrenic nerve that in order to try to resect this for curative intent that the patient would be best served with an open procedure and i made a decision to convert from the da vinci robotic procedure to a posterolateral thoracotomy, so that i could place my hand inside the chest and directly visually and manually inspect and observe the extent of tumor invasion. It was also apparent that the medical aspect of the left upper lobe was quite friable and i felt that being able to place my hand directly on the lobe and retract postoperatively, so i could inspect the anteromedial mediastinum especially in the region of the hilum that this would be a favorable approach. In addition, as i dissected around the hilum and exposed the pulmonary vein draining the upper lobe it became apparent that the tumor or at least the very dense inflammatory reaction of the tumor that would have to be resected with the wedge resection was too close to the pulmonary vein to allow wedge resection because applying the stapler across the lung at this region would most likely interfere or involve the pulmonary vein. Therefore, the decision based on all of these factors was made to convert to thoracotomy. All of the robotic instruments were removed and the da vinci robot was moved away from the patient table in preparation for conversion to thoracotomy. The patient was completely stable at this point and i went out to the waiting room to advise the family on the fact that i was going to convert to thoracotomy and would proceed with upper lobectomy and this might take several more hours based on the operative findings including the invasion into the mediastinum. During the open thoracotomy procedure and in reference to the patient's pulmonary artery, the surgeon noted on the operative report: there was still a significant amount of inflammatory reaction and some lymph node tissue around this trunk, and as i dissected this in preparation for completing the lobectomy i encountered bleeding from the pulmonary artery at the level of the trunk and immediately gained control by placing a soft sponge stick over the site. The surgeon reportedly repaired the point of bleeding of the pulmonary artery using a 4-0 prolene suture. The surgeon noted, after repairing the artery, i removed the right-angle clamp that i had placed temporarily to control the bleeding and as i removed the clamp there was more bleeding directly from the pulmonary artery where this branch bifurcated, and again gained control by direct pressure. In addition, the surgeon indicated, i then inspected the area of the pulmonary artery where the tear was located and again used 4-0 prolene sutures avoiding any tension to repair the pulmonary artery at the site where the main trunk branch to the upper lobe which had been previously divided. This initially provided a good control and the prolene suture was tied snugly and appeared to be hemostatic but after i removed the clamp from the proximal artery this time i encountered very extensive bleeding directly from the pulmonary artery. Also, the surgeon indicated on the operative report, as my assistant held pressure over the pulmonary artery and we had good hemostatis i was able to enter the pericardium and easily visualize the main pulmonary artery intrapericardially. I had another long vascular clamp and i was able to place this across the proximal pulmonary artery in the pericardium under direct vision attempting to do this very cautiously. Unfortunately i encountered more bleeding as i was placing this clamp on the pulmonary artery trunk and this now has become a catastrophic pulmonary artery tear and again try to gain control by gently placing the clamp more proximally without success and it appeared that every small maneuver i made in order to try to gain control or repair of the artery lead to further bleeding. At this point, there was massive bleeding from the pulmonary artery such that the patient developed profound hypertension and ultimately lost blood pressure. I opened up the pericardium widely and placed a clamp as proximally as i could on the left pulmonary artery and all efforts were made to volume resuscitate the patent. I had 4 units of packed red blood cells in the room while we had control prior to this catastrophic event in preparation of very large serious bleeding but yet this happened very quickly and we were not able to resuscitate the patient. Based on the information provided in the operative report, there is no evidence that a malfunction of the da vinci system, an instrument, or an accessory occurred during the reported event. In addition, there is no evidence that the da vinci system, an instrument, or an accessory caused or contributed to the patient's injury to the pulmonary artery and subsequent demise. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2013, isi contacted the clinical sales representative (csr) who initially reported this event. The csr indicated that he was present during the da vinci surgical procedure and that the surgeon's decision to covert to open surgical techniques was due to the size of the patient's lesion. The surgeon found that the lesion was larger than he expected and that the surgeon believed that the lesion was too complex to be removed robotically. The csr indicated that during his attendance of da vinci surgical procedure he did not witness any malfunction of the da vinci surgical system, instruments or accessories while they were in use. The csr indicated that there was no allegation by the site that the da vinci si surgical system, instruments, or accessories caused or contributed to the patient's demise. Per the csr, he was not present during the open procedure and that the cause of the patient's demise was not provided to him by the site. The csr indicated that he was told by the site's vice-president that the medical examiner declined to investigate the reported event to determine the cause of the patient's demise. On (b)(4) 2013, isi contacted the site's risk management department. The risk manager declined to provide a cause of death for the patient or additional information regarding the reported event. Isi's review of the site's system logs for the reported procedure date found no system errors were generated during the surgical procedure that would have caused or contributed to the patient's demise. | DA VINCI SI SURGICAL SYSTEM | Reuse | 3124871 | 3096412 | 2992967 | 2955842201301557 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2013-04-09 | 2013-05-07 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-04-09 | Device Not Retu | No |  | 2013
42030901 | 2013-04-15 | Malfunction |  | During a da vinci si hysterectomy procedure, the user facility allegedly identified a broken separated wire on the mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. Engineering confirmed a broken grip close cable at the distal idlers, which was sticking out at the instrument's wrist. The idler pulley spun freely and was not damaged. Other cables at the instrument's wrist were not damaged. Additional damage found were scratches on the surface of the distal pulley. Engineering was unable to conclude how the pulley damage occurred. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3125889 | 3097496 | 2994049 | 2955842201301558 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-24 | 2013-05-07 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-19 | Malfunction |  | During central processing, the user facility identified a broken separated cable on the mega suturecut needle driver instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was received and evaluated. Engineering confirmed a broken grip close cable at the distal idlers, which was sticking out at the instrument's wrist. The idler pulley spun freely and was not damaged. Other cables at the instrument's wrist were not damaged. Additional damage found were scratches on the surface of the distal pulley. Engineering was unable to conclude how the pulley damage occurred. There were also various scratch marks on the main tube showing light material removal. The scratches were short in length and were not axially aligned with the main tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments the customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3125911 | 3097511 | 2994066 | 2955842201301559 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-05-07 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42000207 | 2013-02-25 | Malfunction |  | It was reported that prior to starting a da vinci s prostatectomy procedure the instrument cannula was checked using the gauge pin. As a result, the gauge pin got stuck in the cannula. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found a hairline indentation near the tip of the cannula shaft, approximately 0. 90 in length. | 8 MM INSTRUMENT CANNULA | ENDOSCOPIC INSTRUMENT ACCESSOR | 3125991 | 3097579 | 2994132 | 2955842201301560 | 1 | NAY | Manufacturer | DistributorHealth Professional | Physician | Initial | 2013-02-26 | 2013-05-07 | No | Yes | Health Professional | 42000207 | Device Returned To Manufacturer | Yes | 2013-02-26 | Device Not Retu | No | Reuse | 2013
| 2013-04-26 | Malfunction |  | (b)(6) underwent the following surgical procedure - robot-assisted total laparoscopic hysterectomy bilateral salpingo-oophorectomy and laparoscopic appendectomy. The pk da vinci dissecting forcep grasper portion of the instrument broke on the jaw of the grasper and one side of the grasper fell free into the abdomen. The 1cm x 5mm piece was successfully retrieved. A postoperative abdominal x-ray was performed to confirm foreign body was removed completely. This device was on its 1st of ten "lives. ". |  | DA VINCI SI SURGICAL SYSTEM PK DISSECTING FORCEP 8MM | PK DISSECTING FORCEP | 3142047 | 3111646 | 3008201 | MW5030124 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-05-07 | 2013-05-07 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2013
42009308 | 2013-04-11 | Malfunction |  | After the da vinci si prostatectomy procedure, the surgical staff found a wire that came off from the pulley of the prograsp forceps instrument when the staff irrigated it. During the procedure, the movement of the instrument was normal. There was no allegation of patient harm or fragments falling into a patient. | The instrument was received and evaluated. Engineering found a derailed grip cable at the distal idler pulley. The grips could still be opened and closed, but movement may not be precise. Additional damage found was deep scratches on the main tube. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage was likely due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3122625 | 3093972 | 2990526 | 2955842201301553 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-05-01 | 2013-05-06 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-29 | Malfunction |  | Event Description: --/-- It was reported that during reprocessing the monopolar curved scissors instrument had a frayed cable. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. These cracks were reported on the initial mdr report. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch up cable was frayed near the proximal clevis cable hole. The frayed strands stuck out at the instruments wrist. The other cables at the wrist were not damaged. An additional observation not reported by the site was a cracked tube. The distal end of the main tube exhibited a couple of hairline cracks parallel to each other, oriented in the axial direction. The cracks overlapped with the section of tube extension that is installed inside the main tube. There was a small bulge in the tube diameter near the cracks. The instrument was disassembled and the cracks were found to align closely with each side of an axial key slot. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3123750 | 3094920 | 2991474 | 2955842201301555 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-04-09 | 2013-05-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
IS2000 A51P8 | 2010-09-01 | Injury | Required InterventionHospitalization | On (b)(4) 2012, isi received a legal complaint alleging that a patient whom underwent a da vinci s nissen fundoplication on (b)(6) 2010 at (b)(6), experienced severe abdominal pain. The legal complaint alleges that on (b)(6) 2010, the patient underwent exploratory laparoscopy and esophagogastroduodenoscopy procedures and it was discovered that the patient's distal esophagus had a perforation. Allegedly, the patient was transferred to the university of iowa hospitals and clinics for further medical management. Allegedly, the patient underwent multiple surgeries to repair the esophageal perforation and the patient required an extended course of treatment due to the severity of the injury. | Isi has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided. A follow-up medwatch report will be submitted to the fda if additional information is received. A review of the patient's operative report for the da vinci s nissen fundoplication procedure performed on (b)(6) 2010 found no indication that the patient experienced any complications during the surgical procedure and there was no indication that any malfunction of the site's da vinci s surgical system, instruments and/or accessories occurred. The patient's operative report stated that the patient complained of significant reflux disease, which caused limitations in the patient's routine lifestyle and due to the patient's condition, the patient was offered the robotic nissen fundoplication procedure. The operative report indicated that all risks, benefits, and alternatives of the procedure were thoroughly discussed with the patient prior to the patient's agreement to the surgery. The patient's consent was obtained prior to the surgical procedure. The operative report indicated that the patient's hernia sac was taken down and a large hiatal hernia defect was appreciated. The patient's hernia sac was completely reduced and excised. The report indicated that initially the patient experienced some bleeding along the undersurface of the liver due to a slight tear that occurred as a result of the retraction of the patient's liver. The bleeding was controlled with the application of pressure and surgicel. It was indicated that the surgeon proceeded to wrap the greater curve of the patient's stomach around the posterior aspect of the ge junction and then repaired the patient's hernia defect using interrupted 2-0 silk sutures in a figure-of-eight fashion to re-approximate the patient's crus muscles. Two posterior and one anterior to the patient's esophagus were placed. It was indicated that a loose wrap was appreciated and the surgeon secured the wrap with interrupted 2-0 silk sutures and after completion of the wrap, his reexamination of the surgical field found that all of the sutures appeared to be stable. Per the information indicated in the patient's operative report, the patient awoke from anesthesia and was taken to the recovery room in stable condition. The patient tolerated the procedure well and there were no complications. Isi's review of the site's system logs for procedures performed on the reported procedure date found that there were no occurrences of any system errors during any of the procedures performed that caused or contributed to the alleged injury sustained by the patient. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3123774 | 3094953 | 2991507 | 2955842201301556 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-09-05 | 2013-05-06 | Yes | No | Health Professional | IS2000 A51P8 | No | No | 2012-09-05 | No | No | Reuse | 2013
400272 | 2013-04-08 | Malfunction |  | Harmonic tip broke off, but was retrieved. Device #1is this a laboratory device or laboratory test?no. |  | HARMONIC ACE | SYSTEM SURGICAL COMPUTER CONTR | 3140807 | 3113346 | 3009902 | 3113346 | 1 | NAY | User Facility |  |  | Initial | 2013-05-06 | 2013-05-06 | No | Yes | Physician | 400272 | No | No Answer Provided |  |  |  |  | 2013
42009308 | 2013-04-08 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted broken wires at the tip of the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Engineering also found indentations on the edge of the distal pulley. Engineering concluded that damage was likely due to mishandling. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3119478 | 3091028 | 2987582 | 2955842201301524 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-24 | Malfunction |  | During a da vinci si sacrocolpopexy procedure, the wires on the mega suturecut needle driver reportedly broke. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint: the grip cable was frayed at the distal idler pulley and was sticking out at the instrument's wrist. Additional observations were found that were not initially reported by the customer. One grip close cable was derailed at the distal idler pulley. The grips can still open and close but movement may not be precise. The cable derailment was likely due to cable losing contact with the pulley during wrist articulation. The fleet angle between the proximal and distal pulleys may have contributed to the cable derailment. The scratches on the distal pulley looked like they came from the cable. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3120092 | 3092029 | 2988583 | 2955842201301512 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-25 | 2013-05-03 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-21 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, customer reported that while they were almost done suturing the patient's uterus, the surgeon noticed that the strings on the maryland bipolar forceps tip snapped making the endowrist instrument defective. The procedure went well and with no incident to the patient. The patient was transferred to the recovery room. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3120431 | 3092240 | 2988794 | 2955842201301525 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-05 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure that a cable was protruding on the distal end of the large needle driver instrument. The instrument was not used. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely and does not exhibit damage. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3120482 | 3092250 | 2988804 | 2955842201301527 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-05 | 2013-05-03 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-11 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the wrist doesn't work because of visible damaged connective rod on the prograsp forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3120442 | 3092267 | 2988821 | 2955842201301526 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-05 | 2013-05-03 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-02 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, customer reported that while they were suturing the patient's uterus, the surgeon noticed that the strings of the maryland bipolar forceps instrument were detached from the tip. The procedure went well and with no incident to the patient. The patient was transferred to the recovery room. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable located at distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 04 in length. No other cable damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3120655 | 3092341 | 2988895 | 2955842201301528 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-02-27 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer reported that while the surgeon was suturing, the wire of the mega needle driver instrument snapped off. The instrument was replaced and the surgical procedure was completed as planned. The procedure went well and with no incident to the patient. The patient was transferred to the recovery room. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3120740 | 3092393 | 2988947 | 2955842201301530 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-06 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the maryland bipolar forceps instrument stopped reacting properly. The surgeon detected inaccurate movements there were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was returned with both pitch cables broken at the distal clevis hub. Both cable segments that contain the crimp were missing. The clevis did not exhibit any damage or wear marks. The other cables at the instruments wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3120731 | 3092414 | 2988968 | 2955842201301529 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42018304 | 2013-04-05 | Malfunction |  | The permanent cautery hook was returned without an initial complaint. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The permanent cautery hook instrument was returned without an initial complaint from the customer. During failure analysis, engineering observed that the instrument was returned with a dislodged luer plate, cracked proximal clevis, and main insulation tube damage. Engineering observed that the luer plate was dislodged from the chassis and sticking out past the chassis back plane. A section of the chassis outer wall feature which holds the luer plate was broken off. Engineering concluded that the wall feature likely broke due to improper cleaning, allowing the luer plate to dislodge. Engineering also observed that the proximal clevis was cracked in the axial direction, halfway between the clevis ears. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This the endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3120980 | 3092581 | 2989135 | 2955842201301531 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42032702 | 2013-03-20 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure the tips locked shut on the medium-large clip applier instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was placed on system and driven. Recognition and engagement test passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Instrument was fully functional. An additional finding not reported was a scratch on the main tube. The distal end of the main tube had a scratch mark showing material removal. The scratch was aligned with the tube. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEDIUMLARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3121076 | 3092652 | 2989206 | 2955842201301532 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42032702 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
40018012 | 2013-03-19 | Malfunction |  | Event Description: --/-- During a da vinci s surgical procedure, the user facility identified slits on the tip cover used with the monopolar curved scissors instrument. The reporter stated that the tip cover installation tool was used when installing the tip cover and that the tip cover was inspected prior to use. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument accessory was returned and evaluated. Engineering confirmed the alleged complaint. Visual inspection of the tip cover showed a couple of tears on one side of the combined silicone/pellethane section. Tears were roughly 0. 050 long, located around 0. 430 below the distal end opening. Tears did not go through the entire wall thickness of the section. Additional observation that was not initially reported by the customer was arcing damage. The combined silicone/pellethane section had a hole burned though it. The hole exhibited localized melting, indicative of arcing. The hole measured roughly 0. 035 x 0. 015 and was located 0. 600 below the distal end of the tip cover. One edge of the hole appeared to have a tear. On (b)(4) 2013, intuitive surgical (isi) spoke with the user site's robotics coordinator. She was unable to provide information about what events led to the damaged tip cover at the time of the call. Isi made several additional attempts to the customer to obtain additional information but had been unable to reach the robotics coordinator. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- Follow-up 1- additional information: on (b)(4) 2013, an intuitive surgical representative spoke with the site's robotics coordinator. She reported that the damaged tip cover was actually discovered when the instrument was being removed from the patient. The surgical staff did not notice any arcing events during the da vinci si hysterectomy surgical procedure and confirmed there was no patient harm during and after the surgical procedure. She was unable to obtain any other information. | TIP COVER ACCESSORY | Reuse | 3121085 | 3092663 | 2989217 | 2955842201301533 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2013-04-03 | 2013-05-03 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No |  | 2013
40018012 | 2013-03-22 | Malfunction |  | During a da vinci s surgical procedure, the user facility identified slits on the tip cover used with the monopolar curved scissors instrument. The reporter stated that the tip cover installation tool was used when installing the tip cover and that the tip cover was inspected prior to use. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. Visual inspection of tip cover showed a couple of axially aligned tears at the distal end opening and one 0. 070 long tear through the combined silicone/pellethane section that is located around 0. 395 below the distal end opening. Additional observation that was not initially reported by the customer was arcing damage. The combined silicone/pellethane section had a hole burned though it. The hole exhibited localized melting, indicative of arcing. The hole measured roughly 0. 010 in diameter and was located 0. 415 below the distal end of the tip cover. One edge of the hole appeared to have a tear. On (b)(4) 2013, intuitive surgical (isi) spoke with the user site's robotics coordinator. She was unable to provide information about what events led to the damaged tip cover at the time of the call. Isi made several additional attempts to the customer to obtain additional information but had been unable to reach the robotics coordinator. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3121117 | 3092670 | 2989217 | 2955842201301534 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-03 | 2013-05-03 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-02-13 | Malfunction |  | It was reported that during a da vinci si gastroplasty the small grasping retractor instrument tips clamped down and would not open. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was placed on a system and driven. Recognition and engagement test passed. The instrument moved intuitively with a full range of motion in all directions. The clamp opened and closed properly. The instrument was fully functional. An additional finding not reported were various scratches on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3121126 | 3092675 | 2989229 | 2955842201301535 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-08 | Malfunction |  | It was reported that during reprocessing the fenestrated bipolar forceps instrument had a frayed cable at the tip. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, the pitch cable was broken at distal end. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the instruments wrist were not damaged. Additional damage found were deep scratch and mechanical indentation. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage may be due to mishandling. The distal idler pulley had a mechanical indentation at the edge. The pulley still rotated freely and the grip cables were not affected by the damaged pulley. Engineering concluded the damage was likely due to mishandling. Electrical continuity test passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3121180 | 3092718 | 2989296 | 2955842201301536 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42020505 | Yes | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-08 | Malfunction |  | It was reported that during reprocessing a cable was frayed at the tip of the fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. For clarification, the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3121202 | 3092742 | 2989296 | 2955842201301538 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-28 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the customer noted that the monopolar curved scissors were dull. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on da vinci robot system for a latex cut test. The instrument did not cut cleanly through (b)(4) latex. Latex got snagged at scissor tips. The blade edges were not damaged, but the edges exhibited wear at the tips, negatively affecting the cut performance. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3121310 | 3092844 | 2989398 | 2955842201301537 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-04-05 | Other | Other | It was reported that during set-up of a da vinci si pyeloplasty procedure, the surgical staff experienced system error code 25580. With the assistance of an isi technical support engineer (tse) the site powered off and emergency powered on the system multiple times, however, the issue persisted. The patient was under anesthesia and the port placements were not yet created when the surgeon made the decision to abort the planned surgical procedure and reschedule it for a later date. No patient harm, adverse outcome, or injury was reported. | The investigation conducted by field service engineering concluded that the system error code 25580 experienced by the customer was associated with a remote arm controller (rac) pca board. The rac consists of five printed circuit assembly boards (pca) which operate together to provide control of the system arms. The system was repaired by replacing the affected rac pca board. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25580 is reported when the da vinci si safety system determines a hardware fault reaction logic occurred on a local rac. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The remote arm controller (rac) board was returned and evaluated. Engineering was unable to replicate the issue experienced by the site. The rac passed functional testing with no issues noted. On (b)(4) 2013, isi contacted the initial reporter regarding the reported event. The initial reporter indicated that there were no complications experienced by the patient as a result of the reported event. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3121324 | 3092858 | 2989412 | 2955842201301540 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-05 | 2013-05-03 | Yes | No | Health Professional | IS3000 A60P8 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2013-03-26 | Death | Death | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the surgeon made the decision to convert to open surgical techniques for an unknown reason and the patient subsequently expired. | Manufacturer Narrative: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) contacted the surgeon involved with this complaint and obtained clarification regarding the reported event. The surgeon indicated that near the end of the case and during application of a hemostatic agent for sealing, he had the needle driver instrument in his right hand and the grasper instrument in his left hand. The surgeon's right hand was out of view and the patient side assistant's 3rd-party laparoscopic suction irrigator instrument was out of view briefly while the camera was moved cephalad to view a different area of the patient's anatomy. At that time, the surgeon noticed pooling blood. When the surgeon moved the camera back to look at the pelvic area, significant bleeding was seen. The surgeon attempted to control the bleeding for 4-5 minutes but was unsuccessful. A vascular surgeon was called in and the case was converted to open traditional surgical techniques. The vascular surgeon repaired the iliac vein and closed the patient. Approximately 1 hour later, the patient was not stable and a secondary operation was performed for further repairs at the same site of the previous bleeding. The details of the secondary operation were not provided. The patient passed away 2 days later due to multiple complications likely due to blood loss. Isi has reviewed the system logs for this procedure and confirmed that no da vinci system errors were associated with the injury to the iliac vein. At this time, there is no evidence that the da vinci si system, or a da vinci instrument or accessory caused or contributed to the patient's injury to the vessel. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2013, isi contacted (b)(6) who initially reported this event. The csm indicated that the surgeon who performed the surgical procedure reported to him that during a da vinci si hysterectomy procedure, performed on (b)(6) 2013, the bedside assistant inadvertently pushed an unknown endowrist instrument installed on a patient side manipulator (psm) arm into the patient's left common iliac vein. The damage to the patient's vessel caused uncontrolled bleeding and approximately 4 minutes after the intra-operative injury occurred the surgeon made the decision to convert the procedure to open surgical techniques to achieve hemostasis. Per the csm, the surgeon indicated to him that no malfunction of the da vinci si surgical system, instruments, or accessories occurred during the surgical procedure that caused or contributed to the injury sustained by the patient. The specific instrument installed on the psm at the time of the patient's injury was not provided to the csm. The csm indicated that he was told by the surgeon that the patient experienced significant blood loss; however, hemostasis was achieved after he (surgeon) converted the procedure to open surgical techniques. The patient expired on (b)(6) 2013 as a result of significant blood loss and multiple organ failure. Per the csr the surgical procedure was recorded by the site; however, the video was not available for review. The csr also indicated that the surgeon told him that a private pathologist conducted an autopsy; however, the autopsy report was not available for review. On (b)(6) 2013, the isi csm assigned to the site also spoke directly to the surgeon. The surgeon informed the csm that the patient underwent a da vinci si radical hysterectomy and that the procedure was completed robotically. The csm indicated the surgeon told him that the patient was stable upon closure at the end of the operation and that the surgeon denied that a malfunction of the da vinci surgical system, instruments, or accessories occurred during the surgical procedure. The patient was reportedly discharged from the hospital on an unspecified date; however, a few days after the patient was discharged from the hospital, the patient returned to the hospital and it was determined that the patient required an open surgical procedure. During the open surgical procedure, the surgeon discovered that the patient had experienced internal bleeding internally for an unknown period of time. The csm indicated that the surgeon reviewed the da vinci procedure video. The surgeon indicated that based on his review of the video, the video did not show that the patient experienced any significant internal bleeding during the da vinci si hysterectomy procedure. The surgeon indicated to the csm that review of the video did not show that any malfunction of the da vinci surgical system, instruments, or accessories occurred. The csm was told by the surgeon that he believes that the internal bleeding experienced by the patient was the result of damage to the patient's left common iliac vein, that occurred after the bedside assistant inadvertently pushed a laparoscopic suction irrigator instrument installed in an assist port into the patient's vessel. On (b)(6) 2013, isi contacted the risk management department at the site. The risk manager indicated that she was unable to disclose any information regarding the cause of the patient's demise and would not be able to provide the patient's medical records. The risk manager indicated that the da vinci surgical system did not cause or contribute to the patient's demise. Isi reviewed the site's system logs for the reported procedure date, showed that there was no malfunction of the site's da vinci surgical system, instruments and/or accessories that would have likely caused or contributed to the reported event. In addition, the site's system logs showed that there was no use of an isi suction irrigator instrument during the surgical procedure. Based on the additional information provided, there is no evidence that the da vinci si surgical system, instruments, or accessories malfunctioned in a way that caused or contributed to the patient's demise. | DA VINCI SI SURGICAL SYSTEM | Reuse | 3121302 | 3092868 | 2989422 | 2955842201301539 | 1 | NAY | Manufacturer | Company RepresentativeHealth Professiona | Other | InitialFollowup | 2013-04-04 | 2013-05-03 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2013-04-04 | No | No |  | 2013
400180 |  | Injury | Other | On (b)(6) 2013, isi became aware of a case report in which a patient who underwent a da vinci prostatectomy procedure sustained injuries. The case report titled iliac vein injury due to a damaged hot shears tip cover during robot assisted radical prostatectomy was cited in a study that was published by the department of urology at loma linda university school of medicine in loma linda, ca titled a prospective analysis of robotic tip cover accessory failure. Reportedly during the surgical procedure, while the surgeon was performing dissection of the patient's contra lateral obturator lymph nodes, he noted that the patient was bleeding from the obturator vein. The surgeon used bipolar coagulation to control the patient's bleeding. Per the information indicated in the case report article, the surgeon initially believed that bleeding was caused by a surgical error. The case report article alleges that upon further dissection of the nodes from the obturator nerve with the monopolar curved scissors (mcs) instrument with the mcs tip cover installed, a small perforation was noted on the right external iliac vein. It was noted by the surgeon that there was a safe distance between the patient's vessel and the mcs instrument. The surgeon repaired the patient's iliac vein using a 5mm metal clip. Per the information indicated in the case report article, inspection of the tip cover accessory after completion of the surgical procedure revealed that the tip cover had two 1 mm holes. The patient was discharged from the hospital 3 days post-op. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Isi has made several attempts to contact the author of the case report to obtain additional information concerning the reported event; however, no additional information has been provided. A follow-up mdr will be submitted if the tip cover is returned (post engineering evaluation) or if additional information is received. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 3121353 | 3092887 | 2989441 | 2955842201301542 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-04-03 | 2013-05-03 | Yes | Yes | Service Personnel | 400180 | No | No | 2013-04-03 | No | Yes | Unkown | 2013
42022703 | 2013-04-08 | Malfunction |  | It was reported that after a da vinci si surgical procedure a wire cable at the tip of the pk dissecting forceps instrument was broken. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. Clevis did not exhibit any damage or wear marks. Electrical continuity test passed. Additional findings not reported were scratches on the surface of the distal pulley. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3121354 | 3092888 | 2989442 | 2955842201301541 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-08 | Malfunction |  | It was reported that during reprocessing fenestrated bipolar forceps instrument a cable was broken and the instrument was not working properly. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported a broken wire, however, for clarification the damaged instrument component was a frayed pitch cable. The instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. Electrical continuity passed. Additional findings not reported was a pitch cable was loose at distal clevis hub. Both cable crimps remained in the clevis. Upon housing removal a pitch cable was found loose at the clamping pulley, but no loose clamping pulley screw was found. There were also scratches on the surface of the distal pulley. The clamping pulleys also exhibited corrosion around the pulleys. Engineering concluded the failure may be cleaning related. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3121369 | 3092907 | 2989296 | 2955842201301543 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-19 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure, the customer noted that the monopolar curved scissors were dull. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. These cracks were reported on the initial mdr report. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the failure mode. Engineering placed the instrument on the da vinci robot system for a latex cut test. The instrument was able to cut cleanly through (b)(4) latex. The blades were not damaged. Instrument passed electrical continuity test. Additional observation not reported by site was a crack on the main tube. Engineering observed that the instrument exhibited a crack on the extension tube side about 0. 350 from the extension tube. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3121340 | 3092909 | 2975341 | 2955842201301544 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-04-08 | 2013-05-03 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
IS3000 A60P4 | 2012-07-05 | No Answer Provided |  | On (b)(6) 2013, intuitive surgical received a legal complaint alleging that a patient who underwent a da vinci si surgical procedure on (b)(6) 2012, expired on (b)(6) 2012. The legal complaint alleges that from (b)(6) 2012, the patient required care and treatment. | Intuitive surgical has made several attempts to obtain additional information concerning the reported event; however, no additional information has been provided. On (b)(6) 2013, the site's risk manager indicated that due to (b)(6) regulations, the hospital is unable to provide additional information concerning the reported event without a subpoena. On (b)(6) 2013, isi contacted the isi clinical sales representative (csr) for the hospital. The csr indicated that she was not the csr for the hospital when the reported event occurred. Approximately two months ago, a physician at the hospital reported to her that an adverse event concerning a da vinci gyn procedure had occurred at the hospital. Per the csr, the physician indicated to her that the procedure did not go well. No other details concerning the reported event were provided to the csr. The csr indicated that due to how long ago the incident occurred, it was her belief that the incident had already been reported to isi. Isi performed a review of the site's system log for the reported procedure date. Review of the system logs found that no malfunction of the da vinci surgical system, instruments or accessories occurred during the surgical procedure. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3121425 | 3092969 | 2989523 | 2955842201301545 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-01 | 2013-05-03 | Yes | No | Health Professional | IS3000 A60P4 | No | No |  | No | No | Reuse | 2013
42020505 | 2013-03-01 | Malfunction |  | During a da vinci si hysterectomy procedure, the fenestrated bipolar forceps instrument reportedly would not close. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. The instrument had a piece removed from the yaw pulley. The damaged yaw pulley allowed the jaw grip to overextend the motion and the jaws not closing properly. Additional observations were found that were not initially reported by the customer. The instrument's jaws were bent, causing side to side misalignment of the grips. There was a 0. 03 offset at the tips. The bent grip had separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. Engineering evaluation concluded that the damages were likely due to mishandling. Electrical continuity testing was performed and passed. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3121451 | 3092977 | 2985119 | 2955842201301546 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-30 | 2013-05-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-04-30 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2013-03-25 | Death |  | It was reported that the patient underwent a da vinci si hysterectomy procedure on (b)(6) 2013. Reportedly 36 hours after successful completion of the surgical procedure, it was discovered that the patient's bowel was punctured. It was reported that the patient expired. | On (b)(6) 2013, isi contacted the csr who initially reported this event. The csr indicated that he was present during the surgical procedure. Per the csr the patient had significant adhesions as a result of a previous open surgical procedure and that the patient's adhesions were taken down using traditional laparoscopic techniques. The surgeon used a 12mm excel ethicon bladeless trocar to create the umbilicus port site. The csr indicated that there was no malfunction of the site's da vinci surgical system, instruments and or accessories during the robotic portion of the surgical procedure and the robotic procedure was successfully completed. The csr indicated that the surgeon reported to him that 36 hours post-op the patient presented symptoms of fever and sepsis. An emergency laparotomy procedure performed on the patient revealed that the patient's colon had a hole. Per the csr, the surgeon told him that the size of the hole was the same size as the trocar that was used during creation of the patient's umbilicus port site. The csr indicated that the surgeon believes that damage to the patient's colon occurred during creation of the umbilicus port site. The date of the patient's demise is unknown. Isi has contacted the site's risk management department several times to obtain additional information concerning the reported event, however, no additional information has been provided. A follow-up medwatch report will be submitted to the fda if additional information is received. Based on the information provided, it was determined that the da vinci surgical system did not cause or contribute to the patient's demise, but likely occurred as a result of the injury sustained by the patient during creation of the patient's umbilicus port site by the surgeon while utilizing a ethicon bladeless trocar. Isi has notified the manufacturer of the trocar regarding this report. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3121435 | 3092988 | 2965815 | 2955842201301547 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-04-05 | 2013-05-03 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2013-04-05 | No | No | Reuse | 2013
42017910 | 2013-02-27 | Injury | Other | On (b)(4) 2013, intuitive surgical received uf/importer report number (b)(4) from the fda concerning a patient injury that occurred during a da vinci si myomectomy procedure. The event details are provided below: large fibroid removed from uterus - fibroid being held with da vinci tenaculum. Monopolar curved scissors used to cut fibroid in half. Tenaculum and scissor never touched and kept good distance apart while cutting. Having trouble getting scissor to work and burn noted to peritoneum. No damage to bowel or bladder - no other procedure needed. During time monopolar scissor appeared to be not working it apparently arced and energy transferred to tenaculum. Tenaculum near the joint of the instrument is what caused the burn on the peritoneum. No one saw it arc - it was assumed by staff and doctor that it arced. Possibly from dissecting adhesions. No visible damage or imperfections noted. Insulation intact. | The monopolar curved scissors instrument has not been returned to isi for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Isi has made several attempts to contact the site concerning the reported event; however, no additional information has been provided. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3121478 | 3092995 | 2989549 | 2955842201301549 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-04-08 | 2013-05-03 | Yes | Yes | Health Professional | 42017910 | No | No | 2013-04-08 | No | No | Unkown | 2013
IS3000 A60P6 | 2013-04-08 | Other | Other | It was reported that during a da vinci si lobectomy procedure, the site experienced a non-recoverable system error code 252. Prior to contacting isi for technical support assistance, the surgeon made the decision complete the planned surgical procedure using open surgical techniques. With the assistance of an isi technical support engineer (tse) the site reviewed their system logs. The system logs showed that system error code 308 also occurred and was associated with the surgeon side console touchpad. The tse instructed the site to power down the system, cycle power the circuit breaker and power the system back on. The system powered on successfully. During a site visit on (b)(4) 2013 by the site's isi clinical sales representative (csr), system error code 252 occurred after the csr powered on the site's da vinci surgical system. No patient harm was reported. | The investigation conducted by isi field service engineering concluded that the non-recoverable system error code 252 and system error code 308 experienced by the customer were associated with the surgeon side console touchpad. The surgeon side console touchpad is located in the middle of the surgeon console armrest and provides the means for selecting various system functions such as system status including instrument arms, camera arm, and energy controls. The system was repaired by replacing the affected touchpad. The system alarms (system generated fault codes) functioned as designed and there was no injury to the patient. Upon determining this condition, the system records the fault and transitions to a safe state. The touchpad was returned for failure analysis evaluation. Engineering was unable to replicate the reported failure mode experienced by the customer. The touchpad passed functional testing on an isi in-house surgeon side cart test system. The touchpad passed a power cycle test (b)(4) and there were no system error codes generated. Engineering evaluation also found that the touchpad's system operation module (som) and cable were almost touching the sheet metal on the touchpad. As a precaution the touchpad has been returned to the original engineering manufacturer to have the som replaced. On (b)(4) 2013, isi contacted the initial reporter regarding the reported event. The initial reporter indicated that the da vinci surgical system was docked to the patient when the system error code 252 occurred. The initial reporter indicated that to the best of her knowledge the patient tolerated the open procedure well and that the patient has not returned to the hospital due to experiencing any post-surgical complications. As of (b)(4) 2013, there have been no reported recurrences of this failure mode at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3121477 | 3093011 | 2989565 | 2955842201301548 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-08 | 2013-05-03 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-05 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the cable on the monoplar curved scissors (mcs) instrument jumped. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported complaint with the following clarification that the cable at the distal idler pulley was frayed. The frayed cable section was approximately 0. 02 in length. The pulley did not exhibit any damage. Engineering evaluation also found that the distal end of the main tube had a hairline crack and multiple deep scratches with some materials removed. On (b)(4) 2013, isi contacted the site's isi clinical sales representative. The csr indicated that he followed up with the site to obtain additional details concerning the reported event. The csr indicated that the site was unable to provide any additional information, as they could not recall the details of the reported event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3121479 | 3093012 | 2989566 | 2955842201301550 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-09 | 2013-05-03 | Yes | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2013-04-10 | Other | Other | It was reported that during a da vinci s sacrocolpopexy procedure, when the system was powered on, the site experienced system error code 23000. With the assistance of an isi technical support engineer (tse) the site performed an emergency power off (epo) of the system and cycled the system breakers; however, system error code 23000 recurred. Review of the site's system logs by the tse found that the system error code 23000 experienced by the site was associated with the left master tool manipulator (mtml) on the surgeon side cart. The tse instructed the site to perform additional troubleshooting techniques; however, the issue persisted. The patient was under anesthesia and the port incisions had been created when the surgeon made the decision to abort the planned surgical procedure and rescheduled it for a later date. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 23000 was associated with the left master tool manipulator (mtml) gimbal. The mtm gimbal is a subsection of the complete mtm arm, consisting of the links associated with the five axes of motion closest to the surgeon's hand. These axes measure the orientation of the mtm handle and the mtm grip angle. The system was repaired by replacing the affected mtml gimbal. The system alarm (system generated fault code) functioned as designed. System error code 23000 occurs when the da vinci safety system determines the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the system records the fault and transitions to a safe state. The mtml gimbal was returned to intuitive surgical for failure analysis investigation. Engineering was able to replicate the reported failure during internal testing and determined that the mtml gimbal's axis 6 potentiometer had malfunctioned, thus generating the system error code experienced by the customer. Engineering evaluation also found that the hard pin on the axis 6 link where the axis 7 link is attached had slipped lower into the housing, which allowed the axis 7 link to go beyond the calibrated hard stop. The axis 6 was replaced to repair the mtml gimbal. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3121464 | 3093014 | 2989568 | 2955842201301551 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-10 | 2013-05-03 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-04-09 | Other | Other | It was reported that during a da vinci si pyeloplasty procedure, the site experienced system error code 22002. The site contacted isi for technical support engineering (tse) assistance. The tse determined that the system error code 22002 experienced by the site was associated with the endoscopic camera manipulator (ecm) arm and was caused by out-of-range readings of a secondary setup joint potentiometer. With the assistance of the tse the site changed the position of the endoscopic camera manipulator (ecm) and restarted the system; however, the issue recurred. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. | The investigation conducted by the field service engineer confirmed that the site experienced system error code 22002 and confirmed that the error code experienced by the site was associated with the 2 potentiometer within the setup joint (suj) of the endoscopic camera manipulators (ecm). The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The setup joints (suj's) are passive manipulator joints that are part of setup structure and are used for pre-operative placement of manipulator arms. Each setup joint has a primary and a secondary (redundant) potentiometer. They are located within the setup joint structure of the robotic arm on the psc and monitor the movement of the setup joint. The system was repaired by replacing the affected potentiometer. The system alarm (the 22002 system error code) functioned as designed and there was no injury to the patient. System error code 22002 appears when the da vinci si safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by the joint's primary control sensor (encoder) and the secondary sensor (potentiometer) were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci si in a nonrecoverable safe state. On april 17, 2013, the isi field service engineer obtained additional information from the site regarding the reported event. The fse indicated that the site reported to him that there were no system issues noted prior to anesthetizing the patient and that the reported issue occurred 40 minutes into the surgical procedure. The fse indicated that the site reported to him that the patient tolerated the open procedure without any problems and that the patient is recovering quickly. The fse indicated that the site reported that the patient has been released from the hospital and has not returned to the hospital due to experiencing any post-surgical complications as a result of the reported event. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3121483 | 3093015 | 2989569 | 2955842201301552 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-09 | 2013-05-03 | Yes | Yes | Health Professional | IS3000 A60P8 | Yes | Yes | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-05 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the pk dissecting forceps instrument was noted to have frayed cables. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was frayed at the distal idler. No damage was found on the pulley and the clevis. The instrument passed electrical continuity testing. Engineering also found the pitch cable was loose at the distal clevis hub. Both cable crimps remained in the clevis and no damage was found on the cable. The housing was removed and the pitch cable was found loose at the clamping pulley, but no loose clamping pulley screw was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3117561 | 3089371 | 2985915 | 2955842201301499 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-05 | 2013-05-02 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-28 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the jaws of the prograsp forceps instrument could not open or close; identified during set up, instrument was not used in the case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on the in-house da vinci robot system and it was driven with no problem. Instrument passed recognition and engagement test. Instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Instrument was found to be fully functional. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3117590 | 3089379 | 2985923 | 2955842201301479 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-05 | 2013-05-02 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-01 | Malfunction |  | It was reported that during a da vinci si radical hysterectomy procedure the prograsp forceps instrument stopped functioning (went loose in the jaws). The staff removed and reinserted the instrument and the issue recurred. The staff inspected the instrument, rotated the discs on the housing, and discovered the jaws/tip acted strange, it was very loose. The staff suspected possibly a loose wire. The instrument was replaced with a backup instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch down cable was broken at the distal clevis hub. The cable segment that contained the crimp was missing. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3118011 | 3089769 | 2986313 | 2955842201301502 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-03 | 2013-05-02 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-04-03 | Malfunction |  | During central processing, the user facility identified a frayed cable on the large needle driver instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. Engineering observed a grip close cable that was broken at the distal idlers pulley. The grip close cable was also sticking out at the instrument's wrist. The idler pulley spun freely but had damages on the edge and on the surface. Engineering concluded that damage was likely due to excess force contact and mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3118065 | 3089815 | 2986339 | 2955842201301503 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-05-02 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-04-04 | Malfunction |  | During central processing, the user facility reportedly identified a broken cable on the large suturecut needle driver instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed that one grip close cable was frayed at the grip hub. The frayed strands were sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were 0. 30-0. 34 in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3118955 | 3090529 | 2987078 | 2955842201301504 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-04 | 2013-05-02 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-04 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-05 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff alleged that the monopolar cautery cord connected to the monopolar curved scissors instrument sparked and then split. The planned surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The monopolar cautery cord was not returned to isi for evaluation. Per the customer reported complaint, engineering evaluation was unable to confirm the alleged issue of the monopolar cautery cord sparking and splitting. The instrument involved with this complaint was returned and evaluated. The instrument was placed on an in-house system and performed cautery functions with no trouble found. Engineering evaluation found a hairline crack at the distal end of the instrument's main tube. On (b)(4) 2013, isi contacted a surgical technician at the site to obtain additional information regarding the reported event. The surgical technician indicated that the surgical staff observed sparks coming from the cautery cord. The surgical staff resolved the issue by replacing the cautery cord and the surgeon was able to proceed with the surgical procedure using the same instrument. The surgical technician indicated that the affected cautery cord was from an unknown vendor. The surgical technician denied that there was any harm, adverse outcome, or injury to the patient or surgical staff. On (b)(4) 2013, an isi field service engineer (fse) indicated that he inspected the da vinci surgical system and was unable to replicate the issue reported by the site. The fse performed an electrical safety test with the system that passed and the values were found to be within tolerance. At the time of the fse's investigation, the affected cautery cord and instruments used during the surgical procedure were not available for investigation. On (b)(4) 2013, isi contacted the initial reporter of this complaint. The initial reporter indicated that no fragments fell into a patient. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3119034 | 3090600 | 2987149 | 2955842201301506 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2013-04-05 | 2013-05-02 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-23 | Malfunction |  | During central processing, the user facility identified a frayed wire at the tip of the fenestrated bipolar forceps instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint: there was a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. Electrical continuity testing was performed and passed. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3119192 | 3090616 | 2987170 | 2955842201301511 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-23 | 2013-05-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-23 | Device Not Retu | No | Reuse | 2013
42805206 | 2013-04-01 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si cholecystectomy procedure, the monopolar cautery hook tip accessory used in conjunction with the monopolar cautery instrument broke off and fell into the patient. The instrument was being used to blunt dissect tissue. The broken tip was retrieved. Upon inspection of the instrument, the threads appeared to be stripped and the rectangular opening that the metal end of the hook inserts into appeared to be distorted as well. The surgeon made the decision to convert the procedure to traditional laparoscopic surgery. No allegation of any harm, injury, or adverse outcome to a patient was reported. | Manufacturer Narrative: --/-- The checkbox on the initial report for returned to manufacturer on should have been unchecked. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Intuitive surgical inc. (isi) has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. | MONOPOLAR CAUTERYINSTRUMENT | Unkown | 3119042 | 3090628 | 2987157 | 2955842201301507 | 1 | NAY | Manufacturer | User facility | Other | FollowupInitial | 2013-04-04 | 2013-05-02 | No | Yes | Health Professional | 42805206 | No | No | 2013-04-04 | Device Not Retu | No |  | 2013
42017207 | 2013-04-02 | Malfunction |  | It was reported that during reprocessing there was an issue with the maryland bipolar forceps instrument. No specific complaint was alleged. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The complaint of a reprocessing issue could not be confirmed as no specific complaint was alleged. A visual inspection of the grips and wrist assembly with no magnification showed no obvious contamination. An additional observation not reported by the site was a frayed cable. The pitch up cable was frayed at the distal clevis hub. The frayed strands stuck out at the wrist. Other cables at the instruments wrist were not damaged. No other damage was found. On (b)(4) 2013, isi attempted to contact the initial reported of this complaint to obtain additional information regarding the reported event. At the time of this report, no additional information has been provided. If additional information becomes available a follow up will be provided if warranted. This the customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3119088 | 3090654 | 2987203 | 2955842201301505 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-02 | 2013-05-02 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-27 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure that the system would not recognize the large needle driver instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was checked for recognition on an in-house is3000 system. The system successfully recognized instrument, showing 3 uses left. The pogo pins did not stick and were not contaminated. Dcp verification of the instrument passed. Engineering was unable to replicate the recognition issue. Additional observations not reported was a frayed cable and damaged distal clevis. The pitch down cable was frayed at the distal clevis hub. The frayed strands stuck out at the instruments wrist. Other cables at the wrist were not damaged. The distal clevis hub exhibited scratch and gouge marks adjacent to the fraying. Engineering concluded the clevis damage was likely due to mishandling/misuse and likely contributed to cable fraying. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3119116 | 3090680 | 2987229 | 2955842201301508 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-03 | 2013-05-02 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-14 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted a broken cable on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3119152 | 3090712 | 2987265 | 2955842201301509 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-02 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-15 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy the jaws of the fenestrated bipolar forceps instrument did not close properly. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The complaint of the jaws not closing properly was confirmed. One grip was bent, causing side to side misalignment of grips. There was a. 032 offset at the instrument tips. Electrical continuity passed. Engineering concluded the damage was likely due to mishandling. Additional damage found was a couple of deep scratches on the distal end of the main tube exhibiting light material removal and a rough surface finish. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3119162 | 3090723 | 2987276 | 2955842201301510 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-05-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-19 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tips of the fenestrated bipolar forceps instrument were not aligned. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip was bent, causing side to side misalignment of the grips. There was a. 020 offset at the tips. The bent grip had separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. Engineering concluded the grip damage may have been caused by mishandling/misuse. An additional finding was various scratches on the main tube showing light material removal. The scratches were scratches are. 185 -. 290 in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3119218 | 3090817 | 2987371 | 2955842201301513 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-05-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
40018012 | 2012-10-19 | Injury | Required Intervention | On (b)(6) 2013, intuitive surgical inc. (isi) received uf importer report (b)(4) from the fda with the following event description: upon closing the patient, it was reported that a cap from the monopolar scissors was missing. Patient required a return to the operating room (or) for removal after tip was located by a computed tomography (ct) scan. What was the original intended procedure? total laparoscopic hysterectomy with davinci assist device usage problem: device malfunction - that is, the device did not do what it was supposed to do. | The tip cover accessory has not been returned to isi for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, isi contacted the initial reporter regarding the reported event. The reporter indicated that prior to use the tip cover was inspected. However, she indicated that she did not know if the installation tool was used to install the tip cover onto the monopolar curved scissors (mcs) instrument or if electrolube was applied to the mcs instrument prior to installation of the tip cover onto the mcs instrument. The initial reporter was unable to confirm if the mcs instrument made contact with any other instrument during the surgical procedure or if the mcs instrument was removed at any time during the surgical procedure. Per the initial reporter, there was no report by the surgical staff that a malfunction of the instrument occurred during the surgical procedure. The initial reporter indicated that a ct scan performed on the patient revealed that the patient had retained the tip cover accessory and that the patient was returned to the operating room the same day for removal of the tip cover accessory. The tip cover was removed laparoscopically. The initial reporter indicated that the patient recovered as planned and was stable upon discharge from the hospital. The reporter indicated that the tip cover accessory and mcs instrument were not available for return to isi for evaluation. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3119290 | 3090826 | 2987380 | 2955842201301514 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-04-02 | 2013-05-02 | Yes | Yes | Health Professional | 40018012 | No | No | 2013-04-02 | No | Yes | Initial | 2013
42022703 | 2013-03-29 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the pk dissecting forceps instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at distal end. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. Engineering also found a deep scratch on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. Instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3119225 | 3090833 | 2975577 | 2955842201301517 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-02 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-04-05 | Malfunction |  | It was reported that during a da vinci si colectomy procedure the small grasping retractor instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the grip close cable was broken at the hypotube inside the distal end of the main tube. No other cable damage was found. Engineering also found the clamping pulley with corrosion. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3119299 | 3090834 | 2985183 | 2955842201301518 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-02 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that both pitch cables were broken at the distal clevis hub. The cable segments that contained the crimp were still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. Engineering also found various scratch marks at the distal end of the main tube showing light material removal. The scratches were. 250 -. 275 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3119300 | 3090835 | 2987389 | 2955842201301519 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the prograsp forceps instrument was noted to have wires exposed at the distal end. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument had a frayed cable at distal idler pulley. The frayed cable section was approximately 0. 05 in length. No damage was found on the pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3119304 | 3090836 | 2987390 | 2955842201301520 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-02 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-04 | Malfunction |  | It was reported that after a da vinci si surgical procedure, frayed wires were identified in central processing on the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The complaint of frayed wires was confirmed. One grip close cable was broken near the proximal pulleys and proximal clevis cable hole. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the instruments wrist. Additional observation not reported by site was proximal clevis wearing. The proximal clevis cable hole exhibited wear on one edge. Other cables at the instruments wrist were not damaged. Evidence is not conclusive but wear on the hole may have contributed to the cable breakage. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3119306 | 3090838 | 2987392 | 2955842201301522 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-04 | 2013-05-02 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-04 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-04-09 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the thoracic grasper instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint; no broken cables were found. Engineering found insulation peeling off shaft. No material was missing because insulation flap covers exposed area. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3119226 | 3090850 | 2987404 | 2955842201301516 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-02 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2013-04-02 | Other | Other | It was reported that during a da vinci si sacrocolpopexy procedure, the surgeon experienced a dark and distorted right eye image through the high resolution stereo viewer (hrsv) on the surgeon side cart (ssc). The site contacted isi for technical support engineering (tse) assistance. Review of the site's system log by the tse found that the site experienced system error code 45310 and that the vision issue experienced by the site was likely due to a faulty camera cable. The tse instructed the site to install a backup camera cable to troubleshoot the issue; however, the site indicated that they did not have a backup camera cable available. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm, adverse outcome, or injury was reported. | The investigation conducted by the field service engineer (fse) concluded that the vision issue experienced by the site was associated with a faulty camera cable. The fse spoke to the site's robotics coordinator and the isi clinical sales representative (csr). They indicated to the fse that after completion of the open procedure, a backup camera cable was located and after replacement of the faulty camera cable the vision issue was resolved. On (b)(4) 2013, the isi clinical sales representative, who was present during the surgical procedure, was contacted to obtain additional information regarding the reported event. The csr indicated that the vision issue experienced by the site occurred after the patient side cart (psc) was docked to the patient, the instruments were installed and after the surgeon sat down at the surgeon side cart to begin the surgery. The surgeon had not yet performed any surgical tasks. According to the csr, the surgical staff was not aware that their backup camera cable was located in their materials department when the vision issue occurred and the patient was under anesthesia for approximately 45 minutes when the surgeon decided to convert the procedure to open surgical techniques. The csr indicated that there were no complications experienced by the patient and that the patient tolerated the open procedure well. Per the csr the patient did not experience any post-surgical complications and the patient is doing ok. As of the date of this report, there have been no reported recurrences of the issue at the hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3119227 | 3090851 | 2987405 | 2955842201301515 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-04-02 | 2013-05-02 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was missing on the opposite side of the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering also found scratch marks on the distal end of the main tube showing light material removal. The scratch was short in length and not axially aligned with the tube. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3119302 | 3090853 | 2987389 | 2955842201301521 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-13 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy the fenestrated bipolar forceps instrument does not close. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip was bent causing side to side misalignment of the grips. There was a. 080 offset at the instrument tips, indicating overloading. Electrical continuity test passed. Additional findings were the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. There were also scratches on the surface of the distal pulley. Lastly, the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3119318 | 3090867 | 2938117 | 2955842201301523 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-05 | 2013-05-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-22 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the fenestrated bipolar forceps instrument was not cauterizing. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the conductor wire was broken from the bipolar pin inside the housing, which has no contact with the patient. The instrument failed electrical continuity testing. Engineering also found deep scratches on the main tube and a bent bipolar pin. The distal end of the main tube had a couple of scratch marks exhibiting light material removal and a rough surface finish. The bottom pin was slightly bent at the midpoint. Engineering concluded that the damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3115352 | 3087271 | 2971873 | 2955842201301462 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-05-01 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-03 | Malfunction |  | During a da vinci s procedure, the user facility allegedly identified broken wires at the tip of the pk dissecting forceps instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3115654 | 3087538 | 2984082 | 2955842201301466 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-05-01 | No | Yes | Health Professional | 42022703 | No | No | 2013-04-03 | No | No | Reuse | 2013
42019408 | 2013-03-25 | Malfunction |  | During a da vinci si tongue base resection procedure, the user facility allegedly identified broken tips on the mega needle driver instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. A broken tip was initially reported by the customer; however, for clarification the damaged instrument component was a broken grip cable by the tip area. Based on this clarification, the customer's reported complaint was confirmed. One grip close cable was broken at the distal idlers and was sticking out at the instrument's wrist. The idler pulley spun freely and was not damaged. Other cables at the instrument's wrist were not damaged. Additional observations not initially reported by the customer were noted. There was an indentation in the middle of the distal pulley with scratches on the surface. Engineering was unable to conclude the cause of the damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3115763 | 3087667 | 2984211 | 2955842201301467 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-05-01 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
40017208 | 2013-03-15 | Malfunction |  | It was reported that during a da vinci surgical procedure, in the body of the maryland bipolar forceps instrument, the pins loosen at the junction of the bipolar cable. The instrument could not be used again and had to be exchanged. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The complaint of the pins loosening at the junction of the bipolar cable was confirmed. The bipolar pins were found loose at the back of the housing because the bipolar insert that supports both pins was dislodged from the plug riser. Electrical continuity test passed. Additional damage found were deep scratches on the main tube. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage was likely due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3115848 | 3087734 | 2975398 | 2955842201301468 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-03 | 2013-05-01 | No | Yes | Health Professional | 40017208 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-03 | Malfunction |  | During a da vinci s surgical procedure, the user facility allegedly identified a frayed cable at the tip of the fenestrated bipolar forceps instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3115887 | 3087741 | 2984305 | 2955842201301469 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-05-01 | No | Yes | Health Professional | 42020505 | No | No | 2013-04-03 | No | No | Reuse | 2013
42017207 | 2013-04-04 | Malfunction |  | During a da vinci si surgical procedure, a cable from the maryland bipolar forceps instrument allegedly broke. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. The pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3116043 | 3087945 | 2984499 | 2955842201301470 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-04 | 2013-05-01 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-04 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-28 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted that a metal fiber was missing at the top of the maryland bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with frayed cable located at the yaw pulley. The frayed cable section was approximately 0. 07 in length. No damage was found on the distal pulley. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3116229 | 3088126 | 2984650 | 2955842201301471 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-02 | 2013-05-01 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-03-29 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified tips on the bowel grasper instrument that did not meet properly. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering did not confirm the alleged complaint. The grip tips were aligned when closed. Additional observation not initially reported by the site was main tube damage. Engineering evaluation observed a deep scratch on the main tube, which was exposing the metal shaft of the main tube. A piece of the insulation coating was chipped off and was missing. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116511 | 3088340 | 2984874 | 2955842201301472 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-05-01 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-28 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the da vinci robot system could not recognize the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was inspected for recognition using an in-house da vinci robot system. The system successfully recognized instrument, showing 4 uses left. The pogo pins did not stick and were not contaminated. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116562 | 3088357 | 2984911 | 2955842201301475 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-05 | 2013-05-01 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-04-02 | Malfunction |  | It was reported that during reprocessing of a bowl grasper instrument the casing was peeling off midway down the shaft. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. Black tube insulation was damaged between the distal and proximal ends. The insulation was gouged in various places and had pieces of insulation lifted up. The metal shaft was exposed as a result. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116529 | 3088358 | 2984912 | 2955842201301476 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-05-01 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-04-02 | Malfunction |  | It was reported that during reprocessing of a bowl grasper instrument the casing was peeling off midway down the shaft. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. Black tube insulation was damaged between the distal and proximal ends. The insulation was gouged in various places and had pieces of insulation lifted up. The metal shaft was exposed as a result. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116554 | 3088388 | 2984912 | 2955842201301473 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-05-01 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-04 | Malfunction |  | During a da vinci s hysterectomy procedure, a spring came out of the mega suturecut needle driver instrument's tip. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. The instrument was found with one grip close cable derailed at the distal idler pulley. Damages were found on the surface of the distal idler pulley. Additional observation not reported by the customer was a frayed grip cable at the distal clevis. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116527 | 3088390 | 2984934 | 2955842201301474 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-04 | 2013-05-01 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-04 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-04 | Malfunction |  | During a da vinci si hysterectomy procedure, a surgical assistant heard a loud noise pop. The surgical assistant reportedly found a visible wire on the fenestrated bipolar forceps and described the instrument's wrist to be limp. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit excessive damage. No other damage was observed. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3116577 | 3088400 | 2984944 | 2955842201301477 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-04 | 2013-05-01 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-04 | Device Not Retu | No | Reuse | 2013
42004906 | 2013-04-02 | Malfunction |  | It was reported that during reprocessing that the tip of the cadiere forceps instrument was broken. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The complaint that the tip of the instrument was broken was not confirmed. However, the finding was a frayed pitch cable. The instrument was found with a frayed pitch cable at distal idler. No damage found on the pulley or clevis. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116569 | 3088409 | 2984963 | 2955842201301478 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-05-01 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-11 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the scrub nurse noticed that one of the wrist cables of the maryland bipolar forceps instrument was damaged. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with frayed pitch cable located at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3116608 | 3088422 | 2984976 | 2955842201301481 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-04-05 | 2013-05-01 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-02 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure a broken cable was identified on the pk dissecting forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported broken wire, however, for clarification the damaged instrument component was a frayed pitch cable. Based on this clarification, the customer reported complaint is confirmed. The instrument was found with a frayed pitch cable at the distal idler. No damage found on the pulley or clevis. An electrical continuity test passed. Additionally, there were scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3116586 | 3088432 | 2984956 | 2955842201301480 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-05-01 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-02 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure that frayed cables were identified on the mega suturecut needle driver instrument. The instrument was not used. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. Additional observations not reported by the site were a scratched pulley and derailed cable. The distal idler pulley on which the broken cable was seated had various scratch marks on the outer surface of the outermost distal idler pulley. The grip open cable on same side as broken cable was derailed from its pulley and seated on outermost distal idler pulley. The derailment of the grip open cable combined with pulley surface scratches suggests the grip close cable was derailed prior to breakage. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116606 | 3088438 | 2984982 | 2955842201301482 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-05-01 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-03 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the customer noted a broken cable on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116639 | 3088466 | 2984990 | 2955842201301483 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-05 | 2013-05-01 | No | Yes | Health Professional | 42030901 | No | No | 2013-04-05 | No | No | Unkown | 2013
42020706 | 2013-04-22 | Malfunction |  | During a da vinci si hysterectomy procedure, a cable on the tenaculum forceps instrument allegedly broke. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. The pitch cable was broken at the distal end. The cable segment that contains the crimp was missing from clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. Additional observation not initially reported by the customer was a discolored chassis. The instrument's chassis exhibited a greyish layer, which was likely due to improper cleaning. No other damage found. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116736 | 3088574 | 2985118 | 2955842201301485 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-22 | 2013-05-01 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted a broken cable on the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was bent grip tips. One grip was found to be bent, causing side to side misalignment of the grips. There was a 0. 0535 offset at the tips. Instrument passed electrical continuity test. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3116737 | 3088575 | 2985119 | 2955842201301484 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-05 | 2013-05-01 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-19 | Malfunction |  | During a da vinci si hysterectomy procedure, the orange sleeve on the monopolar curved scissors allegedly became wrinkled and the user facility was unable to remove the instrument from the trocar. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. Engineering observed a broken piece on the main tube extension at the proximal clevis interface. The broken piece was dislodged from the clevis. Engineering concluded that the main tube extension likely broke due to excessive side loading or other mishandling. Additional observation not initially reported by the customer was main tube damage. The distal end of the main tube exhibited scratch marks with light material removal and had a rough surface finish. Engineering concluded that the main tube damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3116627 | 3088583 | 2985127 | 2955842201301486 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-22 | 2013-05-01 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-24 | Malfunction |  | During a da vinci si prostatectomy procedure, a cable on the prograsp forceps allegedly broke. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. Engineering observed a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116774 | 3088608 | 2985152 | 2955842201301487 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-25 | 2013-05-01 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-13 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the maryland bipolar forceps instrument was noted to slide. The instrument was inspected prior to the robotic surgery. The surgeon was suturing the stump, the surgeon mentioned to the scrub nurse, that there was something wrong with the endowrist; that it slides during suturing of the stump. The scrub nurse noticed that the strings on the instrument tip were snapped making the endowrist defective and it was noted after the procedure when the instrument was being cleaned. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was frayed at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3116787 | 3088622 | 2985166 | 2955842201301489 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-08 | 2013-05-01 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-29 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the mega suturecut needle driver instrument scissor blades were noted to be dull. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. The instrument scissors were not dull. Engineering also found the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had slight damages and on the surface. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116788 | 3088623 | 2985167 | 2955842201301488 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-01 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-03-27 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the tenaculum forceps instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering installed the instrument on an in-house system and the system successfully recognized instrument. The pogo pins were not stuck or contaminated. The dcp (dallas chip programmer) verification of the instrument passed. The instrument was driven and moved intuitively and the grips opened and closed. Engineering found various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the main tube. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116789 | 3088624 | 2985168 | 2955842201301491 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-05 | 2013-05-01 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-08 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical mega suturecut needle driver instrument was noted to have broken cables. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had slight damages on the surface. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116810 | 3088625 | 2985169 | 2955842201301490 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-08 | 2013-05-01 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-04-05 | Malfunction |  | It was reported that during a da vinci si surgical procedure the mega needle driver instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that both grip cables (open/close) attached to one crimp were broken at the distal end. The grip close cable was broken at the distal idler pulley, and the grip open cable was broken near the crimp. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at wrist. The other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116820 | 3088628 | 2985182 | 2955842201301500 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-05 | 2013-05-01 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-04-05 | Malfunction |  | It was reported that during a da vinci si colectomy procedure the small grasping retractor instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering also found scratches on the surface of the main tube. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116821 | 3088629 | 2985183 | 2955842201301501 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-09 | 2013-05-01 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-04-09 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-27 | Malfunction |  | It was reported that during a da vinci s surgical procedure the prograsp forceps instrument was not recognized by the system. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that when the housing was removed the one grip cable was derailed at the back end pulley. The cable was wedged and created tight tension. The other backend cables were undamaged. Engineering also found the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116791 | 3088647 | 2985171 | 2955842201301492 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-05 | 2013-05-01 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-04-02 | Malfunction |  | It was reported that during a da vinci s surgical procedure the mega needle driver instrument wire broke while handling a needle. The instrument was removed and a different instrument was used in its place. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116812 | 3088648 | 2985172 | 2955842201301493 | 1 | NAY | Manufacturer | DistributorHealth Professional | Physician | Initial | 2013-04-05 | 2013-05-01 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-27 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the pk dissecting forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering also found the one grip was bent, causing side-to-side misalignment of the grips. There was a. 067 offset at the tips, indicating overloading at the tip. Engineering concluded that the damage was likely due to mishandling. Engineering also found an indentation on the edge of the distal pulley and some scratches. Engineering concluded the damage was most likely due to mishandling. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3116814 | 3088649 | 2974692 | 2955842201301494 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-05 | 2013-05-01 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-05 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-01 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure the mega suturecut needle driver instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. There was distal clevis wear and pulley scratches. The distal clevis post on the same side as the cable breakage had wear across the top of the post. The outermost distal idler pulley exhibited scratch marks. The evidence was inconclusive, but the clevis and pulley damage suggested the cable derailed and rubbed against these surfaces prior to breakage. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | Reuse | 3116816 | 3088651 | 2985175 | 2955842201301495 | 1 | NAY | Manufacturer | OtherUser facility | Other | InitialFollowup | 2013-05-13 | 2013-05-01 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | No |  | 2013
40027202 | 2013-03-25 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure the harmonic ace curved shears insert was noted to separate after several minutes of use. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the insert with the teflon pad was damaged and with a portion of the pad detached from the clamp arm. Charred marks and slight deformation were found on the pad. No other physical damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | HARMONIC ACE CURVED SHEARS INSERT | Initial | 3116817 | 3088652 | 2985176 | 2955842201301496 | 1 | NAY | Manufacturer | ForeignUser facility | Other | FollowupInitial | 2013-05-13 | 2013-05-01 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2013-05-13 | Device Not Retu | Yes |  | 2013
42020505 | 2013-03-22 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical fenestrated bipolar forceps instrument tips did not meet. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side to side misalignment of the grips. There was a. 035 offset at the tips. The bent grip had separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. Engineering also found tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 110 in length and not aligned with the tube axis. Engineering concluded that the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3116818 | 3088653 | 2985177 | 2955842201301498 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-04 | 2013-05-01 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-04 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-01 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical prograsp forceps instrument came apart. No further details were provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found that the pitch cable was frayed at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 04 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3116819 | 3088654 | 2975549 | 2955842201301497 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-04 | 2013-05-01 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-04 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-03 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the monopolar curved scissors (mcs) instrument orange part was noted to be peeling off the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was found broken and missing a piece at the proximal clevis interface. The clevis was not dislodged from the tube extension. Engineering concluded the tube extension likely broke due to excessive side loading or other mishandling. Engineering also found a derailed grip cable at the clamping pulley and deep scratches on the main tube. The distal end of main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3113136 | 3085061 | 2980244 | 2955842201301452 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-04-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-19 | Malfunction |  | During a da vinci si myomectomy procedure, a cable on the mega suturecut needle driver instrument allegedly broke. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint. One grip close cable was broken at the distal idlers and was also sticking out at the instrument's wrist. The idler pulley spun freely and was not damaged. Other cables at the instrument's wrist were not damaged. Additional observation not initially reported by the site were scratches on the surface of the distal pulley. It is unknown how the scratches occurred. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3113476 | 3085428 | 2981972 | 2955842201301454 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-22 | 2013-04-30 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-22 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-22 | Malfunction |  | During a da vinci si myomectomy procedure, the pk dissecting forceps instrument allegedly was not articulating fully. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint. The pitch cable was found loose at the distal clevis hub. Both of the cable crimps remained in the clevis. When the instrument's housing was removed, the pitch cable was found to be loose at the clamping pulley but no loose clamping pulley screw was found. Electrical continuity testing was performed and passed. Additional observations not initially reported by the site were a frayed cable, scratches on the pulley, and scratches on the main tube. The loose pitch cable was found to be frayed at the distal clevis hub. There were also scratches on the surface of the distal pulley. The distal end of the main tube exhibited various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the main tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3113527 | 3085464 | 2981998 | 2955842201301455 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-26 | 2013-04-30 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-20 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the wrist of the fenestrated bipolar forceps instrument was not rotating. The instrument was not used. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The complaint of the wrist of the instrument not rotating could not be confirmed. The instrument was placed on an is3000 system and driven. Recognition and engagement test passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Additional observations not reported by site were tube damage and a bent bipolar pin. The distal end of main tube has a couple scratch marks with light material removal. The scratches were. 150 in length and not aligned with the tube axis. The bottom bipolar pin was bent upwards. A bipolar cord could not be fully installed as a result. Engineering concluded the damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3113897 | 3085898 | 2982442 | 2955842201301456 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-04-30 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-03-20 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the wires came off the track of the mega needle driver instrument. The instrument was not used. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3113908 | 3085929 | 2982463 | 2955842201301457 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-04-30 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-04-02 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical pk dissecting forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument had a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 04 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3113962 | 3086035 | 2982559 | 2955842201301458 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-04-30 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-26 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a loose wire on the large needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint. One grip close cable was loose at the distal idlers. The idler pulley spun freely and was not damaged. The clamping pulley screws were found to be secured. Additional observations not initially reported by the site were also noted. A grip wire was loose due to a broken grip cable inside the housing by the clamping pulley. The idler pulley spun freely and was not damaged. There were indentations at the edge of the distal pulley: the cause of damage cannot be determined. The distal end of main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the main tube. Engineering concluded that the main tube damage may be due to mishandling. There was evidence of corrosion around the clamping pulley inside the instrument's housing. There was also an orange rusty color like substance below the clamping pulley around the bearing. Engineering concluded that the corrosion may be due to improper cleaning. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3114190 | 3086220 | 2982754 | 2955842201301459 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-04-30 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-02 | Malfunction |  | During a da vinci si procedure, the user facility allegedly identified a exposed wire at the distal end of the fenestrated bipolar forceps instrument. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint. The grip cable was frayed at the distal idler pulley and was sticking out at the instrument's wrist. Other cables at the instrument's wrist were not damaged. Electrical continuity testing was performed and passed. Additional observations not initially reported by the site were noted. One instrument grip was bent, causing side to side misalignment of the grips. There was a 0. 075 offset at the tips, indicating overloading at the tip. There were also scratches on the surface of the distal pulley. The cause of the pulley scratches is unknown. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3114457 | 3086495 | 2983029 | 2955842201301460 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-04-30 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-03-14 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the permanent cautery spatula instrument tip was noted to be bent. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the spatula is broken. The instrument was returned with a piece of the spatula broken off. The broken piece was not returned, but measures roughly. 150 x. 090. The tip fractured at the line where the spatula is formed into the flattened spatula shape. The side view of the fractured surface showed a bend to one side, indicating an applied load perpendicular to the spatula surface. Engineering concluded that the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3114484 | 3086521 | 2983065 | 2955842201301465 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-04-30 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-19 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure the monopolar curved scissors (mcs) instrument shaft was noted to be cracked at the distal end. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was cracked at the distal end, with no material removed. Engineering concluded the damage was likely due to mishandling. Engineering also found a hairline crack at the distal end of the main tube. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3114483 | 3086522 | 2983066 | 2955842201301464 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-04-30 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
40009308 | 2013-03-13 | Malfunction |  | It was reported that during a da vinci s surgical procedure the prograsp forceps instrument was not responding to the system. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering placed the instrument on an in-house is1200 system and it was driven. The instrument was recognized and passed engagement testing. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Functional performance testing did not find any issues. Engineering found the main tube has a couple of wear marks on the distal end showing light material removal and a rough surface finish. The scratches were. 500 - 1. 25 in length and were axially aligned with the tube. Engineering concluded the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3114493 | 3086527 | 2983081 | 2955842201301461 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | PHARMACIST | Initial | 2013-04-03 | 2013-04-30 | No | Yes | Health Professional | 40009308 | Device Returned To Manufacturer | Yes | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-22 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure the maryland bipolar forceps instrument jaws were noted to be misaligned. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of the grips. There was a. 025 offset at the tips. The bent grip had separation of the yaw pulley and the pulley cover at the glue joint, indicating overloading at the tip. Engineering also found tube damage. The distal end of the main tube had a scratch mark with light material removal. The scratch was. 125 in length and not aligned with the tube axis. Engineering concluded the tube damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3114482 | 3086528 | 2983082 | 2955842201301463 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-04-30 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42020906 | 2013-03-19 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the instrument was falling apart. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer complained that the instrument was falling apart but they did not specify which part was broken. Engineering inspected the instrument and some damage was found on the proximal end. The release lever was missing and the housing snaps were broken. The two rear housing pins and snaps were also broken off, indicating the instrument was mishandled. Engineering concluded the housing likely popped off during mishandling, causing the lever to fall out. The housing was then taped back on the chassis without the lever. Additionally, two indentations along the edges of the distal idler pulley were found. Various deep scratch marks exhibiting light material removal and a rough surface finish on the distal and proximal end of main tube were found. Engineering concluded the damage was also likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3109512 | 3081471 | 2978005 | 2955842201301423 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-29 | No | Yes | Health Professional | 42020906 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the robot would not advance the monopolar curved scissors instrument properly. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering did not confirm the alleged complaint. The instrument was checked for recognition and movement on an in-house system. Recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. The instrument's blades opened and closed with no problem and it passed the cut test. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited a deep scratch mark with light material removal and had a rough surface finish. Engineering concluded that the damage may be due to mishandling. Electrical continuity testing was performed on the instrument and passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3109594 | 3081522 | 2978056 | 2955842201301424 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-29 | 2013-04-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-25 | Malfunction |  | During central processing, the user facility identified a frayed cable on the megasuture cut needle driver instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. The pitch down cable was frayed at the distal clevis hub. The frayed strands were sticking out at the wrist. Other cables at the instrument's wrist were not damaged. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3109616 | 3081562 | 2978096 | 2955842201301426 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-29 | 2013-04-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-29 | Malfunction |  | It was observed during central processing that the prograsp forceps instrument had a broken wire. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found one grip close cable frayed at the force amp pulley. Frayed strands were observed to be sticking out at the instrument's wrist. The other cables of the instrument's wrist did not exhibit any damage. Engineering evaluation also found a damaged pulley. The force amp pulley was found to have a small gouge on one edge. The damaged edge was adjacent to where cable fray initiated. Engineering evaluation concluded that the pulley damage was likely due to mishandling/misuse and likely contributed to cable fraying. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, isi contacted the initial reporter of this complaint to obtain additional information regarding the reported event. The initial reporter indicated that nothing fell into a patient and there was no patient harm. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3109609 | 3081570 | 2978104 | 2955842201301425 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-29 | 2013-04-29 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-26 | Malfunction |  | Event Description: --/-- During central processing, the user facility identified a chip in the blade of the monopolar curved scissors instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering confirmed the chip in the instrument's blade. Pitting/chipping was found on one of the instrument's blade at the distal end. Around the blade tip, there was metal material missing causing the blade to not close properly. Engineering concluded that the blade damage may be due to improper cleaning. Additional observation not initially reported by the site were corrosion on the back idler pulleys and deep scratches on the main tube extension. The back idler pulleys showed corrosion around the edges of each pulley. The scratches on the main tube extension exhibited material removal and a rough surface. Engineering concluded that the damages may be due to mishandling. Electrical continuity testing was performed on the instrument and passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3109679 | 3081581 | 2978115 | 2955842201301431 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-04-01 | 2013-04-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42017207 | 2013-03-12 | Malfunction |  | It was reported that during a da vinci si surgical procedure the maryland bipolar forceps instrument would not cauterize. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was placed on an in-house system and successfully passed a cautery test. The conductor wires were not damaged and an electrical continuity test passed. Additional damage found were several deep scratches exhibiting light material removal and a rough surface finish on the distal end of the main tube. Engineering concluded the damage was likely due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3109632 | 3081599 | 2978133 | 2955842201301427 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-29 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-25 | Malfunction |  | During central processing, the user facility identified a fraying cable on the fenestrated bipolar forceps instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. The conductor wires were intact and undamaged but the drive cable was frayed. The pitch down cable was frayed at the distal clevis hub. The frayed strands were sticking out at the wrist. Other cables at wrist are not damaged. Additional observation not initially reported by the site was main tube damage. Engineering evaluation found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were 0. 055-0. 150 in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3109634 | 3081603 | 2978137 | 2955842201301428 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-29 | 2013-04-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-08 | Malfunction |  | During a da vinci si surgical procedure, a cable on the fenestrated bipolar forceps instrument reportedly snapped. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. The pitch down cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3109636 | 3081605 | 2978139 | 2955842201301429 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-29 | 2013-04-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that prograsp forceps instrument had a wire that was separated and hanging out. The planned surgical procedure was completed. There were no missing or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found a broken pitch cable at the distal clevis hub. A missing crimp was observed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering evaluation also found various scratch marks showing light material removal and a rough surface finish on the distal end of instrument's main tube. The scratches were observed to be short in length and not axially aligned with the tube. Engineering evaluation concluded that the main tube damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3109644 | 3081612 | 2978146 | 2955842201301430 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-01 | 2013-04-29 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-25 | Malfunction |  | It was reported that during a da vinci si pancreatectomy procedure, the surgical staff alleged that prograsp forceps instrument broke and was not working. The planned surgical procedure was completed. There were no missing or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation was unable to confirm the alleged issue that the instrument was not working. The instrument was placed on an is3000 system and driven. The instrument passed recognition and engagement tests. The instrument moved intuitively with full range of motion in all directions. The instrument's grips opened and closed properly. Functional performance testing of the instrument did not find any issues. Engineering evaluation found various scratch marks with light material removal on the distal end of the instrument's main tube. The scratches were. 080 -. 100 in length and not aligned with the tube axis. Engineering evaluation concluded that the main tube damage was likely caused by mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3109672 | 3081623 | 2978157 | 2955842201301433 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-30 | 2013-04-29 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-30 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-10 | Malfunction |  | During a da vinci si prostatectomy procedure, the tips of the prograsp forceps instrument allegedly were not closing properly. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3111143 | 3082878 | 2979412 | 2955842201301438 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-25 | 2013-04-29 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-03 | Malfunction |  | Prior to starting a da vinci s surgical procedure, the user facility reportedly identified a twisted jaw on the fenestrated bipolar forceps instrument. It was also reported that the bipolar energy was not working. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint. One of the instrument's grip was bent, causing side to side misalignment of the grips. There was a 0. 103 offset at the tips, indicating overloading at the tip. Electrical continuity testing was performed on the instrument and passed. Engineering concluded that the tip damage may be due to mishandling. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3111234 | 3082977 | 2979514 | 2955842201301439 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2013-04-24 | 2013-04-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-24 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-01-13 | Malfunction |  | During a da vinci si prostatectomy procedure, the instrument axis of the monopolar curved scissors instrument allegedly stopped responding. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint. The roll input disc did not move smoothly when it was manually rotated. Visual inspection showed the front roll axis bearing exhibiting contamination that was possibly causing increased friction. Engineering evaluation concluded that improper cleaning may have contributed to the alleged issue. Additional observation not initially reported by the site were damages to the main tube and banana plug. Engineering evaluation noted that the tube extension exhibited pad print removal on several different locations. The distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. In addition, the banana plug was bent at the back of the chassis. Engineering concluded that the damages may be due to mishandling. Electrical continuity testing on the instrument passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3111554 | 3083193 | 2979727 | 2955842201301440 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-25 | 2013-04-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42017207 | 2011-06-27 | Malfunction |  | During a da vinci si hysterectomy procedure, the bipolar energy of the monopolar curved scissors instrument reportedly was not functioning. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation did not confirm the alleged complaint. Visual inspection of the instrument showed no damage to the conductor wires. The instrument was checked for electrical continuity, which is an indicator of cautery function. The electrical continuity test passed. Functional cautery test using a bipolar cord and generator was performed and the instrument successfully transferred energy to a wet paper towel, causing steam to be generated. Additional observation not initially reported by the site were damages to the bipolar pins and main tube. Both bipolar pins were bent sideways. The bottom pin was bent more than the top pin. A bipolar cord can still be installed but it will not seat flush against the chassis. Engineering evaluation noted that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. Scratches were 0. 065 - 0. 100 in length and were not aligned with the tube axis. Engineering concluded that the damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3111631 | 3083264 | 2979798 | 2955842201301441 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-04-25 | 2013-04-29 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-25 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-29 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the pk dissecting forceps instrument was noted to be broken. The instrument was replaced with another pk dissecting forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. The main tube had various scratch marks on the distal end showing light material removal. The scratches were short in length and not axially aligned with the main tube. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3112067 | 3083680 | 2974692 | 2955842201301442 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-02 | 2013-04-29 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-27 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the mega suturecut needle driver instrument was noted to have a separated wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found indentations on the edge of the distal pulley with some scratches on the surface. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3112091 | 3083685 | 2980219 | 2955842201301443 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-02 | 2013-04-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument separated between the orange part and the shaft. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was broken at the proximal end. The tube wall fractured at the junction between the pad printed section and the shoulder feature. The shoulder exhibited cracks 180 degrees apart that were consistent with excessive lateral force applied. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3112092 | 3083686 | 2980220 | 2955842201301444 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-02 | 2013-04-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-01 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the mega suturecut needle driver instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3112094 | 3083687 | 2980221 | 2955842201301445 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-02 | 2013-04-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-25 | Malfunction |  | It was reported that during a da vinci s pyeloplasty procedure the maryland bipolar forceps instrument was not responding to the system. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage was likely due to mishandling. The instrument passed electrical continuity testing. Engineering also found all four clamping pulleys corroded all around inside the housing. Engineering concluded that the damage was most likely due to improper cleaning. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3112095 | 3083706 | 2980240 | 2955842201301447 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-02 | 2013-04-29 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-11-09 | Malfunction |  | It was reported that during a da vinci s cystectomy procedure the monopolar curved scissors (mcs) instrument was not recognized by the system. The instrument was exchanged into a backup instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the housing end pogo pins were corroded around the pins and one pin was stuck. Due to the corrosion, the pogo pin was jammed down. The instrument passed electrical continuity testing. Engineering concluded the damage was most likely due to improper cleaning. Engineering also found corrosion around the clamping pulleys with the residue covered up the cables created stiffness to the cable movement. Engineering concluded this was most likely due to improper cleaning. Engineering also found an orange rusty like substance around the bearings. This most likely was due to improper cleaning. Engineering also found the distal end of main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3112093 | 3083707 | 2980241 | 2955842201301446 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-02 | 2013-04-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-02 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-03 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the monopolar curved scissors (mcs) instrument orange part was peeling off the tip of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument tube extension was cracked at the proximal clevis interface. Engineering concluded the tube extension likely broke due to excessive side loading or other mishandling. The instrument passed electrical continuity testing. Engineering also found a hairline crack at the distal end of the instrument main tube. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3112098 | 3083710 | 2980244 | 2955842201301448 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-04-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-02 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the mega suturecut needle driver instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damages on the edge and on the surface. Engineering concluded that the damage was likely due to excess contact and mishandling. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3112104 | 3083715 | 2980249 | 2955842201301450 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-04-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-27 | Malfunction |  | It was reported that during a da vinci s sacrocolpopexy procedure the mega suturecut needle driver instrument cable snapped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. Engineering also found scratches on the surface of the main tube. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3112106 | 3083718 | 2980252 | 2955842201301449 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-04-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-19 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure the pk dissecting forceps instrument cable snapped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal end. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. Engineering also found scratch marks on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3112130 | 3083723 | 2980257 | 2955842201301451 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-03 | 2013-04-29 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-03 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-03 | Malfunction |  | Event Description: --/-- It was reported that during reprocessing the da vinci s surgical monopolar curved scissors (mcs) instrument orange part at the end of the instrument was noted to have fallen apart. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to (b)(4) of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was broken and missing a piece at the proximal clevis interface. The missing piece measured roughly about 0. 267 x 0. 253. The clevis was dislodged from the tube extension. The tube extension likely broke due to excessive side loading. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. Insulation damage was found on the main tube by the extension tube side. The length of the damage measured about 0. 202. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3112133 | 3083726 | 2980260 | 2955842201301453 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-04-03 | 2013-04-29 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
| 2013-04-16 | Malfunction |  | During laparoscopic robotic assisted tah, surgeon noted that the endowrist instrument, maryland bipolar forceps 8mm had sparked while in use and the tips were not closing properly. He stated that there was no harm to the patient. The instrument was removed from the field and replaced with another bipolar. Device #1is this a laboratory device or laboratory test?no. |  | DAVINCI | SYSTEM SURGICAL COMPUTER CONTR | 3141020 | 3113448 | 3010004 | 3113448 | 1 | NAY | User Facility |  |  | Initial | 2013-04-29 | 2013-04-29 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2013
| 2013-04-23 | Malfunction |  | Cadiere forceps used in robotic surgery was not working properly. Device #1is this a laboratory device or laboratory test?no. |  | DAVINCI CADIERE | SYSTEM SURGICAL COMPUTER CONTR | 3141049 | 3113453 | 3010009 | 3113453 | 1 | NAY | User Facility |  |  | Initial | 2013-04-29 | 2013-04-29 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
42017207 | 2013-03-22 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted that tips of the maryland bipolar forceps instrument didn't meet. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a small offset of grips at tips. The offset was less than the maximal allowable specification. Additional observations not reported by site were char mark on the yaw pulley and main insulation tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were 0. 100 - 0. 300 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. One yaw pulley cover exhibited a char mark on the outer surface at the grip base. No char marks were present on either of the yaw pulleys between the grips. Instrument passed electrical continuity passes. Internal laboratory testing has determined that the arc track failure mode is most likely to occur when a bipolar instrument is energized with no tissue between the grips. Internal arcing in bipolar instruments results in a potential loss of product function, but no transfer of energy to the patient. Intuitive has concluded that the likelihood of injury due to this failure is remote. User are instructed to retain back-up instruments in case that an instrument fails to perform, and are also instructed in proper technique to minimize the likelihood of internal arcing. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3107100 | 3079341 | 2975855 | 2955842201301413 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-01 | 2013-04-26 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-20 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the customer noted that the pk dissecting forceps instrument would not cauterize, the cord was changed prior to replacing the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with the yaw pulley broken, a piece had broken off and it was removed from the pulley. The conductor wire connection was also broken. Instrument failed electrical continuity test. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3107101 | 3079342 | 2975856 | 2955842201301415 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-01 | 2013-04-26 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42011006 | 2013-04-01 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the customer noted broken cable at the end of the precise bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PRECISE TM BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3107092 | 3079350 | 2975871 | 2955842201301412 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-01 | 2013-04-26 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-25 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted that the shaft of the monopolar curved scissors instrument broke. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension was found to be broken with a missing piece at the proximal clevis interface. The missing piece measured roughly about 0. 186 x 0. 046. The clevis was dislodged from tube extension. Engineering concluded that damage was likely due to excessive side loading. The instrument passed electrical continuity test. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3107086 | 3079356 | 2975877 | 2955842201301414 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-29 | 2013-04-26 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-04-05 | Malfunction |  | The reporter indicated that the prograsp forceps instrument reportedly could not move anymore during a da vinci si prostatectomy procedure. The user facility attempted to use the emergency wrench but the instrument was stuck. The instrument was removed with the trocar. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. The motion in the grip was non-intuitive due to tight cables. The cables were derailed at the distal idler and the clamping pulley. Other backend cables were not damaged. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. Engineering concluded that the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3107946 | 3080036 | 2976570 | 2955842201301402 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-04-19 | 2013-04-26 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-04-19 | Device Not Retu | No | Reuse | 2013
40018010 | 2013-03-27 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgeon observed a spark coming from the monopolar curved scissors (mcs) tip cover accessory installed on the mcs instrument. Examination of the mcs tip cover accessory by the surgical staff found an unspecified broken portion. The planned surgical procedure was completed with a replacement mcs tip cover accessory and without any other problems. No patient harm, adverse outcome, or injury was reported. | The tip cover has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, intuitive surgical inc. (isi) contacted the distributor to obtain additional information regarding the reported event. The mcs tip cover accessory was not available for return to isi for evaluation. The distributor indicated that the mcs tip cover accessory was inspected prior to use and was installed on the mcs instrument using the installation tool. The distributor was unable to indicate if electrolube was applied to the mcs instrument prior to installation of the mcs tip cover accessory. The cannula was inspected with a pin gage prior to use. The distributor was unable to provide the electrosurgical generator unit (esu) settings used during the reported surgical procedure. The distributor indicated that the mcs instrument collided with another instrument during the surgical procedure. However, the distributor was unable to provide details of the damage found on the mcs tip cover accessory. The distributor indicated that the surgical procedure was recorded. However, the distributor declined to provide a copy of the video to isi for evaluation. The distributor denied that the patient sustained any injuries or complications because of the reported mcs tip cover accessory issue. The distributor indicated that the patient was discharged from the hospital with no reported problems. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3109259 | 3081313 | 2977847 | 2955842201301419 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2013-04-01 | 2013-04-26 | No | Yes | Health Professional | 40018010 | No | No | 2013-04-01 | No | Yes | Initial | 2013
42009308 | 2013-03-12 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy the prograsp forceps instrument did not have a full range of motion. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. When articulating the instrument manually, the movement was not precise. The roll input and pitch input disc did not move smoothly when manually rotated. Visual inspection showed the front roll axis bearing and three bearings below clamping pulleys exhibited contamination that was possibly causing increased friction. Engineering concluded improper cleaning may contribute to issue. Additional damage found was deep scratches on the main tube. The distal end of main tube had a couple scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage was likely due to mishandling. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3109305 | 3081332 | 2977832 | 2955842201301420 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | Initial | 2013-03-28 | 2013-04-26 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-03-27 | Malfunction |  | It was reported that during a da vinci si surgical procedure, pieces of insulation fell off the permanent cautery spatula instrument. There were no missing pieces or pieces falling into the patient reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with the ceramic sleeve broken at the base of the spatula, with material removed. Engineering concluded the damage was likely due to excess contact and mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3109312 | 3081347 | 2977861 | 2955842201301416 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-27 | 2013-04-26 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-27 | Malfunction |  | A prograsp forceps instrument was returned to isi without requesting an return materials authorization (rma) or a complaint. The customer did not respond to additional requests for information. | The instrument was returned and evaluated. A complaint for the instrument was not submitted. Engineering found the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3109268 | 3081350 | 2977864 | 2955842201301417 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-27 | 2013-04-26 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-03-05 | Malfunction |  | It was reported that during a da vinci si total hysterectomy the tip of the permanent cautery spatula was chipped. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The complaint of the instrument tip was chipped was confirmed. The instrument was found with the ceramic sleeve broken at the base of the spatula, with a piece removed. Engineering confirmed the damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3109330 | 3081353 | 2977867 | 2955842201301418 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-26 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-04-01 | Malfunction |  | During central processing, the user facility identified frayed wires on the mega suturecut needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering did not confirm the alleged complaint. The grip cable was found broken not frayed at the distal idlers and the cable segment was sticking out at wrist. The idler pulley spun freely but was found with damage. The distal idler pulley had an indentation at the edge. Engineering concluded that the damage may be due to mishandling. Other cables at instrument's wrist were not damaged. Additional observation not initially reported by the site was corrosion on the back idler pulleys. Yellowish residue was found in between each of the pulleys at the edges. Engineering concluded that the damage may be due to improper cleaning. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURE CUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3109359 | 3081381 | 2977915 | 2955842201301422 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-01 | 2013-04-26 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-03-28 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff noticed a fragment from the permanent cautery hook instrument had broken off and fell into the patient. The broken fragment was successfully retrieved and the surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was returned with the yaw pulley boss feature broken off. The broken fragment was not returned to intuitive surgical inc. (isi) for evaluation. The missing fragment measured approximately. 060 diameter by. 080 long. The boss feature that retains the conductor cap in place was found to be fractured at its base. The conductor cap was observed to be loose and was able to be pulled out of distal clevis as a result of the breakage. The boss feature on the opposite side was intact and no damage was observed on the distal clevis. On (b)(6) 2013, isi contacted the clinical coordinator at the site to obtain additional information regarding the reported event. The clinical coordinator indicated that the permanent cautery hook instrument was inspected prior to use and the instrument was used during the entire surgical procedure involving the patient's gallbladder. She indicated that the device did not collide with any other instruments during the surgical procedure and no x-rays were performed on the patient to locate the broken instrument piece. On (b)(6) 2013, isi contacted the isi clinical sales representative (csr) who was present during the reported surgical procedure. The csr indicated that at the end of the surgical procedure the bed side assistant successfully retrieved the fragment from the patient using an unspecified laparoscopic instrument through an existing port site. The csr stated that the patient did not experience any complications because of the reported event and the patient is currently doing well. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3109393 | 3081419 | 2977933 | 2955842201301421 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-26 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
40018012 | 2013-04-03 | Injury | Required Intervention | It was reported that during a da vinci s hysterectomy procedure, while the surgeon was using the monopolar curved scissors instrument with the mcs tip cover installed, the pt's fallopian tube and bowel were burned. | The instrument has not been returned for eval; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering eval) or if additional information is received. On (b)(4) 2013, the isi clinical sales rep (csr) who reported this event was contacted. The csr indicated she was not present during the surgical procedure when the injury to the pt occurred; however, the surgical staff informed her that arcing from the monopolar curved scissors instrument with the tip cover accessory installed, was not observed. The surgical staff indicated that the tip cover was inspected and that the dark gray area of the tip cover appeared to have a "slit. " a colorectal surgeon was consulted to assess and repair the damage to the pt's bowel. The colorectal surgeon performed a resection to repair the damage to the pt's bowel. On (b)(4) 2013, the csr contacted the colorectal surgeon regarding the pt's status and the colorectal surgeon indicated that the pt recovered well. Isi has made several attempts to contact the site to obtain additional information regarding the reported event; however, no additional information has been provided as of the date of this report. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3117159 | 3087051 | 2983595 | 2955842201301436 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-04-03 | 2013-04-26 | Yes | Yes | Health Professional | 40018012 | No | No | 2013-04-03 | Device Not Retu | Yes | Unkown | 2013
| 2013-04-19 | Malfunction |  | Noticed wire was coming through needle driver. No harm to patient. Needle driver broken and wire exposed during use. Replaced with new needle driver. No pt harm. What was the original intended procedure?robotic assisted myomectomy. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST MEGA SUTURECUT | SYSTEMSURGICALCOMPUTER CONTROL | 3124190 | 3095497 | 2992050 | 3095497 | 1 | NAY | User Facility |  |  | Initial | 2013-04-26 | 2013-04-26 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
42022703 | 2013-04-05 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable on the pk dissecting forceps instrument broke and a piece of "plastic" from the instrument fell into the patient. The instrument fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one conductor wire was broken at the yaw pulley exit. The wire was detached from its connection at the grip. The instrument passed electrical continuity testing. The pulley showed no signs of arcing. The yaw pulley was missing a 0. 143" x 0. 034" piece opposite of the conductor break, which allowed the grip to move within the pulley and have a larger range of motion in yaw. Engineering concluded that it was indeterminable as to how the piece of yaw pulley broke. Engineering evaluation also found that the surface of the distal pulley exhibited scratches. The pitch cable at the distal idler was frayed and the pulley and clevis did not exhibit damage. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage to the main tube was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(6) 2013, isi contacted the isi clinical sales representative who reported this event. The csr indicated that he was present during the surgical procedure. Per the csr, the surgeon had completed the surgical task of taking down the patient's infundibulopelvic ligament when the surgeon observed that a foreign object was lying on the patient's tissue in the surgical area where the surgeon was working. The csr indicated that when the surgeon zoomed in to verify what the object was, the surgeon discovered that the foreign object was a fragment from the pk dissecting forceps instrument. The csr indicated that the broken instrument piece was removed laparoscopically using the da vinci surgical system. The procedure was using a replacement instrument of the same type. The csr indicated that he did not observe that the instrument made contact with another instrument during the surgical procedure. Per the csr, there was no report to him by the surgical staff that the patient suffered any complications as a result of the reported event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3132376 | 3096887 | 2993422 | 2955842201301437 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-04-05 | 2013-04-26 | No | Yes | Health Professional | 42022703 | Yes | No | 2013-04-05 | Yes | No | Reuse | 2013
42020706 | 2013-03-27 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wire on the tenaculum forceps instrument broke and pieces from the instrument fell into the patient. The broken instrument pieces were retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Conclusions: the instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp was missing from the clevis. The clevis does not exhibit any damage or wear marks. The other cables at the wrist were not damaged. Engineering evaluation also found that the distal end of the main tube had various scratch marks with light material removal. The scratches were. 086"-. 185" in length and were not aligned with the tube axis. Engineering concluded that the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On april 16, 2013 isi contacted the initial reporter of this event. The initial reporter indicated that the tenaculum forceps instrument was inspected prior to use and there was no damage to the instrument observed. The initial reporter was unable to provide details concerning the surgical task the surgeon was performing when the instrument broke and there is no video recording of the surgical procedure available. Per the initial reporter, the surgeon did not indicate what he believes caused the instrument to break. The initial reporter indicated that no additional surgical procedure was required to remove the broken instrument fragments and the fragments were retrieved laparoscopically using the da vinci surgical system. The initial reporter believes that all of the fragments were retrieved and no x-ray was required and the site believes that all of the fragments were retrieved. The initial reporter indicated that to the best of his knowledge the patient did not suffer any post-surgical complications as a result of the reported event and the patient has not returned to the hospital. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3132428 | 3096888 | 2993423 | 2955842201301435 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-03 | 2013-04-26 | No | Yes | Health Professional | 42020706 | Yes | Yes | 2013-04-03 | Yes | No | Reuse | 2013
42020505 | 2013-04-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, part of the tip on the fenestrated bipolar forceps instrument broke off and fell into the patient. The broken fragment was retrieved from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with one grip broken off from the distal end. The entire fenestrated section of the grip fractured along the most proximal tooth. The fracture surface did not show any obvious voids in material under low magnification. Engineering concluded that the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, isi followed up with the operating room team lead at the site. The team lead indicated that the surgeon was performing dissection using the fenestrated bipolar forceps instrument when the grip on the instrument broke. The team lead indicated that the instrument was inspected prior to use and the instrument did not make contact with any other instrument during the surgical procedure. The team lead indicated that the broken fragment was retrieved laparoscopically using the da vinci si surgical system and the surgical procedure was completed with a replacement fenestrated bipolar forceps instrument. The team lead indicated that the patient did not experience any surgical complications as a result of the reported event and the patient is recovering well. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3132510 | 3096890 | 2993425 | 2955842201301434 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2013-04-03 | 2013-04-26 | No | Yes | Health Professional | 42020505 | Yes | No | 2013-04-03 | Yes | No | Initial | 2013
42019006 | 2013-03-19 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure that the cobra grasper instrument stopped working and the blue housing would not stay closed. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | Conclusions: the instrument was returned and evaluated. The reported complaint of the instrument blue housing would not stay closed was confirmed. Two of four housing snap tabs and all four adjacent housing pins were broken. The housing could be removed as a result of damage. Engineering concluded the damage may be due to mishandling. An additional observation found the instrument pitch cable was frayed at the distal clevis hub. No damage as found at the clevis. The frayed segment was approximately 0. 03" in length. No other cable damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3132141 | 3096934 | 2993487 | 2955842201301432 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-26 | No | Yes | Health Professional | 42019006 | Yes | No | 2013-03-28 | Yes | No | Reuse | 2013
42030901 | 2013-03-29 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted that the cable of the mega suture cut needle driver instrument snapped. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at the wrist. Other cables at wrist were not damaged. Engineering also found indentations in the middle of the distal pulley with some scratch marks on the surface. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 3104456 | 3076247 | 2972757 | 2955842201301383 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-29 | 2013-04-25 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-15 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that a cable on the pk dissecting forceps instrument broke and fragments fell into the patient. The broken fragments were retrieved. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. Frayed cable segments were found sticking out at the instrument's wrist. Engineering evaluation also found deep scratches on the instrument's main tube. The distal end of the instrument's main tube had a few scratch marks exhibiting light material removal and rough surface finish. Engineering evaluation concluded that the main tube damage was likely due to mishandling. The other cables of the instrument's wrist did not exhibit any damage. The instrument passed electrical continuity testing. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Intuitive surgical has attempted to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3105363 | 3077699 | 2974240 | 2955842201301384 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-26 | 2013-04-25 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-28 | Malfunction |  | During central processing, the user facility identified a broken cable on the fenestrated bipolar forceps instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. Engineering found a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3105510 | 3077856 | 2974369 | 2955842201301385 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-25 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-28 | Malfunction |  | During central processing, the user facility identified a broken cable on the megasuture cut needle driver instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint: the grip close cable was broken at the distal idlers and was sticking out at the instrument's wrist. The idler pulley spun freely and did not exhibit any damage. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3105532 | 3077931 | 2974432 | 2955842201301386 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-25 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-28 | Malfunction |  | During central processing, the user facility reportedly identified frayed wires on the tip of the fenestrated bipolar forceps instrument. Nothing reportedly fell into a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint: the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3105567 | 3077951 | 2974472 | 2955842201301387 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-25 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-27 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that a broken cable on the fenestrated bipolar instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be frayed at the distal clevis hub. The frayed segment was approximately 0. 03 in length. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3105615 | 3077972 | 2974485 | 2955842201301388 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-29 | 2013-04-25 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-22 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure, the customer reported that the maryland bipolar forceps instrument refused to open in and out of the robot and then it stayed widely opened. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on the in-house da vinci robot system and it was driven with no problem. Instrument passed recognition and engagement test. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Functional performance testing did not find any issues. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were 0. 070 - 0. 145 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3105830 | 3078174 | 2975325 | 2955842201301390 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-29 | 2013-04-25 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-28 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility reportedly identified a cable that has separated on the pk dissecting forceps. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint: the pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables ta the instrument's wrist were not damaged. Additional observation not initially reported by the site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were. 080 -. 150 in length and were not aligned with the main tube axis. Engineering concluded that the main tube damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3105828 | 3078191 | 2974692 | 2955842201301389 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-25 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
40017208 | 2013-03-13 | Malfunction |  | During a da vinci prostatectomy procedure, the bipolar prongs allegedly came out of the maryland bipolar forceps. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint: the bipolar insert was dislodged from the plug riser and was partially sticking out. The bipolar pins were slightly loose due to the dislodged insert. The insert may have not been fully seated in the chassis or may have been pulled out when the bipolar cord was removed. Inspection of the insert showed the retention tabs pinched inwards, which reduced the engagement with the plug riser. Additional observation not initially reported by the site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were. 060 -. 120 in length and were not aligned with the main tube axis. Engineering concluded that the main tube damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3105919 | 3078286 | 2961163 | 2955842201301391 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-29 | 2013-04-25 | No | Yes | Health Professional | 40017208 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-20 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the prograsp forceps instrument cable for the articulation guidance was noted to be torn. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch up cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3106374 | 3078634 | 2975155 | 2955842201301393 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-28 | 2013-04-25 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-27 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si esophagectomy procedure the mega suturecut needle driver instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip cable was broken at the distal idlers. The idler pulley spun freely but did exhibit some damage. The cable segment stuck out at the wrist. The other cables at the wrist were undamaged. Engineering also found deep scratches on the main tube and mechanical indentations on pulley. The distal end of main tube had scratch marks exhibiting light material removal and a rough surface finish. The scratches were short in length approximately. 22 to. 33. The distal idler pulley exhibited indentations and wear marks along the edges and on the surface. The evidence was inconclusive, but damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | Reuse | 3106356 | 3078635 | 2972757 | 2955842201301392 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2013-05-06 | 2013-04-25 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-05-06 | Device Not Retu | No |  | 2013
42020505 | 2013-03-19 | Malfunction |  | It was reported that during reprocessing the da vinci s surgical fenestrated bipolar instrument jaws were noted not to close properly. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint that the jaws would not close correctly. Engineering found no damage to the grips upon visual inspection. The instrument was placed on an in-house system and driven, recognition and engagement testing passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Engineering found main tube damage and a frayed cable. The instrument had various scratch marks on the distal end of the main tube showing light material removal. The scratches were. 070 -. 190 in length and not aligned with the tube axis. Engineering concluded the damage was likely caused by mishandling/misuse. The pitch down cable was frayed at the distal clevis hub. The frayed strands stuck out at the wrist. The other cables at the wrist were not damaged. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3106382 | 3078652 | 2975165 | 2955842201301394 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-29 | 2013-04-25 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-05 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure the monopolar curved scissors (mcs) instrument was noted to be dull, and not close properly. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering placed the instrument on an in-house system for a latex cut test and the scissors did not cut cleanly through (b)(4) latex. The latex was snagged at scissor tips. Engineering also found a melted tip and main tube damage (scratch marks and cracking). The one blade exhibited melting with thermal deformation at the tip. The deformed tip contributed to cutting issues. Engineering concluded the blade damage was likely due to mishandling/misuse during cautery activation. The distal end of main tube had a scratch mark with light material removal. The scratch was. 150 in length and not aligned with the tube axis. There was also a hairline crack at the distal end of the tube. The crack was oriented in the axial direction. Engineering concluded the main tube damage was likely caused by mishandling/misuse. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. These cracks were reported on the initial mdr report. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3106512 | 3078767 | 2975288 | 2955842201301395 | 1 | NAY | Manufacturer | OtherUser facility | Other | InitialFollowup | 2013-03-30 | 2013-04-25 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42017910 | 2013-03-15 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of instrument not recognized. Engineering installed the instrument on an in-house system and the system successfully recognized the instrument. The pogo pins were not stuck or contaminated. The instrument passed dcp (dallas chip programmer) verification. The instrument was driven and moved intuitively and the grips opened and closed. Engineering was engineering found scratches on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded the damages were likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3106530 | 3078772 | 2975293 | 2955842201301396 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-03-29 | 2013-04-25 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42009308 | 2013-03-29 | Malfunction |  | It was reported that during a da vinci si sacrolpopexy with hysterectomy procedure the surgeon was using the prograsp forceps instrument to manipulate the uterus. The surgeon was lifting on the posterior aspect of the uterus. The prograsp forceps instrument started acting funny and a wire was visualized. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch down cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. Engineering also found various scratch/wear marks at the distal end of the main tube showing light material removal. The scratches were. 050 - 1. 25 in length. Some scratches were axially aligned with the tube, but most were not. Engineering concluded the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3106534 | 3078776 | 2975297 | 2955842201301397 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-03-29 | 2013-04-25 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-20 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the maryland bipolar forceps instrument's grips didn't want to open anymore. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on the in-house da vinci robot system and it was driven with no problem. Instrument passed recognition and engagement test. The grips opened and closed properly. Instrument passed electrical continuity test. Additional observations not reported by site were bent grips tips and deep scratches on main tube. One grip was found to be bent causing side to side misalignment of grips. There was a 0. 026 offset at the tips. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3106551 | 3078792 | 2975325 | 2955842201301398 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-29 | 2013-04-25 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-19 | Malfunction |  | Event Description: --/-- The monopolar curved scissors instrument allegedly stopped moving in the middle of the da vinci si hysterectomy procedure. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the alleged complaint. The instrument was placed on an in house da vinci system and driven: recognition and engagement tests passed. The instrument moved intuitively with full range in all directions. The instrument's blades opened and closed properly and successfully passed the cut test. Electrical continuity testing was performed on the instrument and passed. Additional observations not initially reported by the site were damages to the main tube, banana plug, flush tube, and tube extension. A small crack was found near the tube extension at the distal end. The banana plug was bent at the bad of the chassis. The flush tube had 2 kinks near the luer plate inside the instrument's housing. The tube extension also exhibited spots with pad print removal. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. These cracks were reported on the initial mdr report. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3106571 | 3078820 | 2975341 | 2955842201301399 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-04-23 | 2013-04-25 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
40017208 | 2013-03-12 | Malfunction |  | It was reported that during a da vinci surgical procedure, the customer reported that in the body of the maryland bipolar instrument, the pins were loosen at the junction of the bipolar cable. The instrument could not be used again and had to be exchanged. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the bipolar insert was found to be dislodged from plug riser and it was partially sticking out. The insert could have not be fully seated in the chassis or it could have been pulled out when bipolar cord was removed. Inspection of insert showed retention tabs pinched inwards, which reduces engagement with the plug riser. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were 0. 080 - 0. 200 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3106620 | 3078877 | 2975398 | 2955842201301400 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-29 | 2013-04-25 | No | Yes | Health Professional | 40017208 | Device Returned To Manufacturer | No | 2013-03-29 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-04-01 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted that the thoracic grasper instrument was losing teflon. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument's main insulation tube appeared to have scratches and gouge marks, with materials removed at the distal end. Engineering concluded that damage was likely caused by excess force contact on the main tube surface. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3106665 | 3078903 | 2975444 | 2955842201301401 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-01 | 2013-04-25 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-27 | Malfunction |  | Event Description: --/-- During a da vinci si surgical procedure, the reporter alleged that the monopolar curved scissors instrument was not cutting. Nothing reportedly fell into a patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering did not confirm the alleged complaint. The instrument was placed on an in house da vinci system for a latex cut test. The scissors cut cleanly through 0. 006 latex. The blades were not damaged. Electrical continuity testing was performed on the instrument and passed. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. Engineering concluded that the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3106685 | 3078958 | 2975471 | 2955842201301403 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-04-23 | 2013-04-25 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42020505 | 2013-03-26 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the fenestrated bipolar forceps instrument was not able to grasp tissue. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the grips opened and closed fine when the input discs were manually rotated. The instrument was placed on an-house is3000 system and driven. The instrument passed recognition and engagement testing. The pitch motion was not intuitive, but grips opened and closed properly. Engineering also found a broken pitch cable and tube damage. The pitch up cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were undamaged. The distal end of the main tube had various scratch marks showing light material removal. The scratches were. 085 -. 200 in length and not aligned with the tube axis. Engineering concluded that the damage may have been caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3106767 | 3079029 | 2975543 | 2955842201301404 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-01 | 2013-04-25 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-30 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure the prograsp forceps instrument was noted to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch up cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3106797 | 3079035 | 2975549 | 2955842201301408 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-01 | 2013-04-25 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-04-01 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure the fenestrated bipolar forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3106800 | 3079038 | 2975552 | 2955842201301409 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-01 | 2013-04-25 | No | Yes | Health Professional | 42020505 | No | No | 2013-04-01 | No | No | Unkown | 2013
42019408 | 2013-04-10 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure the mega needle driver instrument cable snapped while the surgeon was suturing mesh. The instrument was removed and replaced with another needle driver instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was undamaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found scratches on the surface of the distal pulley. Engineering found a scratch mark on the distal end of the main tube showing light material removal. The scratch was aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3106814 | 3079041 | 2975555 | 2955842201301410 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-15 | 2013-04-25 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-04-01 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci s surgical procedure the monopolar curved scissors (mcs) instrument orange casing was noted to be cracked. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was cracked at the proximal clevis interface. The clevis was dislodged from the tube extension. Engineering concluded the tube extension likely cracked due to excessive side loading. The instrument passed electrical continuity testing. Engineering also found various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3106769 | 3079047 | 2975568 | 2955842201301405 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2013-04-01 | 2013-04-25 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-27 | Device Not Retu | No |  | 2013
42000606 | 2013-04-01 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical large needle driver instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the cable was frayed. The one grip close cable was frayed at the distal idlers. The frayed cable segment stuck out at the wrist. The other cables at the wrist were undamaged. Engineering also found a damaged pulley. The distal idler pulley on which the frayed cable was seated had a section that was bent inwards. Engineering concluded that the pulley damage was likely due to mishandling and the pulley damage likely contributed to the fraying. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3106792 | 3079052 | 2975573 | 2955842201301406 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-04-01 | 2013-04-25 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-30 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure the pk dissecting forceps instrument was noted to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that both pitch cables were broken at the distal clevis hub. The cable segments that contained the crimp were still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3106785 | 3079056 | 2975577 | 2955842201301407 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-04-01 | 2013-04-25 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-26 | Malfunction |  | It was reported that during reprocessing the da vinci s surgical maryland bipolar forceps instrument jaws would not open. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering placed the instrument on an in-house system and recognition and engagement testing passed. The grips opened and closed properly. Engineering found the instrument had a bent grip and deep scratches on the main tube. One grip is bent, causing side-to-side misalignment of the grips. There was a. 033 offset at the tips. Engineering concluded the damage was likely due to mishandling. The distal end of the main tube had several scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3106822 | 3079088 | 2975602 | 2955842201301411 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-04-01 | 2013-04-25 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-04-01 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-25 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the maryland bipolar forceps instrument cable was noted to be frayed and separated. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the conductor wires were intact and undamaged, but the drive cable was frayed. The pitch up cable was frayed at the distal clevis hub. The frayed strands stuck out at the wrist. The other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3101942 | 3074036 | 2971759 | 2955842201301356 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-25 | Malfunction |  | It was reported that during a da vinci s surgical procedure the pk dissecting forceps instrument was noted to have a cable sticking out. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the idler block pulley. The cable segment at the distal clevis hub was found loose. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3101909 | 3074037 | 2970571 | 2955842201301357 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-27 | 2013-04-24 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-03-19 | Malfunction |  | It was reported that during a da vinci si surgical procedure a cable broke on the mega needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument had a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3101958 | 3074045 | 2962748 | 2955842201301358 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42004906 | 2013-03-14 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the cadiere forceps instrument tip was noted to be bent. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of tip is bent. Engineering found the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp remained in the clevis. The clevis did not exhibit excessive damage. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3101940 | 3074047 | 2970582 | 2955842201301353 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42020503 | 2013-03-16 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the fenestrated bipolar forceps instrument shaft broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was broken at the proximal clevis to main tube interface. The clevis was dislodged from the main tube as a result. The main tube had some missing pieces. Engineering concluded that the damage was likely due to overloading at the tip or other type of mishandling. Engineering also found a deep scratch on the main tube at the distal end exhibiting light material removal and a rough surface finish. Engineering concluded that the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3101941 | 3074048 | 2970583 | 2955842201301354 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42020503 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-25 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the prograsp forceps instrument cable was noted to be frayed and separated. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch down cable was frayed at the distal clevis hub. The frayed strands stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found various scratch marks on the main tube at the distal end showing light material removal. The scratches were. 090 -. 200 in length and not aligned with the tube axis. Engineering concluded the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3101908 | 3074049 | 2970584 | 2955842201301355 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-02-27 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy the mega needle driver instrument was not corresponding to the moment of the surgeon's hands. The instrument was removed and reinstalled to troubleshoot the problem but it remained. It was then noted that the cables of the instrument had broken. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | Instrument was received and evaluated. Engineering found both grip cables were broken at the distal idler pulley. The idler pulley spun freely but was damaged. The cable segments stuck out at the wrist. The pitch cables at the wrist were not damaged. Engineering also found wear marks and indentations on the edges of the distal idler pulley. In addition, the distal end of the main tube had a scratch mark exhibiting light material removal and a rough surface finish. Engineering concluded the damages were likely due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3101975 | 3074063 | 2970598 | 2955842201301359 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure that a wire was sticking out of the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with the grip cable frayed and sticking out at distal clevis. The same cable was derailed off the distal idler pulley. Additional damage found were corroded back idler pulleys. Engineering concluded improper cleaning was the likely cause of the corrosion. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3101976 | 3074064 | 2970599 | 2955842201301360 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-25 | Malfunction |  | It was reported that prior to starting a surgical procedure that a wire frayed and separated on the prograsp forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint of frayed wires. For clarification, the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3101977 | 3074065 | 2970600 | 2955842201301361 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
40000621 | 2013-03-27 | Malfunction |  | It was reported that during a da vinci surgical procedure, the customer noted a bowden torn, not possible to stitch anymore, large needle driver instrument could not close anymore. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at the wrist. Other cables at wrist were not damaged. Engineering also found indentations on the edge of the distal pulley. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3103101 | 3075061 | 2971595 | 2955842201301362 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-27 | 2013-04-24 | No | Yes | Health Professional | 40000621 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-28 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted broken cables on the maryland bipolar forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was found frayed not broken at distal clevis hub. No other cables were found to be damaged at distal clevis. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3103273 | 3075224 | 2971759 | 2955842201301363 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-24 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-25 | Malfunction |  | It was reported that prior to starting a da vinci si total benign hysterectomy procedure the wire separated on the fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported issue. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3103449 | 3075341 | 2971873 | 2955842201301364 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
40000621 | 2013-02-01 | Malfunction |  | It was reported that during a da vinci surgical procedure, a bowden of branch was torn on and the needle couldn't be moved anymore on the large needle driver instrument there were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. The other cables at the wrist were not damaged. There were also scratches on the surface of the distal pulley. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3103674 | 3075481 | 2972013 | 2955842201301365 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-27 | 2013-04-24 | No | Yes | Health Professional | 40000621 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
42027802 | 2013-03-27 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical grasping retractor instrument was noted to have broken wires and the blue housing was coming apart. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal end. The cable segment that contained the crimp was missing from the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. Two of the four housing snap tabs and two pins were broken off housing. The housing was easily removed as a result of the damage. There was a crack on the back end of the housing right above the luer plate. Engineering concluded that damage to the housing was likely due to mishandling. Engineering also found deep scratches on the distal and proximal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3103793 | 3075560 | 2972092 | 2955842201301366 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-27 | 2013-04-24 | No | Yes | Health Professional | 42027802 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-07 | Malfunction |  | It was reported that during a da vinci si cystectomy procedure that the fenestrated bipolar forceps instrument would not cauterize. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint that the instrument would not cauterize. The instrument passed an electrical continuity test, and was able to perform cautery functions on the system with no trouble found; however, the instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3103827 | 3075622 | 2899510 | 2955842201301368 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-18 | Malfunction |  | It was reported that during reprocessing the tip of the prograsp forceps instrument did not move freely. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was placed on a system and driven. Recognition and engagement test passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. Additional observations found was the instrument pitch cable was frayed at the distal clevis hub. The instrument main tube was also found with multiple scratches at the distal end. Engineering concluded the damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3103846 | 3075633 | 2972165 | 2955842201301367 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-03-28 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken ceramic sleeve on the permanent cautery spatula instrument. The instrument reportedly was not used on a patient after the reported issue was identified. Nothing reportedly fell into a patient and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that a piece of the ceramic sleeve was broken off and missing. Engineering concluded that the damaged ceramic sleeve may be due to mishandling. Electrical continuity testing was performed on the instrument and passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3103860 | 3075644 | 2972176 | 2955842201301369 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-24 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-28 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a bent grip on the pk dissecting forceps instrument. The instrument reportedly was not used on a patient after the reported issue was identified. Nothing reportedly fell into a patient and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that one of the instrument's grip was bent, which caused side to side misalignment of the grips. There was a 0. 1725 offset at the tips. Engineering concluded that the damaged grip may be due to mishandling. Additional observations that were not initially reported by the site were frayed pitch cable and deep scratches on the main tube. The instrument had a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The distal end of the main tube exhibited scratches with light material removal and had a rough surface finish. Engineering concluded that the damages may be due to mishandling. Electrical continuity testing was performed on the instrument and passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however,. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3103876 | 3075665 | 2972197 | 2955842201301370 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-24 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-28 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility reportedly identified a bent grip on the pk dissecting forceps instrument. The instrument was not used on a patient after the reported issue was identified. Nothing reportedly fell into a patient and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering did not find any damage with the instrument's grip. Additional observation that was not initially reported by the site was a broken conductor wire at the yaw pulley. The conductor wire was sticking out at the distal clevis. The instrument also had deep scratches on the main tube. The distal end of the main tube had several scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded that the damages may be due to mishandling. Electrical continuity testing was performed on the instrument and failed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3103929 | 3075715 | 2972197 | 2955842201301371 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-24 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-26 | Malfunction |  | It was reported that during reprocessing that the cable was broken on the pk dissecting forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported broken wire, however, for clarification the damaged instrument component was a frayed pitch cable. The instrument was found with a frayed pitch cable at the distal idler. No damage found on the pulley or clevis. Electrical continuity test passed. Additional findings were scratches on the surface of the distal pulley and various scratches on the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3103964 | 3075762 | 2972294 | 2955842201301373 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-26 | Malfunction |  | It was reported that during reprocessing that broken cables were found on the large needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported broken wire, however, for clarification the damaged instrument component was a frayed pitch cable. The instrument was found with a frayed pitch cable at the distal idler. No damage was found on the pulley and clevis. Additionally, an indentation was found at the edge of the distal pulley. Engineering concluded the damage was most likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3103986 | 3075770 | 2972300 | 2955842201301372 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-19 | Malfunction |  | During a da vinci si surgical procedure, the large needle driver instrument reportedly broke during the case. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3104003 | 3075790 | 2972322 | 2955842201301376 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-26 | 2013-04-24 | No | Yes | Health Professional | 42000606 | No | No | 2013-03-26 | No | No | Reuse | 2013
400177 | 2013-03-29 | Malfunction |  | It was reported that during cleaning and sterilization, the customer complained of the 5mm bowel grasper instrument shaft wear. | The investigation conducted by the field service engineer (fse) concluded that shaft wear, reported by the site was associated to the incorrect use of the cannula. The fse visited the site and he observed gouges and scratch marks on the 5mm bowel instrument's shafts. During his visit, fse was also became aware that the customer was not using 5mm cannulas or the 8mm to 5mm cannula reducers. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: always use the appropriate intuitive surgical cannula reducer and seal to ensure intuitive motion when inserting endowrist instruments through cannula with larger inner diameters (e. G. , 5 mm instrument with an 8 mm instrument cannula). Additional observation not reported by site was film build-up on the 5mm instrument's shafts. Fse observed that the customer was manually cleaning the instruments and using an automated ultrasonic cleaner (da vinci sonic pro). The customer was not using a washer disinfector. The film build-up was visible after ultrasonic cleaning. Manual scrubbing and rinsing instrument shafts after ultrasonically cleaning removed the film build-up. The cleaning and sterilization user manual specifically states: o when scrubbing the tip of cautery instruments, take care not to damage the insulation. Isi contacted initial reporter robotics coordinator, she indicated that the issue was identified and corrected and they have not experienced any further issues. She also mentioned that they experienced this issue with a total of 6 bowel graspers instruments. Related mdrs numbers: 2955842-2013-01375, 2955842-2013-01377, 2955842-2013-01378, 2955842-2013-01379 and 2955842-2013-01381. | 5MM BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3104055 | 3075820 | 2972400 | 2955842201301374 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-29 | 2013-04-24 | No | Yes | Health Professional | 400177 | No | No | 2013-03-29 | No | No | Unkown | 2013
400177 | 2013-03-29 | Malfunction |  | It was reported that during cleaning and sterilization, the customer complained of the 5mm bowel grasper instrument shaft wear. | The investigation conducted by the field service engineer (fse) concluded that shaft wear, reported by the site was associated to the incorrect use of the cannula. The fse visited the site and he observed gouges and scratch marks on the 5mm bowel instrument's shafts. During his visit, fse was also became aware that the customer was not using 5mm cannulas or the 8mm to 5mm cannula reducers. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings warning: always use the appropriate intuitive surgical cannula reducer and seal to ensure intuitive motion when inserting endowrist instruments through cannula with larger inner diameters (e. G. , 5 mm instrument with an 8 mm instrument cannula). Additional observation not reported by site was film build-up on the 5mm instrument's shafts. Fse observed that the customer was manually cleaning the instruments and using an automated ultrasonic cleaner (da vinci sonic pro). The customer was not using a washer disinfector. The film build-up was visible after ultrasonic cleaning. Manual scrubbing and rinsing instrument shafts after ultrasonically cleaning removed the film build-up. The cleaning and sterilization user manual specifically states: o when scrubbing the tip of cautery instruments, take care not to damage the insulation. Isi contacted initial reporter robotics coordinator, she indicated that the issue was identified and corrected and they have not experienced any further issues. She also mentioned that they experienced this issue with a total of 6 bowel graspers instruments. Related mdrs numbers: 2955842-2013-01374, 2955842-2013-01377, 2955842-2013-01378, 2955842-2013-01379 and 2955842-2013-01381. | 5MM BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3104036 | 3075824 | 2972400 | 2955842201301375 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-29 | 2013-04-24 | No | Yes | Health Professional | 400177 | No | No | 2013-03-29 | No | No | Unkown | 2013
400177 | 2013-03-29 | Malfunction |  | It was reported that during cleaning and sterilization, the customer complained of the 5mm bowel grasper instrument shaft wear. | The investigation conducted by the field service engineer (fse) concluded that shaft wear, reported by the site was associated to the incorrect use of the cannula. The fse visited the site and he observed gouges and scratch marks on the 5mm bowel instrument's shafts. During his visit, fse was also became aware that the customer was not using 5mm cannulas or the 8mm to 5mm cannula reducers. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings warning: always use the appropriate intuitive surgical cannula reducer and seal to ensure intuitive motion when inserting endowrist instruments through cannula with larger inner diameters (e. G. , 5 mm instrument with an 8 mm instrument cannula). Additional observation not reported by site was film build-up on the 5mm instrument's shafts. Fse observed that the customer was manually cleaning the instruments and using an automated ultrasonic cleaner (da vinci sonic pro). The customer was not using a washer disinfector. The film build-up was visible after ultrasonic cleaning. Manual scrubbing and rinsing instrument shafts after ultrasonically cleaning removed the film build-up. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Isi contacted initial reporter robotics coordinator, she indicated that the issue was identified and corrected and they have not experienced any further issues. She also mentioned that they experienced this issue with a total of 6 bowel graspers instruments. Related mdrs numbers: 2955842-2013-01374, 2955842-2013-01375, 2955842-2013-01378, 2955842-2013-01379 and 2955842-2013-01381. | 5MM BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3104037 | 3075825 | 2972400 | 2955842201301377 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-29 | 2013-04-24 | No | Yes | Health Professional | 400177 | No | No | 2013-03-29 | No | No | Unkown | 2013
42020505 | 2013-03-28 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken cable on the fenestrated bipolar instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed there was a protruding broken conductor wire at the proximal clevis and the wires were broken. Additional observation not initially reported by the site were deep scratches on the main tube and bent grip tips. The distal end of the main tube had several scratch marks exhibiting light material removal and a rough surface finish. One grip was bent causing side to side misalignment of the grips. There was a 0. 037 offset at the tips. Engineering concluded that the damages may be due to mishandling. Electrical continuity testing was performed on the instrument and passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3104034 | 3075828 | 2972318 | 2955842201301380 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-24 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-03-12 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the permanent cautery spatula instrument had limited wrist movement. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint. Engineering found a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3104074 | 3075866 | 2961025 | 2955842201301382 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-28 | 2013-04-24 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No | Reuse | 2013
400177 | 2013-03-29 | Malfunction |  | It was reported that during cleaning and sterilization, the customer complained of the 5mm bowel grasper instrument shaft wear. | The investigation conducted by the field service engineer (fse) concluded that shaft wear, reported by the site was associated to the incorrect use of the cannula. The fse visited the site and he observed gouges and scratch marks on the 5mm bowel instrument's shafts. During his visit, fse was also became aware that the customer was not using 5mm cannulas or the 8mm to 5mm cannula reducers. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - warning: always use the appropriate intuitive surgical cannula reducer and seal to ensure intuitive motion when inserting endowrist instruments through cannula with larger inner diameters (e. G. , 5 mm instrument with an 8 mm instrument cannula). Additional observation not reported by site was film build-up on the 5mm instrument's shafts. Fse observed that the customer was manually cleaning the instruments and using an automated ultrasonic cleaner (da vinci sonic pro). The customer was not using a washer disinfector. The film build-up was visible after ultrasonic cleaning. Manual scrubbing and rinsing instrument shafts after ultrasonically cleaning removed the film build-up. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Isi contacted initial reporter robotics coordinator, she indicated that the issue was identified and corrected and they have not experienced any further issues. She also mentioned that they experienced this issue with a total of 6 bowel graspers instruments. Related mdrs numbers: 2955842-2013-01374, 2955842-2013-01375, 2955842-2013-01377, 2955842-2013-01379 and 2955842-2013-01381. | 5MM BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3104075 | 3075867 | 2972400 | 2955842201301378 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-29 | 2013-04-24 | No | Yes | Health Professional | 400177 | No | No | 2013-03-29 | No | No | Unkown | 2013
400177 | 2013-03-29 | Malfunction |  | It was reported that during cleaning and sterilization, the customer complained of the 5mm bowel grasper instrument shaft wear. | The investigation conducted by the field service engineer (fse) concluded that shaft wear, reported by the site was associated to the incorrect use of the cannula. The fse visited the site and he observed gouges and scratch marks on the 5mm bowel instrument's shafts. During his visit, fse was also became aware that the customer was not using 5mm cannulas or the 8mm to 5mm cannula reducers. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - warning: always use the appropriate intuitive surgical cannula reducer and seal to ensure intuitive motion when inserting endowrist instruments through cannula with larger inner diameters (e. G. , 5 mm instrument with an 8 mm instrument cannula). Additional observation not reported by site was film build-up on the 5mm instrument's shafts. Fse observed that the customer was manually cleaning the instruments and using an automated ultrasonic cleaner (da vinci sonic pro). The customer was not using a washer disinfector. The film build-up was visible after ultrasonic cleaning. Manual scrubbing and rinsing instrument shafts after ultrasonically cleaning removed the film build-up. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Isi contacted initial reporter robotics coordinator, she indicated that the issue was identified and corrected and they have not experienced any further issues. She also mentioned that they experienced this issue with a total of 6 bowel graspers instruments. Related mdrs numbers: 2955842-2013-01374, 2955842-2013-01375, 2955842-2013-01377, 2955842-2013-01378 and 2955842-2013-01381. | 5MM BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3104076 | 3075868 | 2972400 | 2955842201301379 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-29 | 2013-04-24 | No | Yes | Health Professional | 400177 | No | No | 2013-03-29 | No | No | Unkown | 2013
400177 | 2013-03-29 | Malfunction |  | It was reported that during cleaning and sterilization, the customer complained of the 5 mm bowel grasper instrument shaft wear. | The investigation conducted by the field service engineer (fse) concluded that shaft wear, reported by the site was associated to the incorrect use of the cannula. The fse visited the site and he observed gouges and scratch marks on the 5mm bowel instrument's shafts. During his visit, fse was also became aware that the customer was not using 5mm cannulas or the 8mm to 5mm cannula reducers. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - warning: always use the appropriate intuitive surgical cannula reducer and seal to ensure intuitive motion when inserting endowrist instruments through cannula with larger inner diameters (e. G. , 5 mm instrument with an 8 mm instrument cannula). Additional observation not reported by site was film build-up on the 5mm instrument's shafts. Fse observed that the customer was manually cleaning the instruments and using an automated ultrasonic cleaner (da vinci sonic pro). The customer was not using a washer disinfector. The film build-up was visible after ultrasonic cleaning. Manual scrubbing and rinsing instrument shafts after ultrasonically cleaning removed the film build-up. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Isi contacted initial reporter robotics coordinator, she indicated that the issue was identified and corrected and they have not experienced any further issues. She also mentioned that they experienced this issue with a total of 6 bowel graspers instruments. Related mdrs numbers: 2955842-2013-01374, 2955842-2013-01375, 2955842-2013-01377, 2955842-2013-01378 and 2955842-2013-01379. | 5MM BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3104077 | 3075869 | 2972400 | 2955842201301381 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-29 | 2013-04-24 | No | Yes | Health Professional | 400177 | No | No | 2013-03-29 | No | No | Unkown | 2013
42000606 | 2013-03-26 | Malfunction |  | During a da vinci s hysterectomy procedure, the large needle driver instrument reportedly broke during use, which exposed the wires. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering noted a broken grip close cable at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment was sticking out at the instrument's wrist. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3100416 | 3072817 | 2969353 | 2955842201301349 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-27 | 2013-04-23 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
42009303 | 2013-03-26 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility reportedly identified broken wires on the prograsp forceps instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that both pitch cables were broken at the distal end. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the instrument's wrist were not damaged. Additional observation not initially reported by the site were deep scratches on main tube, corroded bearings, and back idler pulleys. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded that the main tube damage may be due to mishandling. A visual inspection of the instrument showed front roll axis bearing, bearings below clamping pulleys, and back idler pulleys exhibiting corrosion that was possibly causing increased friction at the proximal end. Engineering concluded that the damage to the bearings and pulleys may be due to improper cleaning. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3100731 | 3072985 | 2969521 | 2955842201301350 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-26 | 2013-04-23 | No | Yes | Health Professional | 42009303 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-25 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the customer noted a damage wire on the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed grip cable at distal end. The frayed strands stuck out at the wrist. Other cables at wrist are not damaged. Engineering also found indentations on the edges and wear marks of the distal pulley. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3100916 | 3073166 | 2969702 | 2955842201301351 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-27 | 2013-04-23 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
42013904 | 2013-03-27 | Malfunction |  | During central processing, the user facility reportedly identified broken tip scissors from the schertel grasper instrument. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering could not confirm the alleged broken tip scissors complaint since this instrument is a grasper and not a scissors instrument. The instrument's grips opened and closed as expected when the input disc was manually rotated. No visible damage was noted for the instrument's snake wrist or grips. Additional observation not initially reported by the site was main tube insulation damage. The insulation was gouged in a couple of places at the distal end and had a section that was lifted up roughly 0. 650 above the instrument's snake wrist. No pieces were missing since the lifted up piece fully covered the exposed section of the shaft. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | SCHERTEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3101415 | 3073513 | 2970049 | 2955842201301352 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-27 | 2013-04-23 | No | Yes | Health Professional | 42013904 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
410322 | 2013-04-10 | Injury | Required Intervention | Da vinci vessel seal 8 mm - the cautery aspect of device not working properly. Da vinci vessel seal 8 mm - the distal end of device - the spring aspect - malfunctioned. |  | ENDO WRIST VESSEL SEALER | NONE | 3110224 | 3080398 | 2976899 | MW5029931 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-04-17 | 2013-04-23 | No | Yes | Health Professional | 410322 | Yes | Yes |  |  |  |  | 2013
400272 | 2013-04-22 | Malfunction |  | During use of harmonic ace curved shears insert on robotic procedure, the bottom tip of shears broke off and fell into patient cavity. Piece retrieved and entire instrument removed from patient without issue. After incident, spoke to company rep about incident and that it has happened in past once before and he stated that the surgeon is applying too much force while utilizing the shears. |  | HARMONIC ACE CURVED SHEARS INSERT | HARMONIC FOR ROBOT | 3114915 | 3084813 | 2981367 | MW5029943 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-04-23 | 2013-04-23 | No | Yes | Health Professional | 400272 | Yes | Yes |  |  |  |  | 2013
420275 | 2013-03-12 | Malfunction |  | Inserts for the harmonic instrument fell apart during attempted use (x2). Both inserts broke in the same manner while being used by a surgeon very familiar with the device. Physician was able to retrieve broken piece from patient with no residual harm to the patient. What was the original intended procedure?myomectomy. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST HARMONIC ACE | SYSTEMSURGICALCOMPUTER CONTROL | 3117104 | 3088893 | 2985447 | 3088893 | 1 | NAY | User Facility |  |  | Initial | 2013-04-18 | 2013-04-23 | No | Yes | Physician | 420275 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
42034301 | 2013-03-26 | Malfunction |  | During central processing, the user facility reportedly identified that the black insulation on the distal end of the thoracic grasper instrument was peeling. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the black insulation at the distal end of the instrument was peeling. The insulation was gouged on one side and had couple pieces missing (roughly 0. 4 and 01. 25 above the instrument's snake wrist. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3097909 | 3070739 | 2967273 | 2955842201301337 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-26 | 2013-04-22 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42023004 | 2013-03-25 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified broken cables on the tip of the hem-o-lok clip applier instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed the one of the grip close cable was frayed at the distal idler pulley and the frayed strands were sticking out at the wrist. Additional observation not initially reported by the site were derailed cables, main tube damage, and a cracked housing. The same grip cable that was frayed was also derailed from the distal idler pulley. Engineering concluded that the derailment most likely contributed to fraying. The main tube exhibited wearing with material removal and a rough surface finish all along its length. The instrument's housing was discolored and cracked on the top and side surfaces. Engineering concluded that the damage to the housing and main tube may be due to improper cleaning. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE HEMOLOK CLIP APPLIERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3098001 | 3070821 | 2967355 | 2955842201301338 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-25 | 2013-04-22 | No | Yes | Health Professional | 42023004 | Device Returned To Manufacturer | No | 2013-03-25 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-25 | Malfunction |  | During central processing, the user facility identified a broken cable on the large suturecut needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed the one of the grip close cable was broken at the distal idler pulley and the frayed strands were sticking out at the wrist. The idler pulley spun freely and had slight damages on the surface. The other cables at the instrument's wrist were not damaged. | LARGE SUTURECUT NEEDLEDRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3098022 | 3070846 | 2967380 | 2955842201301339 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-25 | 2013-04-22 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-25 | Device Not Retu | No | Reuse | 2013
41032204 | 2013-03-25 | Malfunction |  | It was reported that during a da vinci si colon resection procedure the endowrist one vessel sealer instrument was noted to have wires exposed at the shaft. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument had cables broken at the snake wrist. The first wrist disc at the distal end was found dislocated. The instrument blade remained inside the garage position. The instrument cautery cable was cut prior to returning. The system logs were searched and it was found that multiple incomplete cuts were found in the msc (mastery supervisor controller) logs followed with three system_err_tool_invalid_reason messages and normal usages. | ENDOWRIST ONE VESSEL SEALER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3098175 | 3070906 | 2967440 | 2955842201301340 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-25 | 2013-04-22 | No | Yes | Health Professional | 41032204 | Device Returned To Manufacturer | No | 2013-03-25 | Device Not Retu | No | Initial | 2013
42029602 | 2013-03-25 | Malfunction |  | During central processing, the user facility identified a broken cable on the large suturecut needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering noted a frayed cable at the distal idler pulley. The frayed cable section was approximately 0. 04 in length. No damage was found on the pulley. One grip close cable was derailed at the distal idler pulley. The grips were still able to open and close but the movement may not be precise. No damage was found on the distal idler pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLEDRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3098205 | 3070924 | 2967380 | 2955842201301341 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-25 | 2013-04-22 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-25 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-03-18 | Malfunction |  | During a da vinci si hemicolectomy procedure, the wire on the small grasping retractor instrument reportedly broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering noted a frayed pitch up cable at the distal clevis hub. The frayed strands were sticking out at the instrument's wrist. No other cables were damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3098286 | 3070968 | 2967502 | 2955842201301342 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-26 | 2013-04-22 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-03-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-27 | Malfunction |  | During a da vinci si surgical procedure, the user facility reportedly identified a broken cable (cable was in a loop) on the fenestrated bipolar forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering noted a broken pitch down cable at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. No other cables at the instrument's wrist were damaged. Additional observation not initially reported by the site was main tube and conductor wire damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were. 080 -. 140 in length and were not aligned with the main tube axis. One conductor wire was partially broken. The insulation was damaged near the proximal pulleys, which was adjacent to the cable breakage. Electrical continuity testing on the instrument was performed and passed. Engineering concluded that the main tube and conductor wire damage may have been caused by mishandling/misuse. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3098399 | 3071063 | 2967597 | 2955842201301343 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-27 | 2013-04-22 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-11-07 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the customer noted a broken wire on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was missing. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was corroded clamping pulleys. Upon initial inspection, roll axis had above average of friction. The housing was removed to check roll axis bearings. The bearings rotated properly. All clamping pulleys exhibited excessive amount of white residue around the cable grooves. Engineering concluded that the residue buildup was most likely due to improper cleaning. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3098410 | 3071091 | 2967625 | 2955842201301344 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-25 | 2013-04-22 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-03-25 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-03-25 | Malfunction |  | It was reported that during a da vinci si low anterior resection procedure the small grasping retractor instrument wires broke exposing wires. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at distal end. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. Engineering also found deep scratches on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3098486 | 3071151 | 2967685 | 2955842201301345 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-25 | 2013-04-22 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-03-25 | Device Not Retu | No | Reuse | 2013
42023004 | 2012-11-27 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the customer noted a broken wire on the large hem-o-lok clip applier instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was found to be broken at the distal end. The cable segment that contains the crimp was missing from the clevis. The clevis did not exhibit any damage or wear marks, but engineering found some contamination. Other cables at wrist were not damaged. Additional observation not reported by site was contaminated clamping pulleys and chip. Multiple clamping pulleys exhibited contamination around the groves and the screws. Rti chip exhibited white debris found on all the four pins. Engineering concluded that damage was likely due to improper cleaning. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | HEMOLOK CLIP APPLIER LARGEINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3098579 | 3071241 | 2967775 | 2955842201301346 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-25 | 2013-04-22 | No | Yes | Health Professional | 42023004 | Device Returned To Manufacturer | Yes | 2013-03-25 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-22 | Malfunction |  | During a da vinci si surgical procedure, the user facility reportedly identified a broken cable on the prograsp forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering noted a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3098856 | 3071505 | 2968039 | 2955842201301347 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-27 | 2013-04-22 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-27 | Device Not Retu | No | Reuse | 2013
42011006 | 2013-02-20 | Malfunction |  | It was reported that after a successfully a da vinci si adrenalectomy procedure, when the customer attempted to remove the precise bipolar forceps instrument, it got stuck in the cannula. According to the information provided by the customer, after trying the emergency grip release key, the surgeon had to remove the entire robotic arm, including the port and instrument. No injury to the patient occurred. After the instrument was removed, the surgeon discovered that the instrument exhibited a loose wire. Per the customer, the instrument worked perfectly during the case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The instrument was placed on in-house da vinci robot system and engineering was able to replicate the customer's complaint. The instrument was getting stuck at the cannula due to the pitch cable crimp was sticking out at distal clevis. The instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PRECISE TM BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3099202 | 3071844 | 2968378 | 2955842201301348 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-25 | 2013-04-22 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | Yes | 2013-03-25 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-08 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy a cable broke on the large suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken cable. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had slight damage on the surface. The cable segment stuck out at the instruments wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3094558 | 3067340 | 2963875 | 2955842201301317 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-21 | 2013-04-19 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-15 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the cable broke on the pk dissecting forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3094570 | 3067349 | 2963883 | 2955842201301318 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-21 | 2013-04-19 | No | Yes | Health Professional | 42022703 | No | No | 2013-03-21 | No | No | Reuse | 2013
42020505 | 2013-03-20 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted a broken cable on the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. Engineering also found a bent grip. One grip was found to be bent causing side to side misalignment of grips. There was a 0. 026 offset at the tips. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3095224 | 3068011 | 2964545 | 2955842201301320 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-20 | 2013-04-19 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42805206 | 2013-03-20 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, tip of the monopolar cautery singlesite instrument came off, this was noted before procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with the cautery adapter dislodged from the main tube. The monopolar cautery tip was returned and remained with the cautery adapter. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY SINGLESITEINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3095269 | 3068026 | 2964561 | 2955842201301321 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-19 | No | Yes | Health Professional | 42805206 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-03-20 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the customer noted a broken cable on the grasping retractor instrument, instrument was not used in case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a loose cable at the distal clevis, but no signs of a broken cable were observed initially. After the instrument was disassembled, a broken cable was observed located at the hypotubes, the hypotubes were stuck together at the distal end with dried out biodebris. Engineering concluded that damage was likely due to mishandling. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3095319 | 3068053 | 2964587 | 2955842201301322 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-19 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-07 | Malfunction |  | It was reported that during a davinci s surgical procedure, the customer noted a broken wire on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at the wrist. Other cables at wrist were not damaged. The instrument was also found with frayed pitch cable located at distal idler sticking out. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3095376 | 3068080 | 2964614 | 2955842201301323 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-19 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-03-15 | Malfunction |  | It was reported that during a da vinci s gynecology procedure, the customer noted a loose cable on the tenaculum forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3095434 | 3068112 | 2964646 | 2955842201301324 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-19 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-13 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the maryland bipolar forceps instrument wouldn't react properly to the surgeon control. The instrument was replaced and the surgical procedure was completed as planned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Engineering also found scratches on the surface of the distal pulley. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3095516 | 3068157 | 2964691 | 2955842201301325 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-20 | 2013-04-19 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-03-22 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure one of the cables broke on the tenaculum forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken cable. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. An additional finding was various scratches on the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3095613 | 3068200 | 2964735 | 2955842201301319 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-22 | 2013-04-19 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | Yes | 2013-03-22 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-07 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the customer noted that the prograsp forceps instrument was not functioning properly. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on the in-house da vinci robot system and driven without any problem. The instrument passed recognition and engagement test. The instrument was able to move intuitively with full range of motion in all directions. The grips were able to open and close properly. Instrument was found to be fully functional. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3095648 | 3068234 | 2927232 | 2955842201301326 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-19 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-01-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer reported that during the procedure the monopolar curved scissors instrument did not respond to the movements from the surgeon console. According to the information provided by the customer, the instrument was placed on the patient side manipulator (psm) arm 2 and the da vinci robot system could not recognize the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was inspected for recognition using the in-house da vinci robot system. The robot system successfully recognized the instrument, showing 2 uses left. The pogo pins did not stick and were not contaminated. The instrument was driven and moved intuitively, the grips were able to open and close. Additional observation not reported by site was a dull blade. The instrument was also placed on da vinci robot system for a latex cut test. The instrument's scissors did not cut clean through (b)(4) latex. The latex got snagged at the instrument scissors tips. The blade edges were not damaged, but the edges exhibited wear at the tips, negatively affecting cut performance. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3095750 | 3068278 | 2964813 | 2955842201301327 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-20 | 2013-04-19 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-20 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s surgical procedure the monopolar curved scissors (mcs) instrument orange piece was split near the tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was broken and missing a piece at the proximal clevis interface. The missing piece was found taped to the housing. Engineering concluded the tube extension likely broke due to excessive side loading or other mishandling. Engineering also found deep scratches on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded the damages were likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3095769 | 3068315 | 2961159 | 2955842201301328 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-03-28 | 2013-04-19 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-28 | Device Not Retu | No |  | 2013
42020505 | 2012-12-10 | Malfunction |  | It was reported that during a da vinci si surgical procedure the gynecologist complained that fenestrated bipolar forceps instrument was difficult to use. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal clevis hub, which was causing the lack of movement. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. An electrical continuity test passed. No other damage was found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3096022 | 3068502 | 2965056 | 2955842201301329 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-25 | 2013-04-19 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-25 | Device Not Retu | No | Reuse | 2013
42018305 | 2013-01-24 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the permanent cautery hook instrument was not recognized by the da vinci system robot. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was inspected for recognition on the in-house da vinci robot system. The system successfully recognized the instrument, showing 9 uses left. The pogo pins did not stick and were not contaminated. Instrument passed electrical continuity test. The instrument was driven and moved intuitively, the grips were able to open and close. Additional observation not reported by site was an indentation on the proximal clevis. There was a piece scraped off from the clevis. The scraped off piece measured roughly about 0. 043 x 0. 074. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3096345 | 3068873 | 2965407 | 2955842201301330 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-20 | 2013-04-19 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-14 | Malfunction |  | It was reported that during a da vinci si partial nephrectomy procedure, the jaw on the pk dissecting forceps instrument broke off and fell into the patient. The broken instrument piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found one broken grip jaw at the distal clevis at the grip's base. The broken grip tip was returned taped to the instrument's housing. No broken cables or additional damage was found on the distal end of the instrument; however, biodebris was present. Engineering concluded that the damage was likely due to excess force applied to the jaw or other type of mishandling. The instrument passed electrical continuity testing. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Isi has made several follow-up attempts to obtain additional information concerning the reported event; however, no additional information has been received. A follow-up medwatch report will be submitted to the fda if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3096392 | 3068897 | 2965432 | 2955842201301331 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-20 | 2013-04-19 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-09-20 | Injury | HospitalizationLife ThreateningRequired Intervention | On (b)(4) 2013, isi received a legal complaint alleging that a patient who underwent a da vinci si hysterectomy procedure on (b)(6) 2012, at (b)(6), experienced post-operative complications of bilateral renal obstruction and acute renal failure. | A review of the patient's medical records indicated that the patient underwent a robotic total laparoscopic hysterectomy procedure on (b)(6) 2012. The patient's pre-operative diagnosis was pelvic pain and menometrorrhagia. The patient's operative report indicated that her urine was clear at the end of the procedure and she tolerated the procedure well and went to recovery in satisfactory condition. There was no indication in the operative report that a malfunction of the da vinci si system, an instrument, or an accessory occurred. The patient was discharged from the hospital on (b)(6) 2012 in good condition according to the patient's discharge summary report. The patient's discharge summary report indicated that she underwent a robotic total laparoscopic hysterectomy under general anesthesia without complications. The discharge summary report also indicated the patient was ambulating well, voiding well, afebrile. The patient's medical records indicate that on (b)(6) 2012, she was admitted to the hospital's emergency room (er) with a complaint of anuria and a diagnosis of acute renal failure and bilateral renal obstruction. A ct scan performed that same day showed bilateral hydronephrosis. On (b)(6) 2012, the patient underwent placement of bilateral percutaneous nephrostomy tubes. The operative report indicated that the patient's renal failure quickly resolved and she was discharged from the hospital on (b)(6) 2012. The patient's medical records indicated that on (b)(6) 2012, the patient was readmitted to the hospital with a complaint of no drainage from her right kidney. A nephrostogram performed on the patient revealed that the nephrostomy tube was no longer in the kidney. The nephrostomy tube was replaced and she was discharged from the hospital on (b)(6) 2012. The patient's medical record indicated that she was scheduled to undergo bilateral ureteral reimplants. On (b)(4) 2013, isi contacted the corporate director of risk management for the hospital to obtain additional information regarding the reported events. The corporate director of risk management indicated that no malfunction of the da vinci si system, instruments, and accessories during the hysterectomy procedure performed on (b)(6) 2012, was reported to him by the surgical staff. He indicated that the surgical procedure was not recorded. He indicated that he was unable to give a possible cause of the patient's reported injuries and that he was unable to provide any additional clinical information regarding the reported events. Isi's review of the site's system log for (b)(6) 2012, revealed two surgical procedures that were not performed by the surgeon involved with this complaint. However, the site's system log revealed that on (b)(6) 2012, a hysterectomy procedure was performed by the surgeon involved in this complaint and there were no occurrences of any system errors that caused or contributed the alleged injuries sustained by the patient. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3096497 | 3068979 | 2965513 | 2955842201301333 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-20 | 2013-04-19 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-03-20 | No | No | Reuse | 2013
IS1200 A43P9 | 2012-08-06 | Injury | HospitalizationLife ThreateningRequired Intervention | On (b)(4) 2013, isi received a legal complaint alleging that a patient who underwent a da vinci hysterectomy procedure on (b)(6) 2012, at (b)(6), developed post-surgical symptoms of abdominal pain and distention, soft tissue swelling in the lower abdomen and labial area, anemia, nausea, and fever. The legal complaint also alleges that the patient was readmitted to the hospital post-operatively from robotic surgery for repair of a bladder laceration on (b)(6) 2012. | A review of the patient's operative report for the da vinci hysterectomy performed on (b)(6) 2012, indicated that during the surgical procedure the power was lost twice with prompt return of generator power and battery-operated power with the robot with no compromise to the surgical field or the patient. The operative report also indicated that after completion of the surgical procedure, the patient was awakened and taken to the recovery room in stable condition. A review of the patient's medical records indicated that a ct scan was performed on the patient on (b)(6) 2012, and it was discovered that the patient sustained a probable bladder laceration with intra-abdominal urine extravasation. On (b)(6) 2012, the patient was readmitted to the hospital and an exploratory laparotomy procedure performed on that same day confirmed that the patient sustained a laceration to her bladder; however, the origin was undetermined. The bladder laceration was repaired per urology. Based on the patient's discharge summary report, the patient had slow return of gi function post-operatively with improvement in pain and resolution of fever. The patient was discharged from the hospital on (b)(6) 2012, with instructions on catheter care and postoperative care. No further clinical information was provided. On (b)(4) 2013, isi contacted the clinical sales representative (csr) who was present during the da vinci hysterectomy on (b)(6) 2012. The csr indicated that he did not observe any malfunction of the da vinci system, instruments, or accessories during the procedure. The csr also indicated that he did not witness the patient's bladder being damaged at any time during the surgical procedure. On (b)(4) 2013, isi contacted the risk manager of the hospital to obtain additional information regarding the reported events. The risk manager was unwilling and unable to provide any information regarding the reported events. A review of the system error logs for the da vinci surgical system with a procedure date of (b)(6) 2012, confirmed that 20 minutes into the procedure, the surgeon console lost ac power and switched to battery backup with an alarm. Ac power returned approximately 15 seconds later. The first alarm was cleared; however approximately 5 minutes later, the ac power was lost again, but was restored 8 seconds later. The second alarm was cleared approximately 1 minute later. No others errors were noted to have occurred. | DA VINCI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3096510 | 3068994 | 2965529 | 2955842201301332 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-20 | 2013-04-19 | Yes | No | Health Professional | IS1200 A43P9 | No | No | 2013-03-20 | No | No | Reuse | 2013
IS3000 A60P8 | 2012-08-31 | Injury | Required InterventionHospitalization | On (b)(4) 2013, isi received a legal complaint alleging that a patient who underwent a da vinci si myomectomy procedure on (b)(6) 2012 at (b)(6) developed post-surgical symptoms of chest pain, dyspnea, abdominal pain, and intra-abdominal pelvic abscesses. The legal complaint alleges that the patient developed chest pain post-operatively on (b)(6) 2012. Allegedly, oxygen was administered to the patient and the patient's chest pain was relieved spontaneously after 30 minutes. The legal complaint alleges that on (b)(6) 2012 a chest ct angiogram (cta) performed on the patient showed no evidence of pulmonary embolus and that same day the patient was discharged from the hospital. On (b)(6) 2012, the patient allegedly developed shortness of breath (sob), chest pain, and abdominal pain and on (b)(6) 2012, the patient was admitted into the hospital's emergency room due to worsening symptoms of sob. Allegedly a chest x-ray performed showed that the patient had small pleural effusions with associated atelectasis. The patient was treated for a possible respiratory tract infection. On (b)(6) 2012, the patient allegedly developed a recurrent fever and on (b)(6) 2012, an abdominal ct performed on the patient showed that the patient had multiple intra-abdominal and pelvic abscesses. On (b)(6) 2012, the patient underwent a ct-guided drainage for 3 of the abscesses. Allegedly, due to the patient's recurrent fever, on (b)(6) 2012, another ct was performed on the patient. The ct scan showed that the patient had persistent fluid collections. The legal complaint alleges that on (b)(6) 2012, the patient underwent an exploratory laparotomy, small bowel resection, lysis of adhesions, and placement of jackson-pratt drains. The patient was admitted into the icu post-operatively and was then transferred to the regular floor on (b)(6) 2012. Allegedly, 14 days post operatively, the patient remained stable and afebrile. On an unspecified date/time, the patient was discharged to an acute care nursing facility on a full liquid diet and antibiotic therapy. | On (b)(4) 2013, isi contacted the risk manager at the site to obtain additional information regarding the reported events. The risk manager declined to provide any information regarding the patient or details of the reported events. However, the risk manager indicated that there were no reports by the surgical staff of a malfunction of the da vinci system, instruments, or accessories during the myomectomy procedure performed on (b)(6) 2012. A review of the patient's operative report for the da vinci si myomectomy procedure performed on (b)(6) 2012 did not indicate that the patient experienced any complications during the surgical procedure. The patient's operative report indicated that fragments of all the fibroids were removed, the pelvis was thoroughly irrigated, all debris was removed, and good hemostasis was noted. The disposition on the operative report noted that the patient was extubated and sent to the recovery room in stable condition. Three separate surgeries were performed on the system on the same day, and insufficient information is available to determine which log applies to this event, although no errors in function were noted in all logs for that day. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3096525 | 3069004 | 2965539 | 2955842201301334 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-20 | 2013-04-19 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-03-20 | No | No | Reuse | 2013
|  | Injury | Required Intervention | On march 27th, 2013, isi became aware of an on-line article that was posted on (b)(6). The article was published on march 26th, 2013, by the new scientist magazine titled robo-surgeon da vinci faces lawsuits. The article refers to a study that was published in the journal of robotic surgery, dated july 2008, volume 2, issue 2, pp 95-96. Only the abstract of the study was provided and is as follows: we report a rare case of a da vinci robotic arm failure during a laparoscopic robot-assisted radical prostatectomy. The articulation joint of an endowrist needle driver was broken and positioned at such an angle that made it impossible to remove through the trocar. In addition, it was later discovered that a small piece of the instrument was detached and remained inside the abdomen of the patient without even having been identified on subsequent radiological evaluation. In order to remove the broken instrument, we had to uninstall it from the robot arm and a bigger incision had to be made in the abdominal wall of the patient. The operation was completed without any other incidents. Testing the broken instrument for integrity is recommended to avoid this rare complication. | Based on the limited information provided regarding the reported event, isi was unable to contact the hospital where the event occurred. Therefore, the root cause of the reported failure mode cannot be determined. A follow-up mdr will be submitted if additional information is received. Inefficient information was available, so system logs could not be reviewed for this event and no instrument matching this event was returned to isi for evaluation. | ENDOWRIST NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3096679 | 3069181 | 2965715 | 2955842201301335 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-03-27 | 2013-04-19 | Yes | Yes | Health Professional |  | No | No | 2013-03-27 | No | No | Unkown | 2013
IS3000 A60P3 | 2013-03-25 | Injury | Life Threatening | It was reported that during a da vinci si pulmonary lobectomy procedure, the patient experienced uncontrolled bleeding. The surgeon made the decision to convert the surgical procedure to open surgical techniques. | On (b)(4) 2013, isi contacted the isi clinical sales representative (csr) who was present during the surgical procedure. The csr indicated approximately 5 hours into the surgical procedure, while the surgeon was transecting the patient's pulmonary vein or pulmonary artery to remove a lobe, the surgeon noted that the patient was bleeding. The csr indicated that the specific vessel that the patient was bleeding from is unknown, however, based on the location of the anatomy that the surgeon was working, he believes that the patient was bleeding from a vessel branch of the pulmonary vein of the left upper lobe. The csr indicated that he does not recall which instruments were installed at the time that the issue occurred; however, the surgeon attempted to control the patient's bleeding using ethicon clips installed on an ethicon clip applier instrument used through an assist port. The surgeon's attempt to control the patient's bleeding was unsuccessful, and the surgeon made the decision to convert the procedure to open surgical techniques to control the patient's bleeding and to complete the surgical procedure. The csr indicated that no malfunction of the da vinci surgical system, instruments and/or accessories occurred when the issue occurred. The csr indicated that approximately 1 week later, he spoke to the circulating nurse and she indicated to him that the patient recovered well post-operatively and was discharged from the hospital. On (b)(6) 2013, isi contacted the surgeon regarding the reported event. The surgeon indicated that the patient underwent a da vinci si lobectomy procedure due to t1 stage cancer. The surgeon indicated that towards the end of the surgical procedure, the patient began to experience bleeding from a branch of the pulmonary artery of the left upper lobe and that he attempted to repair the affected area robotically; however, the ventilation tube that was inserted inside of the patient's bronchus dislodged, causing the patient's lung to inflate. The surgeon indicated that the physician's assistant at the patient side cart experienced difficulty reinserting the ventilation tube and due to inflation of the patient's lung it was difficult to visualize the surgical area to allow for repair of the affected anatomy using the da vinci surgical system; thus he made the decision to convert the surgical procedure to open to allow for repair of the affected vessel and completion of the surgical procedure. The surgeon indicated that he repaired the affected area with a suture. The surgeon did not indicate the cause of the bleeding experienced by the patient; however, he indicated that no malfunction of the site's da vinci surgical system, instruments and/or accessories occurred while he was performing the surgical procedure and the bleeding experienced by the patient was unrelated to the da vinci si surgical system. The surgeon indicated that the patient's hospitalization was not prolonged and the patient was discharged from the hospital approximately 5 days post-op. The surgeon indicated that he has had two post-surgical follow up visits with the patient and the patient is recovering well. Based on the additional information provided concerning the reported event, it has been determined that the site's da vinci si surgical system did not malfunction in a way that caused or contributed to the injury sustained by the patient. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3096795 | 3069280 | 2965815 | 2955842201301336 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-03-25 | 2013-04-19 | Yes | No | Service Personnel | IS3000 A60P3 | No | No | 2013-03-25 | No | No | Reuse | 2013
| 2012-08-24 | Malfunction |  | During robotic procedure, surgeon saw spark while using a maryland robotic instrument. The maryland was immediately removed. No patient harm. Sparking was at the instrument piece but not the robotic equipment. Device #1is this a laboratory device or laboratory test?no. |  | DAVINCI MARYLAND BIPOLAR FORCEPS | SYSTEMSURGICALCOMPUTER CONTROL | 3117172 | 3088940 | 2985484 | 3088940 | 1 | NAY | User Facility |  |  | Initial | 2013-04-19 | 2013-04-19 | No | Yes | Physician |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
40017208 | 2013-02-25 | Malfunction |  | It was reported that during a da vinci prostatectomy procedure that the maryland bipolar forceps instrument could be connected, but cannot be moved to working position. With a replacement instrument the robot arm works without any problems. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported issue. The instrument was placed on a system and driven. Recognition and engagement test passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. The instrument was fully functional. An electrical continuity test passed. An additional finding not reported was scratches on the main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3092188 | 3065267 | 2961802 | 2955842201301296 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-20 | 2013-04-18 | No | Yes | Health Professional | 40017208 | Device Returned To Manufacturer | Yes | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-22 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical pk dissecting forceps instrument was noted to have a broken cable sticking out at the distal end. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was frayed at the distal idlers. The idler pulley spun freely. The frayed cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found a derailed cable. The same grip close cable that was frayed was also derailed from and exposed on the outside of the distal idler pulley. The yaw motion was non-intuitive as a result. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3093161 | 3066015 | 2962550 | 2955842201301297 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-22 | 2013-04-18 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-22 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-01-29 | Malfunction |  | Event Description: --/-- During a da vinci si procedure, the user facility reportedly identified dull scissors on the monopolar curved scissors instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering confirmed the dull scissors complaint. The returned instrument was placed on an in house da vinci system for latex cut test. The instrument's scissors did not cut cleanly through 0. 006 latex. The latex snagged at the scissor tips. The blade edges were not damaged but exhibited wear the tips, negatively affecting cut performance. There was also an indentation in the form of a line on the blade's surface. Electrical continuity testing was performed on the instrument and passed. Additional observation not initially reported by the site was a bent banana plug and damaged main tube. The banana plug was found bent at the back end of the instrument's housing. In addition, the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that damages to the scissor's blade, main tube, and banana pin may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3093213 | 3066047 | 2962582 | 2955842201301298 | 1 | NAY | Manufacturer | OtherUser facility | Other | FollowupInitial | 2013-03-21 | 2013-04-18 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42034301 | 2013-03-22 | Malfunction |  | The instrument was returned without complaint details, no further information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was received and evaluated. The instrument was returned without details, no further information was provided. Engineering found a broken cable. The one drive cable was broken within the snake wrist. The broken cable segments stuck out at the wrist. The wrist was unable to be properly straightened due to the cable breakage. No other damage was found. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3093277 | 3066085 | 2962619 | 2955842201301299 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-22 | 2013-04-18 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-03-22 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-15 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted that a wire snapped on the mega needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely, but it exhibited some damage on the edge and on the surface. The cable segment stuck out at the wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3093350 | 3066134 | 2962668 | 2955842201301301 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-18 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-21 | Malfunction |  | Event Description: --/-- Prior to starting a da vinci si surgical procedure, the user facility identified a broken cable on the mega needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- Device evaluation: on (b)(4) 2013, intuitive surgical, inc. Received the instrument involved with the complaint. Engineering confirmed one broken grip close cable near the proximal pulleys and proximal clevis cable hole. The idler pulley spun freely and did not exhibit damage. The cable segment was sticking out at the instrument's wrist. No major wear was found on the proximal clevis cable hole. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | Reuse | 3093375 | 3066165 | 2962699 | 2955842201301302 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-03-21 | 2013-04-18 | No | Yes | Health Professional | 42030901 | Yes | No | 2013-03-21 | Device Not Retu | No |  | 2013
42030901 | 2013-03-21 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken cable on the mega needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering found a broken grip close cable near the proximal pulleys and proximal clevis cable hold and was sticking out at the instrument's wrist. The idler pulley spun freely and did not exhibit damage. No major wear was found on the proximal clevis cable hold. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3093348 | 3066169 | 2962699 | 2955842201301300 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-18 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-03-19 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the customer noted a broken cable on the mega needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely, but it exhibited some damage on the edge. The cable segment stuck out at the wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3093409 | 3066213 | 2962748 | 2955842201301303 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-18 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-22 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical mega suturecut needle driver instrument was noted to have a frayed cable at the distal end. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found various scratch marks at the distal end of the main tube showing light material removal. The scratches were. 090 -. 130 in length and not aligned with the tube axis. Engineering concluded that the damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3093427 | 3066214 | 2962749 | 2955842201301304 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-22 | 2013-04-18 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-22 | Device Not Retu | No | Reuse | 2013
42020504 | 2013-03-14 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the customer stated that tips of the fenestrated bipolar forceps instrument would not open and close. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was able to open and close without any problem. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3093724 | 3066574 | 2963128 | 2955842201301305 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-18 | No | Yes | Health Professional | 42020504 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-11 | Malfunction |  | Event Description: --/-- It was reported that during cleaning and sterilization, the customer noted broken cables on the monopolar curved scissors instrument. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was found to be loose at the distal idlers. The idler pulley was able to spin freely and it was not damaged. Upon opening of the housing, engineering found a broken grip cable by the clamping pulley. The clamping pulley screws were found to be secured. The idler block showed no sign of damaged. Instrument passed electrical continuity test. Engineering also found indentations marks on the edge of the distal pulley. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3093778 | 3066646 | 2952177 | 2955842201301306 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-03-19 | 2013-04-18 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42022703 | 2013-02-28 | Malfunction |  | It was reported that during a da vinci s robotic pelvic exploration, the customer reported that the pk dissecting forceps instrument wrist was unable to articulate to the left or the right. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3093886 | 3066719 | 2963253 | 2955842201301307 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-18 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-13 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the pk dissecting forceps instrument was not recognized by the system. The instrument was not used in the case. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was successfully checked for recognition on an in-house system. The pogo pins were not stuck and are were not contaminated. Dcp verification of instrument passed. An additional observation not reported was the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3093915 | 3066727 | 2963261 | 2955842201301308 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-18 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-04 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the tips on the pk dissecting forceps instrument did not meet. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent causing side to side misalignment of grips. There was a 0. 102 offset at the tips indicating overloading at the tip. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3093979 | 3066753 | 2963288 | 2955842201301309 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-18 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42011006 | 2013-03-25 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the precise bipolar forceps instrument was noted to have frayed cables at the distal end. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument had damage on the distal end, a broken pitch cable, not frayed. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering also found several scratch marks on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded that the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3093961 | 3066770 | 2963304 | 2955842201301310 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-25 | 2013-04-18 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2013-03-25 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-19 | Malfunction |  | It was reported that during a da vinci si surgical procedure the mega suturecut needle driver instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument had a frayed grip cable. The grip cable was frayed at the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at wrist were not damaged. Engineering also found an indentation on the edge of the distal pulley. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3094055 | 3066820 | 2963355 | 2955842201301311 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-25 | 2013-04-18 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-03-25 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-13 | Malfunction |  | It was reported that during a da vinci si surgical procedure that a loose wire was on the prograsp forceps instrument there were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. No loose cables were found on the distal end. The instrument was placed on system and driven. Recognition and engagement test passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. The instrument was fully functional. An additional finding not reported was deep scratches on the main tube. The distal end of the main tube had scratch marks exhibiting light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3094357 | 3067116 | 2963651 | 2955842201301313 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-18 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-03-13 | Malfunction |  | It was reported that during a da vinci si surgical procedure the bowl grasper instrument will not close. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on a system and driven. Recognition and engagement test passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. The instrument was fully functional. A finding not reported was tear marks on the insulation coating on the distal end of the main tube. Evidence is not conclusive, but the possibility of material removal exists because the metal surface of the main tube was exposed. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3094343 | 3067123 | 2949959 | 2955842201301315 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-18 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42034403 | 2013-03-06 | Malfunction |  | It was reported that during a da vinci si surgical procedure with the curved bipolar dissector instrument, it did not move when using the console. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was placed on a system and driven. Recognition and engagement test passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. The instrument was fully functional. An electrical continuity test passed. An additional finding was various scratches on the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | CURVED BIPOLAR DISSECTOR INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3094332 | 3067133 | 2963667 | 2955842201301312 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-18 | No | Yes | Health Professional | 42034403 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-22 | Malfunction |  | Event Description: --/-- It was reported that during reprocessing the da vinci si surgical mega suturecut needle driver instrument was noted to have a frayed cable. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of frayed cable. Engineering found one pitch cable was derailed at the back of the idler pulley right above the idler block inside the housing, which has no contact with the patient. The instrument grip still pitched up and down, but movement may not be precise. No other damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument was received for evaluation on (b)(4) 2013, but has not yet been evaluated; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted when the instrument has been evaluated, or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | Reuse | 3094342 | 3067141 | 2962749 | 2955842201301314 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-03-22 | 2013-04-18 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-22 | Device Not Retu | No |  | 2013
42034301 | 2013-03-13 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical thoracic grasper instrument insulation was noted to have gouges. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the main tube insulation had scratches and gouge marks showing material removed at the distal end, approximately 0. 74 from the snake wrist. Engineering concluded the damage was likely caused by excess force contact on the main tube surface. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3094366 | 3067148 | 2963702 | 2955842201301316 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-25 | 2013-04-18 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-03-25 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-15 | Malfunction |  | During central processing, the user facility identified a derailed cable on the prograsp forceps instrument. It was reported that nothing fell in a patient and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering found both grip cables on one side of the distal clevis were derailed from the distal idler pulleys. The grip open cable was seated on the outermost pulley and the grip close cable was exposed on the outside of the pulley. The yaw motion was non-intuitive as a result of the derailed cables. Other cables at the instrument's wrist were not damaged. Additional observation not initially reported by the site was deep scratches on the main tube. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal. The scratches were 0. 090 - 0. 110 in length and not aligned with the tube axis. Engineering concluded that the damage may be due to mishandling. Electrical continuity testing was performed on the instrument and passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3090356 | 3063345 | 2959900 | 2955842201301266 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-21 | 2013-04-17 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-06 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer stated that the fenestrated bipolar forceps instrument was not firing. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument cauterized fine when placed in the in-house da vinci robot system. An observation not reported by the customer was frayed pitch cable. The instrument was found with a frayed pitch cable located at the distal clevis hub. There was no damage found at the clevis. No other cable damage was found. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3090515 | 3063576 | 2960091 | 2955842201301267 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-18 | 2013-04-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-19 | Malfunction |  | It was reported that during the da vinci s prostatectomy procedure, the da vinci robot system did not recognize the maryland bipolar instrument installed on patient side manipulator (psm) arm 2. The surgical's staff attempted to solve this problem by resetting the instrument, cleaning the contact pin and resetting the sterile adapter; however, this problem could not be solved. Next, the staff reset the instrument to another psm arm; however, the da vinci robot system did not recognize it. The staff exchanged it into a backup instrument, and the planned procedure was completed without problem. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was checked for recognition using an in-house da vinci robot system. The system successfully recognized the instrument, showing 4 uses remaining. The pogo pins did not stick and were not contaminated. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3090744 | 3063891 | 2960406 | 2955842201301268 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2013-03-18 | 2013-04-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-05 | Malfunction |  | It was reported that prior to starting a da vinci si prostatectomy procedure, the customer noted that the articulation cable of the monopolar curved scissors instrument was broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at the wrist. Other cables at wrist were not damaged. Engineering also found indentations on the edge of the distal pulley. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3090864 | 3064026 | 2960541 | 2955842201301269 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-18 | 2013-04-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-18 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility identified a broken cable on the pk dissecting forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was broken at the distal end. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. Electrical continuity testing was performed on the instrument and passed. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091412 | 3064464 | 2954833 | 2955842201301275 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-22 | 2013-04-17 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-22 | Device Not Retu | No | Reuse | 2013
42018404 | 2013-03-12 | Malfunction |  | During a da vinci si hysterectomy procedure, the end of the spatula tip reportedly broke off the permanent cautery spatula instrument and then became stuck inside the trocar. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering noted that the proximal clevis was broken from the main tube and the pitch cable was broken at the distal clevis hub. The cable segment with the cable crimp remained in the clevis hub. Engineering concluded that the damage may be likely due to mishandling (excess force applied normal to the clevis). Additional observation not initially reported by the site was physical damages to the proximal and distal clevis. Engineering also concluded that the clevis damage may be likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091413 | 3064465 | 2961025 | 2955842201301271 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-17 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-19 | Injury | Required Intervention | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, a piece of the cable on the pk dissecting forceps instrument snapped off and fell into the patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found a broken pitch cable broke off at the distal end of the instrument. The cable segment that contains the crimp was missing from the clevis. The clevis did not exhibit any damage or wear marks. The other cables of the instrument's wrist did not exhibit any damage. Engineering evaluation also found damage on the instrument's main tube. The distal end of the instrument's main tube had a scratch mark exhibiting light material removal and a rough surface finish. Engineering concluded that the damage may be due to mishandling. No further damages were found on the device. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, the initial reporter of this incident contacted isi and provided additional information regarding the reported event. The initial reporter indicated that she was not present in the surgery unit when the alleged issue occurred, she does not recall the specific date of the surgical procedure and that she thinks that the procedure was performed approximately two days prior to reporting the incident to isi. The initial reporter indicated that at the beginning of a da vinci si ob/gyn procedure, while the surgeon was using the pk dissecting forceps instrument, a fragment from the instrument broke off and fell into the patient. The fragment was retrieved laparoscopically from the patient with a grasper instrument and that the hospital did not perform an x-ray to confirm removal of the fragment. The initial reporter indicated that the surgeon was able to complete the planned surgical procedure and the patient's condition was fine during and after the surgery. No further clinical information was provided by the site. | PK DISSECTING FORCEPS INSTRUMENT | Unkown | 3091347 | 3064469 | 2960984 | 2955842201301272 | 1 | NAY | Manufacturer | User facility | Other | FollowupInitial | 2013-03-22 | 2013-04-17 | Yes | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-22 | Device Not Retu | No |  | 2013
42017704 | 2013-03-01 | Malfunction |  | It was reported that during a da vinci si gastric bypass procedure the jaws of the bowl grasper instrument would not close. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was placed on an is3000 system and driven. Recognition and engagement test passed. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. Functional performance testing did not find any issues. An additional observation not reported by the site was tube damage. Black tube insulation was damaged at the distal end. The insulation was gouged on one side and had a. 070 x. 020 piece missing roughly. 365 above the instruments snake wrist. The insulation also exhibited various other scratches and gouges at the distal end without material removal. Engineering concluded the damage was likely due to mishandling/misuse. No other damage found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3091411 | 3064496 | 2961011 | 2955842201301276 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-18 | 2013-04-17 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-19 | Malfunction |  | It was reported that during reprocessing a broken cable was found on the maryland bipolar forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The complaint of a broken cable was confirmed with the following clarification, the cable was frayed. The instruments conductor wires were intact and undamaged, but the drive cable was frayed. The pitch up cable was frayed at the distal clevis hub. The frayed strands stuck out at the instruments wrist. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091399 | 3064497 | 2961012 | 2955842201301277 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-18 | 2013-04-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-18 | Malfunction |  | During a da vinci si surgical procedure, the user facility identified a crooked bipolar connector pin on the fenestrated bipolar forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering noted that the bottom bipolar pin was bent upwards and sideways. The pin exhibited various gouge/scratch marks. A bipolar cord could not be fully installed due to the pin damage. Electrical continuity testing was performed on the instrument and passed. Engineering concluded that the damage may be due to mishandling/misuse. Additional observation not initially reported by the site was deep scratches on the main tube. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal (scratches were 0. 060 - 0. 450 in length). A couple scratches were axially aligned, but the majority were not. Engineering concluded that the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091402 | 3064499 | 2961014 | 2955842201301274 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-22 | 2013-04-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-22 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-03-19 | Malfunction |  | During a da vinci si hysterectomy procedure, the mega needle driver instrument reportedly stopped working. The user facility then noticed that the instrument's wires were broken. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering noted a broken grip cable at the distal idler and was sticking out at the instrument's wrist. The idler pulley spun freely but did not exhibit any wear marks. No other cables at the instrument's wrist were damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3091414 | 3064506 | 2961026 | 2955842201301270 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-17 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-21 | Malfunction |  | During a da vinci si hysterectomy procedure, the user facility identified a wire hanging out of the fenestrated bipolar forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering noted that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still inside the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091417 | 3064509 | 2961029 | 2955842201301273 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-22 | Malfunction |  | During a da vinci si hysterectomy procedure, the user facility identified a cable hanging out from the maryland bipolar forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that both of the pitch cables were broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091528 | 3064625 | 2961160 | 2955842201301278 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-22 | 2013-04-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2013-03-22 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-12 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the maryland bipolar forceps instrument did not grasp tissue very well. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of the grips. There was a. 048 offset at the tips, indicating overloading at the tip. The instrument passed electrical continuity testing. Engineering also found various scratch marks at the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091553 | 3064630 | 2961150 | 2955842201301280 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-18 | Malfunction |  | During a da vinci si surgical procedure, the user facility identified a ripped string on the fenestrated bipolar forceps instrument before insertion. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that the conductor wires were intact and undamaged; however, the drive cable was frayed. The pitch down cable was frayed at the distal clevis hub and the frayed strands were also sticking out at the wrist. The pitch up cable was also frayed at the clevis hub. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091537 | 3064631 | 2961014 | 2955842201301282 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-03-22 | 2013-04-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-22 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-19 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the mega suturecut needle driver instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but has slight damages on the edge and on the surface. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3091538 | 3064633 | 2961153 | 2955842201301284 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-03-06 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the tenaculum forceps instrument was non-intuitive there were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal end, causing a lack of intuitive motion. The cable segment that contains the crimp was missing from the clevis. Clevis did not exhibit any damage or wear marks. Other cables at the instruments wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3091540 | 3064634 | 2952093 | 2955842201301286 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-17 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-20 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure the monopolar curved scissors (mcs) instrument orange piece split near the tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was broken and missing a piece at the proximal clevis interface. The missing piece was not returned to isi. Engineering concluded the tube extension likely broke due to excessive side loading or other mishandling. Engineering also found deep scratch marks on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded that the damages were likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091562 | 3064639 | 2961159 | 2955842201301290 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-11-17 | Malfunction |  | It was reported that during a da vinci si surgical procedure the maryland bipolar forceps instrument energy cable was suddenly cut. The instrument continued to work but without energy. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one conductor wire was broken at the yaw pulley exit. The wire was detached from the connection at the grip. The instrument failed electrical continuity testing. The yaw pulley showed no signs of arcing. Engineering also found scratches on the surface of the distal pulley. Engineering also found various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091563 | 3064640 | 2961180 | 2955842201301291 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-20 | 2013-04-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-21 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the pk dissecting forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip cable was frayed, not broken, at the distal idler pulley. The frayed strands stuck out at the wrist. No damage was found on the distal idler pulleys. The other cables at the wrist were not damaged. The instrument passed electrical continuity testing. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091564 | 3064641 | 2961181 | 2955842201301292 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-21 | 2013-04-17 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
40017208 | 2013-03-12 | Malfunction |  | It was reported that during a da vinci s surgical procedure the maryland bipolar forceps instrument was missing the bipolar prongs. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the bipolar pins were missing from the back end. Some bipolar cords are a tight fit on the bipolar pins and may require a high removal force, potentially causing the pins to pull out. The instrument failed electrical continuity testing. Engineering also found the bipolar insert was missing from the back end. This might have happened with the missing pins when removed by high force, potentially causing the insert and the pins to pull out. Engineering also found the various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091552 | 3064648 | 2961163 | 2955842201301279 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-17 | No | Yes | Health Professional | 40017208 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-18 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the large suturecut needle driver instrument cables frayed. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3091554 | 3064650 | 2961165 | 2955842201301281 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-17 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-19 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the pk dissecting forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that both conductor wires were broken at the yaw pulley exit. The wires were detached from its connection at the grip. The instrument failed electrical continuity testing. The yaw pulley showed no signs of arcing but one was broken at the distal clevis. The yaw pulley was missing a small piece and was dislodged from the distal clevis. Engineering concluded it was not clear how the piece of yaw pulley broke. Engineering also found deep scratches on the main tube at the distal end of the main tube and had a scratch mark exhibiting light material removal and a rough surface finish. Engineering concluded that the damages were likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091555 | 3064651 | 2960984 | 2955842201301283 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-17 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-19 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical large needle driver instrument was noted to have a frayed wire at the distal end. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was frayed at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 04 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3091539 | 3064652 | 2961167 | 2955842201301285 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-19 | 2013-04-17 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No | Reuse | 2013
42000207 | 2012-11-12 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical 8mm cannula instrument was noted to have a crack along the distal end. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the distal end of the cannula tube was cracked on one side. The crack started at the rim, and was roughly 1. 25 long, oriented in the axial direction. The cannula gage pin did not fit through the distal end opening; the crack had caused deformation of the cannula wall. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CANNULA INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3091541 | 3064654 | 2961169 | 2955842201301287 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-20 | 2013-04-17 | No | Yes | Health Professional | 42000207 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-05 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the monopolar curved scissors (mcs) instrument orange sleeve was cracked as the staff was trying to remove it from the cannula. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was cracked at the proximal clevis interface. The clevis was dislodged from tube extension. The tube extension was likely cracked due to excessive side loading. The instrument passed electrical continuity testing. Engineering also found various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091559 | 3064657 | 2961172 | 2955842201301288 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-19 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical maryland bipolar forceps instrument was noted to have frayed wires. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was frayed at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. Engineering also found a scratch mark on the distal end of the main tube showing light material removal. The scratch was short in length and not axially aligned with the tube. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091545 | 3064659 | 2961174 | 2955842201301289 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-08 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the pk dissecting forceps instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was checked for recognition on an in-house system. The system successfully recognized the instrument. The pogo pins were not stuck and or contaminated. Engineering also found deep scratches on the distal end and the proximal end of the main tube exhibiting light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091593 | 3064670 | 2961205 | 2955842201301293 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-17 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-06 | Malfunction |  | It was reported that during a da vinci s hysterectomy with bilateral salpingo oophorectomy procedure the large needle driver instrument was noted to have broken cables near the tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3091600 | 3064678 | 2961213 | 2955842201301294 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-03-21 | 2013-04-17 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-19 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure that the instrument was noted that a wire at the tip of the instrument at the pulley system was broken. The surgeon did not note any unusual movements or issues handling the instrument. No patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3091787 | 3064897 | 2961432 | 2955842201301295 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-03-19 | 2013-04-17 | No | Yes | Health Professional | 42022703 | No | No | 2013-03-19 | No | No | Unkown | 2013
| 2013-04-08 | Injury | Required Intervention | During a laparoscopic myomectomy, the robotic harmonic scalpel metal tip broke. X-ray was taken and revealed no retained foreign body. |  | DAVINCI | HARMONIC ACE CURVED SHEARS | 3104784 | 3075011 | 3035198 | 3075011 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-04-11 | 2013-04-17 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2013
42000606 | 2013-02-28 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the large needle driver instrument jaws would not close. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering was not able to confirm the customer reported complaint that the jaws would not close. Engineering found a frayed grip cable. Manually, we are able to close and open the jaws. The grip cable was frayed at the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at the wrist are not damaged. Engineering also found there was an indentation on the edge of the distal pulley. There were scratch marks on the distal end of the main tube showing light material removal. The scratch was short in length and not axially aligned with the tube. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3087232 | 3059176 | 2930598 | 2955842201301250 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-18 | 2013-04-16 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42004906 | 2013-01-29 | Malfunction |  | It was reported that during a da vinci si surgical procedure the cadiere forceps instrument malfunctioned. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis does not exhibit excessive damage. Engineering also found the main tube with deep scratches at the distal end. The scratches were perpendicular to the main tube orientation. Engineering concluded that the damage was likely due to mishandling and excess force contact with the main tube. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3087235 | 3059179 | 2955693 | 2955842201301251 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-18 | 2013-04-16 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-12 | Malfunction |  | It was reported that during a da vinci si sacrocolpexy procedure the maryland bipolar forceps instrument pulley had sprung during dissection. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was loose at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering found a broken pitch cable by the clamping pulley inside the housing, which has no contact with the patient. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. There were various scratch marks on the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3087236 | 3059180 | 2955694 | 2955842201301252 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-18 | 2013-04-16 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-20 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken cable on the mega needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering found a grip close cable broken at the distal idlers pulley and was sticking out at the instrument's wrist. The idler pulley spun freely but had slight damages on the edge and on the surface. The other cables at the instrument's wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 3088018 | 3060965 | 2957479 | 2955842201301253 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-12 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility reportedly identified a broken wire on the prograsp forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still there but was sticking out of the distal clevis. The clevis did not exhibit any damage or wear marks. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. Engineering concluded that the damage may be due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3088057 | 3060991 | 2957505 | 2955842201301255 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-16 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-12 | Malfunction |  | During a da vinci s hysterectomy procedure, a wire on the fenestrated bipolar forceps instrument reportedly broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Additional observation not initially reported by the site was deep scratches on the main tube. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. Engineering concluded that the damage may be due to mishandling. Electrical continuity testing was performed on the instrument and passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3088131 | 3061059 | 2957573 | 2955842201301256 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-20 | 2013-04-16 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-20 | Malfunction |  | During a da vinci si hysterectomy procedure, the or staff reportedly found a wire disconnected from the mega suturecut needle driver instrument. The or staff also indicated that the broken portion of the instrument was retained within the instument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found one grip close cable was broken at the distal idlers and was sticking out at the instrument's wrist. The idler pulley spun freely and was not damaged. Other cables at the instrument's wrist were not damaged. Additional observation not initially reported by the site were scratches on the surface of the distal pulley. Engineering was unable to conclude how the damage occurred. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3088748 | 3061815 | 2958335 | 2955842201301257 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-20 | 2013-04-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
42001105 | 2013-03-21 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a cannula that was reportedly out of round and was scratching insulation from the instruments. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the returned cannula was out of round. The 0. 226 diameter gage pin would not fit through the distal end opening, suggesting an out of round condition. A small burr was found on the inner surface of the cannula tube near the distal tip. Engineering concluded that the damage may be due to mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | 5MM CANNULA | ENDOSCOPIC ACCESSORY | 3088898 | 3061928 | 2958448 | 2955842201301258 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-16 | No | Yes | Health Professional | 42001105 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42001105 | 2013-03-21 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a cannula that was reportedly out of round and was scratching insulation from the instruments. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the returned cannula was out of round. The 0. 226 diameter gage pin would fit through the distal end opening; however, the cannula's tube was found bent. Rotation of the cannula while laying on its side showed a slight bend near the midpoint of the tube. Visual inspection of cannula showed no damage at the distal end opening. Engineering concluded that the damage may be due to mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | 5MM CANNULA ACCESSORY | ENDOSCOPIC ACCESSORY | 3088938 | 3061948 | 2958448 | 2955842201301259 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-16 | No | Yes | Health Professional | 42001105 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-03-21 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified the insulation from the thoracic grasper stripped off as a result of an out of round cannula. The reporter indicated nothing fell in a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the black tube insulation was damaged at the distal end. The insulation was gouged on one side and had various sections lifted up. A 0. 025 x 0. 025 piece was missing roughly 0. 230 above the instrument's snake wrist. The site also returned a couple 5mm cannulas exhibiting damage which may have contributed to tube insulation damage (refer to mfr report 's 2955842-2013-01258 and 2955842-2013-01259. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3088943 | 3061953 | 2958534 | 2955842201301260 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-16 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-03-21 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified the insulation from the thoracic grasper stripped off as a result of an out of round cannula. The reporter indicated nothing fell in a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the black tube insulation was damaged at the distal end. The insulation was gouged on one side and had various sections lifted up. A 0. 125 x 0. 04 piece was missing directly above the instrument's snake wrist. Site also returned a couple 5mm cannulas exhibiting damage which may have contributed to tube insulation damage (refer to mfr report 's 2955842-2013-01258 and 2955842-2013-01259. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3089005 | 3062017 | 2958534 | 2955842201301262 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-16 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-03-21 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified the insulation from the thoracic grasper instrument stripped off as a result of an out of round cannula. The reporter indicated nothing fell in a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the black tube insulation was damaged at the distal end. The insulation was gouged on one side and has a 0. 275 x 0. 185 piece was missing roughly 0. 32 above the instrument's snake wrist. Site also returned a couple 5mm cannulas exhibiting damage which may have contributed to tube insulation damage (refer to mfr reports 2955842-2013-01258 and 2955842-2013-01259). No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3089033 | 3062021 | 2958534 | 2955842201301263 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-16 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-03-21 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified the insulation from the thoracic grasper instrument stripped off as a result of an out of round cannula. The reporter indicated nothing fell in a patient. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the black tube insulation was damaged at the distal end. The insulation was gouged on one side and has a 0. 115 x 0. 060 piece was missing roughly 0. 650 above the instrument's snake wrist. Site also returned a couple 5mm cannulas exhibiting damage which may have contributed to tube insulation damage (refer to mfr report 's 2955842-2013-01258 and 2955842-2013-01259. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3089038 | 3062039 | 2958534 | 2955842201301264 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-16 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-06 | Malfunction |  | During a da vinci s surgical procedure, the user facility reportedly felt that the maryland bipolar forceps instrument was not very weakly and found that the wire was loose. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found a frayed pitch cable at the distal clevis hub. The cable segment containing the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. No other damage was found at the clevis or with any other cables. Electrical continuity testing was performed on the instrument and passed. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3089110 | 3062062 | 2958579 | 2955842201301265 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-21 | 2013-04-16 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the monopolar curved scissors (mcs) instrument was noted to have a broken cable near the tip of the instrument. The tip cover was not damaged. The instrument was removed and replaced. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The instrument passed electrical continuity testing. No other damage was found. Engineering also found print removed from extension tube. Engineering concluded the damage was likely due to improper cleaning. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's pad printing, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3089042 | 3062074 | 2958594 | 2955842201301261 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-03-18 | 2013-04-16 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2012-03-14 | Injury | Required InterventionHospitalization | On (b)(6) 2013, isi received a legal complaint alleging that a patient whom underwent a da vinci si gynecological procedure on (b)(6) 2012 at (b)(6) medical center in (b)(6), experienced pain, suffering and bleeding post the surgical procedure. The legal complaint alleges that upon the patient's follow up with her treating physician, the physician indicated to her that the post-surgical complications she was experiencing were a normal part of the recovery from her surgery. The legal complaint alleged that on (b)(6) 2012 the patient met with her treating physician due to extensive gynecological bleeding. It was alleged that the treating physician directed the patient to the emergency room at (b)(6) medical center. It was discovered that the patient required an emergency surgical procedure to correct a surgical defect that was caused by the site's da vinci si surgical system. This was not discovered during the patient's surgical procedure performed on (b)(6) 2012. Reportedly, following the emergency surgical procedure, the patient was hospitalized and confined to her bed for many days. | Isi has made several attempts to contact the site's risk management department to obtain additional information concerning the reported event; however, no additional information has been provided. A follow-up medwatch report will be submitted to the fda if additional information is received. Isi's review of the site's system log for the reported procedure date found that there were no occurrences of any system errors that caused or contributed to the alleged injury sustained by the patient. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3089323 | 3062317 | 2958817 | 2955842201301254 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-03-18 | 2013-04-16 | Yes | No | Health Professional | IS3000 A60P8 | No | No | 2013-03-18 | No | No | Reuse | 2013
| 2010-04-15 |  | Life Threatening | After a biopsy, pt was told she had stage 1 uterine cancer. She was advised to undergo a hysterectomy preferably through a minimally invasive method called the da vinci robotic arm system invented by intuitive surgical. On (b)(6) 2010, pt came for the procedure to be done. She was asked to sign some paperwork, which she did. Yet pt said, no one ever told her the procedure she was about to undergo was risky. She expressed her concerns to her doctor (who was indeed the head doctor) about the effects the procedure may have on her with respect to her age, and her other health issues; she was reassured by the doctor that all would be well. The procedure began with a pelvic wash. Rare cancer cells were found in the wash but somehow because the robotic arm was wrongly positioned, it spread the cancer cells to her fallopian tubes and abdominal area. Four day post - surgery pt was told that the cancer cells were now all over her abdominal area, and that she was now a stage 4 uterine cancer pt. In addition she was advised to immediately start undergoing chemotherapy lest she'll die in maximum 2 years. Obviously she was furious and told the doctor she'll sue. The doctor, said pt would have to prove that she didn't have the cancer cells in her abdominal area prior to the surgery. It's been 3 years since the surgery. Pt refused chemotherapy but is taking a vitamin c iv therapy twice a week. Pt contacted the local media (tv and newspaper) to report her case, but was turned down. Even malpractice attorneys refused taking her case. It is worth noting that on her pathology report, the pathologist wrote that the stage iii cancer level was probably caused by the, da vinci robotic arm system. |  | DA VINCI ROBOTIC ARM | ROBOTIC SURGICAL SYSTEM | 3094776 | 3065991 | 2962525 | MW5029777 | 1 | NAY | Voluntary |  | Patient | Initial | 2013-04-16 | 2013-04-16 | Yes | No | Health Professional |  |  | No |  |  |  |  | 2013
42022703 | 2013-03-18 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a broken cable on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3085893 | 3057925 | 2954833 | 2955842201301240 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-18 | 2013-04-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-06 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted that the wire snapped during case on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely, but it exhibited some damage on the surface. The cable segment stuck out at the wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3086064 | 3058023 | 2954537 | 2955842201301241 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-18 | 2013-04-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-21 | Malfunction |  | The user facility reportedly identified cables sticking out of the large needle driver instrument's wrist during central processing. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering found a frayed pitch cable at the distal clevis. No other damage was found on the returned instrument, including the pulley and clevis. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3086177 | 3058145 | 2954659 | 2955842201301242 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-15 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-03-21 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified a broken cable on the mega needle driver instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the grip close cable was broken at the distal idlers and was sticking out at the instrument's wrist. The idler pulley pin spun freely and was not damaged. The other cables at the instrument's wrist were not damaged. Additional observation not initially reported by the site was indentation on the edge of the distal pulley with scratches on the surface. Engineering concluded that the pulley damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3086276 | 3058276 | 2954790 | 2955842201301243 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-21 | 2013-04-15 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-18 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted a broken cable on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode of broken cable. The instrument was placed on the in-house da vinci robot system and driven. Instrument passed recognition, engagement and electrical continuity tests. Instrument moved intuitively with full range of motion in all directions. The grips were able to open and close properly. The instrument was found to be fully functional. Additional observation not reported by site was char marks on one side of the yaw pulley. Internal laboratory testing has determined that the arc track failure mode is most likely to occur when a bipolar instrument is energized with no tissue between the grips. Internal arcing in bipolar instruments results in a potential loss of product function, but no transfer of energy to the patient. Intuitive has concluded that the likelihood of injury due to this failure is remote. User are instructed to retain back-up instruments in case that an instrument fails to perform, and are also instructed in proper technique to minimize the likelihood of internal arcing. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3086351 | 3058319 | 2954833 | 2955842201301244 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-18 | 2013-04-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-11 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted a broken cable on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3086733 | 3058654 | 2955168 | 2955842201301245 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-18 | 2013-04-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-18 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted a broken cable on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with frayed pitch cable at distal idler. There was no damage found on the pulley or the clevis. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3086751 | 3058680 | 2954833 | 2955842201301246 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-18 | 2013-04-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-21 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified wires that were stretched out at the jaw of the pk dissecting forceps instrument. No missing or fallen pieces were reported. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was frayed at the distal end. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing was performed on the instrument and passed. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3086903 | 3058850 | 2955364 | 2955842201301247 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-11 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the tips did not line up on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument's tip were able to meet with no problem. Instrument passed electrical continuity test. Additional observation not reported by site was a broken pitch cable. Engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3086922 | 3058855 | 2955369 | 2955842201301248 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-18 | 2013-04-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-05 | Malfunction |  | During a da vinci si hysterectomy procedure, the user facility reportedly identified frayed wires at the end of the pk dissecting forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that the conductor wires were intact and undamaged but the pitch down cable was frayed at the distal clevis hub. The frayed strands were sticking out at the instrument's wrist. Other cables at wrist were not damaged. Additional observation not initially reported by the user facility was damaged distal clevis. The clevis hub exhibited a scratch mark adjacent to the frayed pitch cable. Engineering concluded that the damage to clevis may be likely due to mishandling/misuse, which likely contributed to cable fraying. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3086926 | 3058862 | 2955376 | 2955842201301249 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-20 | 2013-04-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No | Reuse | 2013
410322 | 2013-04-15 | Malfunction |  | During robotic procedure, the vessel sealer was connected to robotic arm. While in pt, the machine / computer reported malfunction, alarmed with exposed blade reported. However, no blade noted by any staff. Vessel sealer removed from pt and removed from field. No harm to pt. |  | ENDOWRIST ONE VESSEL SEALER | VESSEL SEALER | 3097343 | 3068376 | 2964910 | MW5029828 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-04-15 | 2013-04-15 | No | Yes | Health Professional | 410322 | Yes | Yes |  |  |  |  | 2013
42020505 | 2013-03-06 | Malfunction |  | It was reported that during reprocessing the jaws were bent on the fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument jaws were bent. One grip was bent, causing a side to side misalignment of the grips. There was a. 098 offset at the tips, indicating overloading at the instruments tip. Electrical continuity test passed. An additional finding was various scratches on the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the material removal if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3082327 | 3054323 | 2950837 | 2955842201301222 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-12 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-06 | Malfunction |  | It was reported that during a da vinci si surgical procedure a loose cable was found on the pk dissecting forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the loose cable. The pitch cable was broken at the distal clevis hub. The cable segments that contain the crimps were still retained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3082328 | 3054324 | 2950838 | 2955842201301223 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-12 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-08 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the cable was hanging at the tip of the mega suturecute needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the cable was hanging at the tip. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the surface and on the edge. The cable segment stuck out at the instrument wrist. The other cables at wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3082422 | 3054418 | 2950932 | 2955842201301224 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-12 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-20 | Malfunction |  | The user site reported an unspecified reprocessing issue with the prograsp forceps instrument. The instrument reportedly was not used on a patient after the reported issue was identified. There was no allegation of patient harm or fragments falling into a patient. | The instrument was returned and evaluated. Engineering was not able to confirm any reprocessing issue with the returned instrument. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3083244 | 3055430 | 2914394 | 2955842201301225 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-14 | 2013-04-12 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-14 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-20 | Malfunction |  | The site reported an unspecified reprocessing issue with the prograsp forceps instrument. The instrument reportedly was not used on a patient after the reported issue was identified. There was no allegation of patient harm or fragments falling into a patient. | The instrument was returned and evaluated. Engineering was unable to confirm any reprocessing issue with the returned instrument. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3083279 | 3055465 | 2914394 | 2955842201301226 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-14 | 2013-04-12 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-14 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-20 | Malfunction |  | The reporter alleged an unspecified reprocessing issue with the large needle driver instrument. The instrument reportedly was not used on a patient after the reported issue was identified. There was no allegation of patient harm or fragments falling into a patient. | The instrument was returned and evaluated. Engineering was unable to confirm any reprocessing issue with the returned instrument. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3083342 | 3055516 | 2952030 | 2955842201301227 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-14 | 2013-04-12 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-14 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-03-06 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, loose wires were found on the tenaculum forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the loose wire. The instrument was found with the pitch cable broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3083368 | 3055579 | 2952093 | 2955842201301228 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-12 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-03-15 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed sparks and smoke coming from an exposed cable on the permanent cautery hook instrument. The instrument was reportedly removed safely and the planned surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering evaluation found the conductor wire of the instrument damaged and sticking out of the conductor wire cap. The wire insulation was found to be damaged. However, the instrument passed an electrical continuity test. No damage was found on the instrument's yaw pulley. On (b)(4) 2013, intuitive surgical inc. (isi) contacted the isi clinical sales representative (csr) of the site and obtained additional information regarding the reported event. The csr did not know if the instrument was inspected prior to use or if the instrument collided with any other instruments during the surgical procedure. The csr indicated that it was unknown if the instrument was removed at any time during the surgical procedure. Per the csr, the surgical procedure was not recorded. The surgical procedure was completed using a replacement permanent cautery hook instrument. The csr confirmed that there was no harm to the patient as a result of the reported event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3083874 | 3056023 | 2952537 | 2955842201301229 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-15 | 2013-04-12 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-03-18 | Malfunction |  | The tenaculum forceps instrument was returned without a specific reason or initial complaint. No patient harm, adverse outcome or injury was reported. Based on visual inspection, the pitch cable appeared to be broken at the distal clevis hub. | The instrument was returned and evaluated. Engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was missing. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3084140 | 3056222 | 2952736 | 2955842201301230 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-18 | 2013-04-12 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-08 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical fenestrated bipolar forceps instrument cable was noted to be bulging and fraying. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was found loose and frayed at the distal clevis hub. Both cable crimps remained in the clevis. Upon removing the housing, the pitch cable was found loose at the clamping pulley, and derailed at the idler block pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3084310 | 3056325 | 2952839 | 2955842201301231 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-12 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-25 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the mega suturecut needle driver instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found scratches on the surface of the distal pulley. Evidence was inconclusive how damage occurred. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3084322 | 3056330 | 2952844 | 2955842201301232 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-18 | 2013-04-12 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-14 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure the cable broke on the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the broken cable. One grip cable was broken at the distal idlers. The idler pulley spun freely but did exhibit some wear marks and an indentation. The cable segment stuck out at the instruments wrist. The other cables at the wrist were not damaged. Additionally, engineering found damage to the blades. Both blades had indentations along the edges. Engineering concluded the damage was likely due to mishandling. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported cable malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3084286 | 3056374 | 2952888 | 2955842201301233 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-12 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
40018010 |  | Injury | Required Intervention | On (b)(6) 2013, isi became aware of a video recording of a da vinci hysterectomy procedure performed at (b)(6) in 2012, titled perils of insulation failure during robotic hysterectomy. The video showed while the surgeon was performing the colpotomy portion of the surgical procedure, arcing from the wrist of the monopolar curved scissors (mcs) instrument with the mcs tip cover accessory installed occurred, resulting in a superficial burn to the patient's sigmoid colon. The planned surgical procedure was completed with a replacement tip cover accessory. After completion of the surgical procedure, the surgeon consulted with the hospital's general surgery department for assistance with repair of the patient's sigmoid colon. With the guided assistance of a general surgeon repair of the patient's sigmoid colon was performed using the da vinci surgical system. As a precaution, the surgeon reinforced the affected area by placing two imbricating sutures. Per the information provided by the narrator of the video, the patient did very well post-operatively without any complications. | The tip cover accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3084511 | 3056509 | 2953023 | 2955842201301234 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-03-13 | 2013-04-12 | Yes | Yes | Health Professional | 40018010 | No | No |  | No | Yes | Unkown | 2013
IS2000 A51P8 | 2013-03-14 | Other | Other | It was reported that during a da vinci s prostatectomy procedure, the site experienced system error message, invalid cannula detection while installing an instrument onto a patient side manipulator (psm) arm. The site contacted isi for technical support engineering (tse) assistance. The surgeon confirmed with the tse that the clamps on the cannula mount were closed and were not opening after installation of the instrument. The tse instructed the site to push the cannula towards the psm with the cannula mount clamps closed while installing the instrument; however, the issue recurred. With the assistance of the tse, the site performed additional trouble-shooting techniques; however, the issue persisted. The patient was under anesthesia and the port incisions had been created when the surgeon made the decision to abort the planned surgical procedure and reschedule to a later date. | The investigation conducted by field service engineering determined that system error message invalid cannula detected experienced by the site was due to a faulty cannula mount on the patient side manipulator (psm) arm 1. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The psm cannula mount accessory is attached to the psm and is designed to hold the cannula securely in place. The system was repaired by replacing the affected cannula mount. On (b)(4) 2013, isi contacted the initial reporter of this event. The initial reporter indicated that the patient underwent the rescheduled surgical procedure on (b)(6) 2013. The initial reporter indicated that to the best of her knowledge the rescheduled procedure went well and the patient is doing well. The initial reporter indicated that to the best of her knowledge, the patient has not returned to the hospital due to experiencing any post-surgical complications as a result of the reported event. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3084523 | 3056533 | 2953047 | 2955842201301235 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-14 | 2013-04-12 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2013-03-14 | Device Not Retu | No | Reuse | 2013
41032204 | 2013-01-18 | Injury | Required Intervention | It was reported that during a da vinci si sigmoid colectomy procedure, while the surgeon was using the endowrist one vessel sealer instrument, the error message error on instrument occurred. The site selected the menu option solve problem on the touchscreen monitor, however, the site was unable to open the grips on the instrument. Reportedly, the patient experienced bleeding. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the blade exposed; however, there was no damage to the blade. The tracks exhibited some residue. Review of the instrument's icb log and msc log confirmed that the instrument's blade jammed a couple of times and multiple attempts were made to retract the blade. On (b)(4) 2013, intuitive surgical contacted the initial reporter of this incident. The initial reporter indicated that the patient's bleeding occurred while the surgeon was sealing the patient's vessel using the vessel sealer instrument. The initial reporter indicated that removal of the instrument from the patient found that the blade between the grips of the instrument was exposed. The initial reporter indicated that there was no permanent injury to the patient, as the surgeon was able to successfully stop the patient's bleeding using the vessel sealer instrument. The initial reporter indicated that the procedure was completed with an instrument of the same type. No other information was provided. Intuitive surgical has contacted the initial reporter of this incident several times to obtain additional information based on the additional information initially provided on (b)(6) 2013. As of the date of this report, no additional information has been provided by the site. | ENDOWRIST ONE VESSEL SEALER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3084536 | 3056552 | 2953066 | 2955842201301236 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-14 | 2013-04-12 | Yes | Yes | Health Professional | 41032204 | Device Returned To Manufacturer | Yes | 2013-03-14 | Device Not Retu | No | Unkown | 2013
42019408 | 2013-03-15 | Malfunction |  | It was reported that during reprocessing a cable was broken on the mega needle driver instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the cable was broken. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the edge and surface. The broken cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3084537 | 3056555 | 2953069 | 2955842201301237 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-12 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-14 | Malfunction |  | It was reported that a large needle driver instrument had broken wires right out of the box. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The instrument had been used one time. Per the customer reported complaint, engineering confirmed the instrument had broken wires. One grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the instruments wrist. Other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3084567 | 3056564 | 2953078 | 2955842201301238 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-14 | 2013-04-12 | No | Yes | Health Professional | 42000606 | Yes | No | 2013-03-14 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-03-18 | Malfunction |  | The permanent cautery spatula instrument was returned without a specific reason or initial complaint. No patient harm, adverse outcome or injury was reported. Based on visual inspection, the instrument appeared to have a broken ceramic sleeve at the tip. | The instrument was returned and evaluated. Engineering observed that the ceramic sleeve and one side of the distal clevis were found to be broken. The cables were not damage, but other damage was found linked to the broken distal clevis. The instrument also exhibited a broken piece on the conductor cap. Engineering concluded that damage was likely due to mishandling possible caused by excessive side loading at the distal tip during use. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3084716 | 3056709 | 2953223 | 2955842201301239 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-18 | 2013-04-12 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-12 | Malfunction |  | Event Description: --/-- Prior to starting a da vinci si procedure, the user facility reportedly identified a broken blade from the monopolar cautery instrument. The instrument was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | Manufacturer Narrative: --/-- Correction: the initial mdr submitted on (b)(4) 2013 has been reclassified as a non-reportable event since the recurrence of the failure modes are not expected to be likely to cause or contribute to a death, serious injury or serious deterioration in the state of health of a patient. No adverse events have occurred related to these failure modes. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering confirmed the broken blade on the instrument. The instrument blade was found with an indentation near the blade's tip. The scissors did not fully close due to the blade damage. As a result, the blades did not cut through their entire length and the cut test failed. One blade's edge had an indentation between the midpoint and tip that acted as a mechanical stop for the other blade when closing. Engineering concluded that damages may be due to mishandling. Electrical continuity testing was performed on the instrument and passed. An additional finding not reported by the site was corrosion found on the dark tube reinforcement ring. There were several spots on the ring, which revealed the metal surface located right above the tube extension. No other damage was found with the instrument. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3080538 | 3052820 | 2949334 | 2955842201301193 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-05-01 | 2013-04-11 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-05-01 | Device Not Retu | No |  | 2013
42011006 | 2013-03-14 | Malfunction |  | During a da vinci si surgical procedure, the user facility identified a frayed cable on the precise bipolar forceps instrument. No missing or fallen pieces were reported. There was no allegation of patient harm, adverse outcome or injury. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was broken at the distal end. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. Electrical continuity testing was performed on the instrument and passed. Additional observation not initially reported by the site was bent bipolar pins. One of the two pins at the back of the housing was bent. Engineering concluded that the bent pins may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3080528 | 3052833 | 2949347 | 2955842201301192 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-14 | 2013-04-11 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2013-03-14 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-03-12 | Malfunction |  | Prior to starting a da vinci si procedure, the user facility reportedly identified damaged insulation on a monopolar cautery instrument. The instrument was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed the damaged insulation on the instrument. Engineering found that the distal end of the instrument's main tube exhibited various scratch marks and gouge marks with light material removal. Engineering concluded that the damage was most likely caused by mishandling (excess force contact on the main tube surface). The endowrist instruments instructions for use (ifu) specifically states:general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3080558 | 3052845 | 2949359 | 2955842201301194 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-12 | 2013-04-11 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-03-12 | Malfunction |  | Prior to starting a da vinci si procedure, the user facility reportedly identified a broken wire on the tenaculums forceps instrument's wrist. The instrument was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that one of the pitch cables was broken at the distal clevis hub. The cable segment containing the crimp still remained in the clevis. The clevis did not exhibit any excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3080571 | 3052870 | 2949398 | 2955842201301195 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-12 | 2013-04-11 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-03-12 | Malfunction |  | Prior to starting a da vinci si procedure, the user facility reportedly identified a broken wire on the tenaculums forceps instrument's wrist. The instrument was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the returned instrument had a frayed pitch cable at the distal end: the frayed strands were approximately 0. 0985 in length. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damaged cables were found. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited a couple scratch marks with light material removal and had a rough surface finish. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3080591 | 3052884 | 2949398 | 2955842201301196 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-12 | 2013-04-11 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42018405 | 2012-12-13 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the cautery seal was broken on the permanent cautery spatula instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a piece of the ceramic sleeve was broken off. The broken piece was not returned. Engineering concluded the damage was likely due to mishandling. The instrument tip looked like it had some build up residue leaving it discolored. An electrical continuity test passed. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3080638 | 3052943 | 2949457 | 2955842201301197 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-11 | No | Yes | Health Professional | 42018405 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42805206 | 2013-03-15 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility was unable to screw the spatula onto the monopolar cautery instrument. The instrument reportedly was not used on a patient after the reported issue was identified. There was no allegation of patient harm or fragments falling into a patient. | The instrument was returned and evaluated. Engineering confirmed the alleged complaint of not being able to screw on the spatula onto the instrument. A visual inspection of the monopolar adapter showed both vertical legs on the electrical contact were bent downwards. As a result, a disposable cautery tip cannot be fully installed. Electrical continuity testing was performed and passed. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3080659 | 3052974 | 2949488 | 2955842201301198 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-11 | No | Yes | Health Professional | 42805206 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
40017208 | 2013-02-27 | Malfunction |  | It was reported that after a da vinci surgical procedure the cable became loose on the maryland bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the loose cable. The bipolar insert was missing from the back end. The pins and conductor wires were sticking out of the chassis. The insert may have not been fully seated in the chassis or may been pulled out when bipolar cord was removed. Some bipolar cords are a tight fit on the bipolar pins and may require a high removal force, potentially causing the insert to pull out. Electrical continuity test passed. An additional finding was various scratches on the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3080673 | 3052975 | 2949489 | 2955842201301199 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-13 | 2013-04-11 | No | Yes | Health Professional | 40017208 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-13 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure a broken wire was found on the pk dissecting forceps instrument. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a wire was broken. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3080737 | 3052997 | 2949511 | 2955842201301201 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-11 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-13 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure broken wires were found on the large suturecut needle driver instrument. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed wires were broken. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the surface. The cable segment stuck out at the instrument wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3080713 | 3053007 | 2949521 | 2955842201301200 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-11 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-13 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure the instrument was noted to have a broken tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument blade was damaged at the tip. The one instrument blade had an indentation at the tip that acted as a mechanical stop for the other blade when closing. The evidence was inconclusive, but the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3081108 | 3053280 | 2949794 | 2955842201301202 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2013-05-02 | 2013-04-11 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-05-02 | Device Not Retu | No |  | 2013
42018306 | 2013-02-11 | Malfunction |  | It was reported that during reprocessing the da vinci s surgical permanent cautery hook instrument was noted to be broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint, the hook was still attached to the base on inspection. Engineering found a broken distal clevis. One of the clevis pieces broke off and was returned with the instrument. The broken piece measured roughly about 0. 281 x 0. 186. The instrument passed electrical continuity testing. Engineering also found various scratch marks on the distal end the main tube showing light material removal. The scratches were short in length and not axially aligned with the main tube. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3081201 | 3053339 | 2949853 | 2955842201301204 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-13 | 2013-04-11 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-15 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical fenestrated bipolar forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp remained in the clevis. The clevis did not exhibit excessive damage. Engineering also found the one grip was bent, causing side-to-side misalignment of the grips. There was a. 04 offset at the tips. The bent grip had a separation of the yaw pulley and the pulley cover at the glue joint, indicating overloading at the tip. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3081258 | 3053374 | 2949888 | 2955842201301203 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-11 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-05 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the prograsp forceps instrument wire was noted to be off track causing the forceps not be aligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip cable was derailed off the pulley at the distal clevis. The grips still opened and closed, but no yaw movement possible. The cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. Engineering found multiple deep scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length, suggesting instrument collisions or other type of mishandling might have caused them. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3081305 | 3053433 | 2949947 | 2955842201301205 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-11 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-05 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the pk dissecting forceps instrument wire was noted to be off track. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint, wire off track. Engineering was unable to duplicate the customer reported complaint. Engineering found the instrument had a frayed pitch cable at the distal end. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage nor wear marks. No other cables were damaged. Engineering also found bent grips. The one grip was bent, causing side-to-side misalignment of the grips. There was a. 0540 offset at the tips. Engineering concluded the damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's pitch cable, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3081331 | 3053438 | 2928158 | 2955842201301206 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-11 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-03-13 | Malfunction |  | It was reported that during a da vinci si surgical procedure the bowel grasper instrument coating was noted to be coming off the main tube. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the main tube insulation appeared to have scratches and gouge marks, with materials removed at the distal end, approximately 4. 46 from the snake wrist. Engineering concluded the damage was likely caused by excess force contact on the main tube surface, mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER 5MM INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3081322 | 3053445 | 2949959 | 2955842201301207 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-13 | 2013-04-11 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-01 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si total benign hysterectomy procedure the large suturecut needle driver instrument was noted to have broken wires. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | Reuse | 3081374 | 3053455 | 2949969 | 2955842201301209 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-04-15 | 2013-04-11 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-04-15 | Device Not Retu | No |  | 2013
42020706 | 2013-01-18 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the tenaculum forceps instrument was noted to have a separated wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3081383 | 3053462 | 2949976 | 2955842201301210 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-11 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-14 | Malfunction |  | It was reported that during reprocessing the da vinci s surgical prograsp forceps instrument was noted to have broken cables. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument had a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3081372 | 3053473 | 2949987 | 2955842201301208 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-14 | 2013-04-11 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-14 | Device Not Retu | No | Reuse | 2013
42011006 | 2013-03-13 | Malfunction |  | It was reported that during a da vinci si surgical procedure the cable frayed on the precise bipolar forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the frayed cable. The instrument was found with a frayed grip cable at the distal clevis and proximal clevis. A small indentation was found along the edge of distal idler pulley, which is the likely the cause of frayed cable. Frayed strands stuck out at the instruments wrist. The other cables at the wrist were not damaged. Additional damage found was a bent bipolar pin and char marks on the yaw pulley. One bipolar pin was bent at the back of the housing. Engineering concluded the damage was likely due to mishandling. Both yaw pulleys exhibited charring and localized melting at the grip base. Electrical continuity test passed. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported cable malfunction if to recur could cause or contribute to an adverse event. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3081421 | 3053487 | 2950001 | 2955842201301211 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-13 | 2013-04-11 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42017206 | 2013-02-06 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the maryland bipolar forceps instrument was not recognized by the system. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was successfully checked for recognition on an in-house system. The pogo pins did not stick and were not contaminated. Dcp verification of the instrument passed. The instrument was driven and moved intuitively and the grips opened and closed. Additionally, engineering found various scratch marks on the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3081781 | 3053777 | 2950291 | 2955842201301212 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-11 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-08 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the large suturecut needle driver instrument was noted to have a separated wire hanging out of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damages on the edge and on the surface. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3081880 | 3053858 | 2950372 | 2955842201301213 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-11 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
40018012 | 2013-03-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff observed arcing from the side of the monopolar curved scissors (mcs) tip cover accessory installed on the mcs instrument. The surgical staff also indicated that the tip of the mcs tip cover accessory tore. The planned surgical procedure was completed with a replacement mcs tip cover accessory and no patient harm, adverse outcome or injury was reported. There were no missing or fallen pieces reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering evaluation found a hairline/mechanical tear at the distal end of the tip cover accessory. The tear was approximately 0. 16 in length and a hole was found near the tear hole indicating that an arcing event occurred. The instruments and accessories user manual specifically states: general precautions and warnings - inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. Do not use another instrument to clean the tip cover accessory. To prevent damage to the tip cover accessory insulation, always straighten the instrument wrist under endoscopic visualization before removing the instrument through the cannula - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, intuitive surgical inc. (isi) contacted the initial reporter of this complaint to obtain additional information regarding the reported event. The reporter indicated that the mcs instrument, mcs tip cover accessory, and cannula were inspected prior to use. The mcs tip cover accessory was installed using the installation tool and electrolube was applied to the mcs instrument. The initial reporter was unable to provide additional information regarding the reported event. On (b)(4) 2013, isi contacted the clinical sales representative (csr) for the site and obtained additional information regarding the reported event. The csr indicated that the surgical staff noticed the arcing event during the surgical procedure and alerted the surgeon. The csr indicated that the mcs instrument was removed and the mcs tip cover accessory was replaced. The csr denied that the procedure was recorded. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3081927 | 3053918 | 2950432 | 2955842201301215 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-12 | 2013-04-11 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | Yes | Initial | 2013
40027202 | 2013-03-12 | Injury | Required Intervention | Event Description: --/-- It was reported that during a da vinci si myomectomy procedure, the surgical staff noticed that one of the jaws of the harmonic ace curved shears insert accessory broke off and fell into the patient. The broken fragment was retrieved and the surgical procedure was completed. There was no allegation of any harm, injury, or adverse outcome to a patient. | Manufacturer Narrative: --/-- The instrument accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument accessory is returned (post engineering evaluation) or if additional information is received. On (b)(6) 2013, intuitive surgical inc. (isi) contacted the initial reporter of this complaint. The reporter indicated that the harmonic ace curved shears instrument and harmonic ace curved shears insert accessory were inspected prior to starting the surgical procedure. The reporter indicated that 45-60 minutes into the procedure, the surgical staff saw the fragment of the instrument break off and fall into the patient. The reporter also indicated that a bedside assistant immediately removed the fragment using a laparoscopic assist port. The reporter denied that the instrument collided with any other instruments during the procedure. She also denied that the instrument was removed at any time during the procedure. The reporter stated that the patient did not develop any post-operative complications due to the reported issue. She also denied that an x-ray was used to locate the fragment. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation confirmed that one of the instrument's jaws broke off. The instrument was returned with a. 660 long segment of the curved blade detached from the distal end. The curved blade was found to be fractured within the shaft, adjacent to the clamp arm pin. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | Unkown | 3081953 | 3053946 | 2950460 | 2955842201301216 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2013-05-06 | 2013-04-11 | Yes | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2013-05-06 | Device Not Retu | Yes |  | 2013
42017910 | 2013-03-01 | Malfunction |  | It was reported that during a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument was no cutting. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument was placed on an-house system for a latex cut test. The scissors did not cut cleanly through (b)(4) latex. The latex snagged at the scissor tips. The blade edges were not damaged, but edges exhibited wear at the tips and on the blade edge, negatively affecting the cut performance. Engineering also found a bent banana plug and deep scratches on the main tube. The banana plug was found bent at the back end of the housing. Engineering concluded that damage was likely due to mishandling. Engineering found a couple of scratch marks on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3081960 | 3053956 | 2950470 | 2955842201301214 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-11 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the mega suturecut needle driver instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint; the instrument was checked for recognition on an in-house system and was successfully recognized. The pogo pins were not stuck and or contaminated. Dcp (dallas chip programmer) verification of the instrument passed. The instrument was driven and moved intuitively and the grips opened and closed. Engineering found various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the main tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3081988 | 3053976 | 2950490 | 2955842201301218 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-15 | 2013-04-11 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-13 | Malfunction |  | It was reported that during reprocessing that the plastic by the tip was found cracked on the monopolar curved scissors. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported cracked tip; however for clarification, the nonconformance was a cracked main tube at the distal end. The main tube exhibited a crack from the retaining ring of the shaft propagating downward linearly. The crack was approximately 0. 565. Engineering concluded the main tube likely cracked due to excessive side loading or other type of mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3081986 | 3053993 | 2950507 | 2955842201301217 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-14 | 2013-04-11 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-14 | Device Not Retu | No | Reuse | 2013
| 2013-03-13 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the pk dissecting forceps instrument was not functioning properly. The surgeon felt one of the cables was broken causing the jaws to open to wide. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3082231 | 3054202 | 2950716 | 2955842201301220 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2013-03-13 | 2013-04-11 | No | Yes | Health Professional |  | No | Yes | 2013-03-13 | No | No | Reuse | 2013
42030901 | 2013-03-13 | Malfunction |  | It was reported that during a da vinci si surgical procedure a cable snapped on the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3082211 | 3054206 | 2950720 | 2955842201301219 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-14 | 2013-04-11 | No | Yes | Health Professional | 42030901 | No | No | 2013-03-14 | No | No | Reuse | 2013
42017910 | 2013-02-20 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the tips almost fell off and the orange casing was broken and bent. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the orange casing was broken. The tube extension was broken and missing a piece at the proximal clevis interface. The broken piece measured roughly about 0. 256 x 0. 206. The clevis was dislodged from the tube extension. The tube extension likely broke due to excessive side loading. Electrical continuity test passed. An additional finding was various scratches on the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported material removal if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3082259 | 3054251 | 2950765 | 2955842201301221 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-15 | 2013-04-11 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-22 | Malfunction |  | It was reported that during a da vinci si surgical procedure the mega suturecut needle driver instrument stopped working during the case. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument stopped working. The scissors didn't fully close due to blade damage. Both blade edges had indentations between the midpoint and tip that acted as a mechanical stop for both blades when closing. Engineering concluded damage to the blades were likely due to misuse. An additional finding were various scratches on the main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. 62 - the instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3077033 | 3049763 | 2905594 | 2955842201301169 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-11 | 2013-04-10 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-28 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure when the fenestrated bipolar forceps instrument was inserted, it was noticed that the forceps were not aligned on the instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint was confirmed. One grip was bent, causing side to side misalignment of grips. There is a. 093 offset at the tips, indicating overloading at the instrument tip. An electrical continuity test passed. Engineering concluded damage may be due to mishandling. An additional finding were various scratches on the main tube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3077049 | 3049776 | 2946290 | 2955842201301170 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-08 | Malfunction |  | It was reported that prior to starting a da vinci si procedure the maryland bipolar forceps instrument is blocked and it does not move. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument. The pogo pins did not stick and were not contaminated. Dcp verification of instrument passed. Engineering found the instrument tip was bent. One grip was bent, causing side to side misalignment of grips. There was a. 110 offset at the tips, indicating overloading at the instruments tip. Electrical continuity test passed. Additionally, the instrument was found with a frayed pitch cable at distal idler. No damage was found on the pulley and clevis. The last finding were scratches on the main tube. The distal end of the main tube had various scratch marks with light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3077111 | 3049815 | 2946329 | 2955842201301171 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-05 | Malfunction |  | It was reported that during a da vinci si surgical procedure the monopolar curved scissors instrument could not move and the instrument bent when pulling out the trocar. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument couldn't move. The tube extension was broken and missing a piece at the prox clevis interface. The broken piece measured roughly about 0. 299 x 0. 119. The clevis was dislodged from the tube extension. The tube extension likely broke due to excessive side loading. An electrical continuity test passed. An additional finding were various scratches on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3077131 | 3049829 | 2946343 | 2955842201301172 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42805312 | 2013-03-05 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure the single site hem-o-lok medium large clip applier instrument would not close. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument was not closing. Engineering found the welding connection of the wire was broken off by the grip area resulting in the instrument not being able to close. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | HEMOLOK MEDIUMLARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3077348 | 3050041 | 2946555 | 2955842201301173 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42805312 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-06 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the robot would not recognize the prograsp forceps instrument. Impossible to dock the instrument as defective. Prograsp has not yet been used the 10 lives. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was checked for recognition on an in-house system. The system successfully recognized the instrument. The pogo pins did not stick and were not contaminated. Dcp verification of the instrument passed. The instrument was driven and moved intuitively and the grips opened and closed. An additional observation not reported by site were various scratches on the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3077533 | 3050175 | 2946689 | 2955842201301174 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
40027202 | 2013-03-12 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the jaws fell off the harmonic ace curved shears insert. There were no missing pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported issue. The insert was returned with the curved blade broken off at the distal end. The broken piece measured roughly. 660 x. 085. The blade fractured within the shaft, at the clamp arm pin. The insert was disassembled to find wear marks on the blade adjacent to the clamp arm pin. The wear marks indicated contact between the blade and pin, which likely contributed to fracture. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3077615 | 3050346 | 2950460 | 2955842201301175 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 40027202 | Yes | No | 2013-03-12 | Device Not Retu | Yes | Initial | 2013
42030901 | 2013-03-11 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy, the cable broke on the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the broken cable. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the surface. The broken cable segment stuck out at the instruments wrist. The other cables at wrist were not damaged. This report does not admit that the report or information submitted under. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3077671 | 3050375 | 2946889 | 2955842201301176 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42017206 | 2013-02-25 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the release levers on the maryland bipolar forceps instrument's housing were broken. The instrument was not used on a patient and the planned surgical procedure was completed. There was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the instrument's housing (2 of the 4 housing snap tabs and 3 of the adjacent housing pins) were broken. The instrument's housing could be removed as a result of the damage. The instrument was also missing one of the release levers. The instrument's housing most likely popped off during mishandling, causing release levers to fall out. Electrical continuity testing was performed on the instrument and passed. Additional findings not reported by the site were damages to the luer plate and the bipolar pin. Engineering found that the luer plate was dislodged from the chassis and sticking out past the chassis back pushed into the backend. A section of the chassis wall feature that holds the luer plate was broken off. Additional finding was a bent polar pin. The top bipolar pin was pushed down to the bottom pin. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Evidence not conclusive, but all of the damages to the instrument's housing, luer plate, bipolar pin and main tube suggests mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3077707 | 3050445 | 2946959 | 2955842201301177 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-11 | 2013-04-10 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | No | 2013-03-11 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-11 | Malfunction |  | The user facility reportedly identified a broken cable from the mega suturecut needle driver instrument after it was cleaned and sterilized. The instrument reportedly was not used on a patient after the reported issue was identified and here was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that one of the grip cables was broken at the distal idlers pulley and was sticking out at the instrument's wrist. The idler pulley spun freely but had damages on the surface and edge. No other cables at the instrument's wrist were damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3077783 | 3050516 | 2947030 | 2955842201301178 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-11 | 2013-04-10 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-11 | Device Not Retu | No | Reuse | 2013
42020405 | 2013-03-12 | Malfunction |  | Prior to starting a da vinci si procedure, the user facility reportedly identified a broken wire from the atrial retractor instrument. The instrument was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that one of the grip cables was frayed at the distal idlers pulley and was sticking out at the instrument's wrist. No other cables at the instrument's wrist were damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | ATRIAL RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3078076 | 3050744 | 2947258 | 2955842201301180 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42020405 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-07 | Malfunction |  | The user facility reported a broken wire on the prograsp forceps instrument. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the instrument had a frayed cable at the distal idler pulley. The frayed cable section was approximately 0. 08 in length. No damage was found on the pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3078132 | 3050805 | 2947319 | 2955842201301181 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-10 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-28 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure the pk dissecting forceps instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was missing from clevis. The clevis did not exhibit any damage or wear marks. Engineering also found deep scratches on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3078240 | 3050879 | 2947393 | 2955842201301183 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42019007 | 2013-02-07 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical cobra grasper instrument port fell in while being cleaned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the luer plate was dislodged from the chassis and stuck in past the chassis backend. Engineering concluded that the damage was most likely due to improper cleaning. Engineering also found various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3078252 | 3050891 | 2947405 | 2955842201301182 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42019007 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-01-31 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the grasping retractor instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument failed recognition testing on an in-house system. The si in-house system failed to recognize the instrument on multiple installation attempts. Instrument information system fund the dallas chip and loopback path was not being detected or functioning, causing the failure. Engineering also found a broken pitch cable at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's pitch cable, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3078273 | 3050911 | 2947425 | 2955842201301184 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-03 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure the orange sleeve of the monopolar curved scissors (mcs) instrument was noted to be cracked. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the tube extension was broken and missing a piece at the proximal clevis interface. The clevis was dislodged from the tube extension. Engineering concluded that the tube extension likely broke due to excessive side loading or other mishandling. Engineering also found several deep scratch marks on the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3078404 | 3051023 | 2947537 | 2955842201301186 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-14 | Malfunction |  | It was reported that during a da vinci si total gastrectomy procedure the fenestrated bipolar forceps instrument was not moving well and the surgical staff noted wires were out of place. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. Engineering also found various scratch marks on the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3078495 | 3051108 | 2947622 | 2955842201301185 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-13 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the mega suturecut needle driver instrument was noted to have a broken cable. The surgical staff replaced the instrument with a brand new one and the surgery was completed. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3078653 | 3051243 | 2947757 | 2955842201301190 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-10 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-12 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci si surgical procedure the mega suturecut needle driver instrument was noted to have a frayed cable. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | Reuse | 3078644 | 3051257 | 2947771 | 2955842201301187 | 1 | NAY | Manufacturer | User facility | Other | FollowupInitial | 2013-04-10 | 2013-04-10 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-04-10 | Device Not Retu | No |  | 2013
42805912 | 2013-03-12 | Malfunction |  | It was reported that during a da vinci si surgical procedure the plastic part that meets at the shaft of the crocodile grasper instrument tip broke into two pieces. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CROCODILE GRASPER SINGLESITEINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3078647 | 3051261 | 2947775 | 2955842201301188 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42805912 | No | No | 2013-03-12 | No | No | Unkown | 2013
42017207 | 2013-03-12 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the maryland bipolar forceps instrument was noted to have a frayed cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. Engineering also found a bent bipolar pin and deep scratches on the main tube. The bottom pin at the back of the housing was bent. Engineering concluded that damage was likely due to mishandling. The distal end of the main tube had two scratch marks exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3078659 | 3051271 | 2947785 | 2955842201301189 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-12 | 2013-04-10 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
| 2013-03-11 | Injury | OtherRequired Intervention | It was reported that during a da vinci si cardiac procedure, when the site moved the touchscreen monitor, the touchscreen monitor went dark. The site proceeded with the case using an external monitor. The surgeon converted the procedure to traditional open techniques due to an injury to the patient's heart. Reportedly, damage to the patient's heart was unrelated to the vision issue experienced by the site. | An investigation conducted by an isi field service engineer (fse) determined that the vision issue experienced by the site was associated with the pmva pca board and the digital video interface (dvi) cable. The pmva is a printed circuit assembly (pca) which receives the stereo video images from the camera controller and outputs the audio and video for the touchscreen monitor. The dvi cable connects the video output from the pmva to the touchscreen monitor. Review of the site's system log by the fse found no system errors that would have caused or contributed to the patient's injury. The system was repaired by replacing the affected pmva pca board and the dvi cable. The pmva pca board was returned and evaluated. Engineering evaluation was able to replicate the customer reported failure mode, as functional testing of the pmva pca board found that the fiber panel leaf (fpl) board on the pmva pca board was faulty. The fpl was replaced to repair the pmva pca board. 62 - on (b)(6) 2013, isi contacted the isi clinical sales representative (csr) who initially reported this event. The csr indicated that he was not present during the surgical procedure; however, he was contacted by the surgical staff concerning the vision issue that occurred during the surgical procedure. A patient injury unrelated to a malfunction of the da vinci si surgical system, instruments and/or accessories had occurred. The csr indicated that the surgical staff reported that while the surgeon was inserting a 5mm obturator through the patient's body wall, the surgeon inadvertently punctured the patient's heart, causing the surgeon to convert the planned surgical procedure to open surgical techniques to repair the injury and complete the surgical procedure. The csr indicated that based on the information provided by the surgical staff it is unclear if the da vinci surgical system was docked to the patient when the issue occurred. The obturator was not returned to isi for failure analysis evaluation. On (b)(4) 2013, isi contacted the surgeon's office and spoke to the surgeon's office assistant. The surgeon's assistant indicated that based on the patient's operative report notes, the endoscopic camera manipulator (ecm) was docked to the patient when the obturator was inserted. The surgeon's office assistant indicated that repair to the patient's heart was performed and that she was unable to provide other details concerning the reported event based on the patient's operative report and indicated that the surgeon would be able to provide the specific event details of the reported event. No other information was provided. As of the date of this report, no response from the surgeon has been received. A follow-up medwatch report will be submitted to the fda if additional information is received. Based on the information as provided, it was determined that the da vinci si surgical system did not malfunction in a way that would cause nor contribute to the patient's injury; however, it was determined that the patient injury occurred as a result of the surgeon's manipulation of the obturator into the patient's body wall. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | OBTURATOR | ENDOSCOPIC ACCESSORY | 3078844 | 3051473 | 2947987 | 2955842201301191 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-03-11 | 2013-04-10 | Yes | Yes | Health Professional |  | No | No | 2013-03-11 | No | No | Unkown | 2013
42022703 | 2013-03-11 | Malfunction |  | The user facility reportedly identified a frayed cable from the pk dissecting forceps instrument during central processing. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the returned instrument had a frayed pitch cable at the distal clevis hub. No damage was found at the clevis or with other cables. Electrical continuity testing was performed on the instrument and passed. Additional observation not initially reported by the site was pulley scratches and main tube damage. The surface of the distal pulley had scratches. In addition, the distal end of the instrument's main tube exhibited a couple scratch marks with light material removal and had a rough surface finish. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3070145 | 3043268 | 2939783 | 2955842201301158 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-11 | 2013-04-09 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-11 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-06 | Malfunction |  | During a da vinci si prostatectomy procedure, the wire on the prograsp forceps instrument reportedly broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that the returned instrument had a broken pitch cable at the distal clevis hub. The cable segment containing the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Additional observation not initially reported by the site was main tube damage. The distal end of the instrument's main tube exhibited a couple scratch marks with light material removal and had a rough surface finish. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3070234 | 3043329 | 2939843 | 2955842201301159 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-12 | 2013-04-09 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-13 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified tips on the mega needle driver instrument that were allegedly out of alignment. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the instrument's tips were out of alignment. One of the grip was bent causing side to side misalignment of the grips. There was a 0. 069 offset at the tips, indicating overloading at the tip. Electrical continuity testing was performed and passed. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3070249 | 3043367 | 2939881 | 2955842201301160 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-09 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-11 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified that the service diathermy electrode from the fenestrated bipolar forceps was allegedly coming away from the main body. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the electrode was coming away form the main body. The bipolar insert was broken off from the back end. One pin broke off with the insert and was returned with the instrument. The insert may have not been fully seated in chassis or may been pulled out when bipolar cord was removed. Some bipolar cords were a tight fit on the bipolar pins and may require a high removal force, potentially causing the insert to pull out. Electrical continuity test was performed on the instrument and failed. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3070285 | 3043412 | 2939926 | 2955842201301161 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-09 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-12 | Malfunction |  | During a da vinci si surgical procedure, the user facility identified a frayed cable on the mega suturecut needle driver instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed a frayed cable at the distal idlers pulley. The frayed cable section was approximately 0. 08 in length. There was also slight damage on the pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3071698 | 3044558 | 2941072 | 2955842201301164 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-09 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-12 | Malfunction |  | The user facility reported a torn cable on the prograsp forceps instrument. The reporter was unable to report if the reported issue was identified during a procedure or during central processing. There was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the instrument had a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INTRUMENT | ENDOSCOPIC INSTRUMENT | 3071625 | 3044566 | 2941080 | 2955842201301162 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-13 | 2013-04-09 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-13 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user facility identified broken wires on the mega needle driver instrument. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that one of the grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damages on the surface. The cable segment was found sticking out at the instrument's wrist. No other cable damages were found at the instrument's wrist. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3071671 | 3044593 | 2941107 | 2955842201301163 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-13 | 2013-04-09 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-13 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-12 | Malfunction |  | During a da vinci si hysterectomy procedure, the user facility identified broken wires at the end of the prograsp forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3071712 | 3044630 | 2941144 | 2955842201301165 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-14 | 2013-04-09 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-14 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-01 | Malfunction |  | During a da vinci si hysterectomy procedure, the user facility identified broken wires at the end of the pk dissecting forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering did not find broken wires at the end of the returned instrument; however, a frayed pitch cable with cable sticking out at the distal end of the instrument was observed. No additional damage was found around the area of the frayed cable. Other cables at the distal clevis were not damaged. Electrical continuity testing was performed on the instrument and passed. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3071784 | 3044700 | 2941214 | 2955842201301166 | 1 | NAY | Manufacturer |  | Other | Initial | 2013-03-14 | 2013-04-09 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-14 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-03-15 | Malfunction |  | During a da vinci s hysterectomy procedure, the user facility identified a broken cable on the tenaculum forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3071805 | 3044744 | 2941258 | 2955842201301167 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-15 | 2013-04-09 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | Yes | 2013-03-15 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-12-19 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was noted that the orange sleeve was cracked on the monopolar curved scissors instrument. The instrument was not used in the case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument's tube extension pad printing was found scratched and partially removed. Engineering concluded that the damage was likely due to mishandling during cleaning. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3071856 | 3044756 | 2941270 | 2955842201301168 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-04-09 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-11 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical mega suturecut needle driver instrument was noted to have a wire sticking out. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip cable was broken at the distal clevis. The broken segment stuck out approximately. 4760 at the wrist. The idler pulleys on the distal end spun freely but did exhibit some damage. One pulley had several indentations along the edges and the other had wear marks from the grip cable rubbing against it. The other cables at the wrist were not damaged. Engineering also found a damaged blade and a deep scratch mark on the main tube. The instrument blade appeared to have indentations along the edges of one blade. The distal end of the main tube had a scratch mark exhibiting light material removal and a rough surface finish. Engineering concluded that damage to the blade was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3068812 | 3041563 | 2938077 | 2955842201301153 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-11 | 2013-04-08 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-11 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have a wire visible at the end of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the wire is visible in the end of the instrument. Both of the pitch cables were broken at the distal clevis hub. The cable segments that contained the crimp were still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. Engineering also found a pogo pin that was stuck. The instrument could not be read on an in-house system or dcp (dallas chip programmer) due to one the pogo pin was stuck in the down position. The pogo pin did not pop up when manually actuated. It appeared that something was wedged between the pin and the barrel. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3068788 | 3041566 | 2938101 | 2955842201301154 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-11 | 2013-04-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-03-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-11 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have a wire visible at the end of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the wire was visible in the end of the instrument. The instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3068832 | 3041587 | 2938101 | 2955842201301155 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-11 | 2013-04-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-03-11 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-28 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si total benign hysterectomy procedure the pk dissecting forceps instrument was noted to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | Reuse | 3068850 | 3041599 | 2938113 | 2955842201301156 | 1 | NAY | Manufacturer | User facility | Other | FollowupInitial | 2013-04-12 | 2013-04-08 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-12 | Device Not Retu | No |  | 2013
42020505 | 2013-03-11 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the fenestrated bipolar forceps instrument was noted to have frayed wires near the tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the damaged instrument component was a broken pitch cable. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering also found indentations and scratches on the surface of the distal pulley. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3068853 | 3041603 | 2938117 | 2955842201301157 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-11 | 2013-04-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-03-11 | Device Not Retu | No | Reuse | 2013
42003605 | 2013-02-27 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the debakey forceps instrument housing was noted to be loose. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that one of the four housing snap tabs and two adjacent housing pins were broken. The housing could be removed as a result of the damage. Engineering concluded the instrument was likely mishandled. Engineering also found various scratch marks on the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | DEBAKEY FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3068869 | 3041625 | 2938139 | 2955842201301152 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-11 | 2013-04-08 | No | Yes | Health Professional | 42003605 | Device Returned To Manufacturer | No | 2013-03-11 | Device Not Retu | No | Reuse | 2013
| 2013-03-28 | Injury | Required Intervention | During a robotic-assisted laparoscopic multiple myomectomy, the davinci mega suturecut needle driver's wire came off inside the pt's abdomen. X-ray taken per hospital policy showed no metallic radiopaque foreign object but very tiny foreign bodies may not be radiographically visible. |  | DAVINCI | MEGA SUTURECUT NEEDLE DRIVER | 3085762 | 3054810 | 3147675 | 3054810 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-04-02 | 2013-04-06 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2013
42023004 | 2013-02-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the large clip applier instrument would not straighten out. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument would not straighten. The instrument had a dislodged proximal clevis at the main tube interface. Damage is likely due to overloading at the tip. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE CLIP APPLIER | ENDOSCOPIC INSTRUMENT | 3064159 | 3036661 | 2933175 | 2955842201301126 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-06 | 2013-04-05 | No | Yes | Health Professional | 42023004 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-05 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure a wire broke during the case. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a wire was broken. The instrument was found with a frayed grip cable and was derailed at the distal idler pulley. The frayed cable section was approximately 0. 09 in length. Slight damage was found on the surface of the pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3064161 | 3036662 | 2933176 | 2955842201301127 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-06 | 2013-04-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure broken wires were found on the mega needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed broken wires. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. The other cables at the wrist were not damaged. Scratches on the surface of the distal pulley were also found additionally, various scratches on the distal end of the main tube showing light material removal were found. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3064162 | 3036663 | 2933177 | 2955842201301129 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-06 | 2013-04-05 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-13 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, wires broke on the large suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed broken wires. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instrument wrist. The other cables at the wrist were not damaged. Additionally, a derailed grip cable at the distal idler pulley was found. The cable derailment was likely due to cable losing contact with the pulley during wrist articulation. The fleet angle between the proximal and distal pulleys may have contributed to the derailment. This the customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3064178 | 3036721 | 2933235 | 2955842201301128 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-06 | 2013-04-05 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-02-26 | Malfunction |  | Prior to starting a surgical procedure, the user facility identified a mega needle driver instrument's jaws that allegedly were not closing. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering was not able to confirm the alleged issue where the instrument's jaws were not closing. Recognition and engagement tests were performed on the instrument and passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Additional observation not initially reported by the site was a frayed cable at the distal idler pulley. The frayed cable section was approximately 0. 08 in length. There was also damages found on the edge of the pulleys, suggestive of mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3066189 | 3039108 | 2935622 | 2955842201301130 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-06 | 2013-04-05 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-05 | Malfunction |  | Event Description: --/-- During a da vinci si prostatectomy surgical procedure, the monopolar curved scissors instrument 's cable allegedly broke while the tip cover was still on it. In addition, the instrument's scissors reportedly no longer works. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering inspected the returned instrument and found one of the grip close cables broken at the distal idlers and was sticking out at the instrument's wrist. The idler pulley spun freely and was not damaged. The other cables at the instrument's wrist were not damaged. Electrical continuity test was performed on the instrument and passed. Additional observation not initially reported by the site was pulley and main tube damage. Engineering found scratches on the surface of the distal pulley. In addition, the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3066250 | 3039140 | 2935654 | 2955842201301131 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-03-06 | 2013-04-05 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42018406 | 2013-03-08 | Injury | OtherRequired Intervention | Event Description: --/-- It was observed during reprocessing of the permanent cautery spatula instrument, the tip of the instrument was noted to be missing. The site's central processing department notified the surgeon of the missing fragment. An x-ray performed on the patient found that the missing fragment was retained inside of the patient; however, the surgeon made the decision not to retrieve the broken instrument piece from the patient. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation confirmed the alleged issue that the tip of the instrument broke off. The instrument was found with the tip of the spatula broken. The broken surface was found to be not smooth. No damage was found on the ceramic sleeve or the pulley of the instrument. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, intuitive surgical inc. (isi) contacted the isi clinical sales representative (csr) and obtained additional information regarding the reported event. The csr indicated that he was not present in the operating room when the hysterectomy procedure was being performed. After the procedure was completed, central processing noticed that the tip of the permanent cautery spatula instrument was missing. In turn, central processing notified the surgeon and csr of the missing fragment. The surgeon reviewed an x-ray that was performed on the patient and confirmed that the missing fragment was still in the patient. The surgeon made the decision not to retrieve the missing fragment. The csr provided a photographic image of the affected permanent cautery spatula instrument. Based on the photographic image, engineering evaluation noted that the spatula fractured near the forming line (where the tip transitions from a rod to a flattened shape) and the approximate size of the piece that broke off is roughly. 150 x. 090. Engineering evaluation concluded that the damage was likely due to overloading of the tip. On (b)(4) 2013, isi contacted the robotics coordinator from the site. She indicated that the subject instrument involved in this complaint was inspected prior to use. The hysterectomy procedure was not recorded. The robotics coordinator denied that the permanent cautery spatula instrument collided with any other instruments during the procedure. At the time of contact with the robotics coordinator, she denied that the patient had developed any post-operative complications. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 3066281 | 3039162 | 2935676 | 2955842201301133 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2013-03-08 | 2013-04-05 | Yes | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No |  | 2013
42805206 | 2013-03-08 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user site identified a cautery hook that will not screw on the monopolar cautery instrument. Reportedly the instrument was not used on a patient and there was no allegation of harm or injury to a patient. | The instrument and cautery hook were returned and evaluated. The cautery hook installed onto an instrument successfully with no problem found. No physical damage was found. The instrument was found with damages on the tip of the cautery adapter. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3066290 | 3039189 | 2935703 | 2955842201301134 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-08 | 2013-04-05 | No | Yes | Health Professional | 42805206 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-13 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the csr manager found a broken cable on the pk dissecting forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering inspected the returned instrument and found a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp was missing one in the clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity test was performed on the instrument and passed. Additional observation not initially reported by the site was main tube damage. The distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3066285 | 3039194 | 2932530 | 2955842201301132 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-08 | 2013-04-05 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-07 | Malfunction |  | Prior to starting a da vinci si surgical procedure, the user site found frayed cables on the mega suturecut instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that one of the grip cables was broken at the distal idlers and was sticking out at the instrument's wrist. The idler pulley spun freely but had damages on the edge, likely due to excess force contact on the edge. No other cables at the instrument's wrist were damaged. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3066299 | 3039205 | 2935719 | 2955842201301136 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-07 | 2013-04-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-07 | Device Not Retu | No | Reuse | 2013
420172 | 2013-03-06 | Injury | Other | It was reported that during a da vinci s hysterectomy procedure, tissue would stick to the jaws on the maryland bipolar forceps instruments, causing a tear to the patient's tissue when the surgeon opened and pulled away the jaws on the instrument. It is the site's belief that a sufficient amount of anti-stick lubrication was used on the instruments. Reportedly, the jaws on the maryland bipolar forceps instrument were not fully opening and contributed to the tissue sticking to the jaws on the instruments. Medwatch report 2955842-2013-01135 was submitted for concerning the second maryland bipolar forceps instrument. | The instrument has not been returned for evaluation; however, a field investigation performed by an isi field service engineer (fse) was unable to replicate the issue experienced by the customer. The fse found that the site's da vinci si surgical system functioned within specifications. The fse also observed a da vinci hysterectomy procedure at the site; however, prior to the surgeon starting the procedure, the fse observed that the site's electrosurgical unit's cabling layout needed to be rerouted. The fse indicated to the customer that the layout of the cabling may introduce interference within the circuit. The fse also noted that the monopolar cable from the esu to the surgeon side cart had exposed wires and recommended that they replace the monopolar cable in addition to the bipolar cable. During the fse's observation of the surgical procedure, while the surgeon was performing dissection, during the 1st half of the procedure, the surgeon noted that the instrument seemed to be functioning fine; however, the surgeon eventually started to experiencing sticking issues. The fse found no issues with the site's da vinci s surgical system when the issue occurred. The fse checked the site's esu settings and they were set at 35 watts. The fse recommended to the site that they replace their esu as a precaution, and advised that if the esu is not operating to specification, it may provide intermittent results. On april 4, 2013, isi contacted the initial reporter of this event. The initial reporter indicated that the planned surgical procedure was a da vinci s hysterectomy with lymph node dissection procedure and that the surgeon was performing lymph node dissection using the maryland bipolar forceps instrument when the surgeon noted that there was a delay when opening the instrument's jaws. The initial reporter indicated that the patient's lymph node and surrounding tissue stuck to the maryland bipolar forceps instrument, causing a tear to the patient's lymph node and surrounding tissue. The initial reporter indicated that the maryland bipolar forceps instrument was replaced; however, after some time of use, the issue recurred while the surgeon was using the replacement instrument. The initial reporter indicated that the surgeon was able to complete the surgical task using the replacement instrument. The initial reporter indicated that no repair of the patient's lymph node and surrounding tissue were required, the planned surgical procedure was completed and the patient did not suffer any post-surgical complications as a result of the reported event. The initial reporter stated that the patient has not returned to the hospital due to experiencing any post-surgical complications as a result of the reported event and that based on the fse's recommendations, the hospital has replaced their esu and the esu cables. The initial reporter indicated that since replacement of the recommended devices, the hospital has not experienced any other issues related to the reported issue. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | MARYLAND BIPOLAR FORCEPS INSTRRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3066381 | 3039225 | 2935739 | 2955842201301138 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-06 | 2013-04-05 | Yes | Yes | Health Professional | 420172 | No | No | 2013-03-06 | No | No | Reuse | 2013
42004906 | 2013-03-08 | Malfunction |  | Intuitive surgical received a cadiere forceps instrument from the user facility without any complaint information. There was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering inspected the returned instrument and found one of the grip close cables was broken and was sticking out at the instrument's wrist. The idler pulley spun freely and exhibited pulley face and rim damage. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3066324 | 3039233 | 2935747 | 2955842201301137 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-08 | 2013-04-05 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-19 | Malfunction |  | It was reported that during a da vinci si surgical procedure the large suturecut needle driver instrument stopped working after a few uses. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument stopped working. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The broken cable segment stuck out at the instruments wrist. Other cables at wrist were not damaged. The instrument had no uses remaining. An additional finding were various scratches on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3066309 | 3039256 | 2935770 | 2955842201301139 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-07 | 2013-04-05 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-07 | Device Not Retu | No | Reuse | 2013
420172 | 2013-03-06 | Injury | Other | It was reported that during a da vinci s hysterectomy procedure, tissue would stick to the jaws on the maryland bipolar forceps instruments, causing a tear to the patient's tissue when the surgeon opened and pulled away the jaws on the instrument. It is the site's belief that a sufficient amount of anti-stick lubrication was used on the instruments. Reportedly, the jaws on the maryland bipolar forceps instrument were not fully opening and contributed to the tissue sticking to the jaws on the instruments. Medwatch report 2955842-2013-01138 was submitted concerning the first maryland bipolar forceps instrument. | The instrument has not been returned for evaluation; however, a field investigation performed by an isi field service engineer (fse) was unable to replicate the issue experienced by the customer. The fse found that the site's da vinci si surgical system functioned within specifications. The fse also observed a da vinci hysterectomy procedure at the site; however, prior to the surgeon starting the procedure, the fse observed that the site's electrosurgical unit's cabling layout needed to be rerouted. The fse indicated to the customer that the layout of the cabling may introduce interference within the circuit. The fse also noted that the monopolar cable from the esu to the surgeon side cart had exposed wires and recommended that they replace the monopolar cable in addition to the bipolar cable. During the fse's observation of the surgical procedure, while the surgeon was performing dissection, during the 1st half of the procedure, the surgeon noted that the instrument seemed to be functioning fine; however, the surgeon eventually started to experiencing sticking issues. The fse found no issues with the site's da vinci s surgical system when the issue occurred. The fse checked the site's esu settings and they were set at 35 watts. The fse recommended to the site that they replace their esu as a precaution, and advised that if the esu is not operating to specification, it may provide intermittent results. On (b)(4) 2013, isi contacted the initial reporter of this event. The initial reporter indicated that the planned surgical procedure was a da vinci s hysterectomy with lymph node dissection procedure and that the surgeon was performing lymph node dissection using the maryland bipolar forceps instrument when the surgeon noted that there was a delay when opening the instrument's jaws. The initial reporter indicated that the patient's lymph node and surrounding tissue stuck to the maryland bipolar forceps instrument, causing a tear to the patient's lymph node and surrounding tissue. The initial reporter indicated that the maryland bipolar forceps instrument was replaced; however, after some time of use, the issue recurred while the surgeon was using the replacement instrument. The initial reporter indicated that the surgeon was able to complete the surgical task using the replacement instrument. The initial reporter indicated that no repair of the patient's lymph node and surrounding tissue were required, the planned surgical procedure was completed and the patient did not suffer any post-surgical complications as a result of the reported event. The initial reporter stated that the patient has not returned to the hospital due to experiencing any post-surgical complications as a result of the reported event and that based on the fse's recommendations, the hospital has replaced their esu and the esu cables. The initial reporter indicated that since replacement of the recommended devices, the hospital has not experienced any other issues related to the reported issue. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | MARYLAND BIPOLAR FORCEPS INSTRRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3066373 | 3039268 | 2935739 | 2955842201301135 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-06 | 2013-04-05 | Yes | Yes | Health Professional | 420172 | No | No | 2013-03-06 | No | No | Reuse | 2013
42030901 | 2013-03-07 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the needle holder broke on the mega suturecut needle driver instrument after the first point was made during the vaginal suture intervention there were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the needle holder broke after the first point. One grip cable was broken at the distal idlers, which was likely the cause the distal end was not working properly. The idler pulley spun freely but exhibited some wear. The cable segment stuck out at the instrument wrist. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3066379 | 3039274 | 2935788 | 2955842201301140 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-07 | 2013-04-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-07 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-02-25 | Malfunction |  | It was reported that during reprocessing that a piece of insulation was missing from the monopolar cautery instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument was missing insulation. The main tube insulation exhibited damage along the proximal end with tube insulation removed. Engineering concluded damage may be due to mishandling. An additional observation not reported was a damaged prong at the tip of the instruments snake wrist where the cautery spatula tip was installed. One of the electrical assembly contact prongs was missing and no longer seated in the track. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3066395 | 3039285 | 2935799 | 2955842201301141 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-07 | 2013-04-05 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-02-25 | Malfunction |  | It was reported that during reprocessing that insulation was damaged on the monopolar cautery instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed insulation was damaged. The main tube insulation exhibited damage along the distal end. Insulation was found with several gouge marks within the area. The marks were short in length and not axially aligned with the tube. The main tube also exhibited a couple of different damaged sections with tube insulation removed. Engineering concluded the damage may be due to misuse. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3066369 | 3039310 | 2935799 | 2955842201301144 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-07 | 2013-04-05 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-06 | Injury | Required Intervention | It was reported that during a da vinci s hysterectomy procedure, the surgical staff noticed that the axial part for opening and closing of the prograsp forceps instrument broke off and fell into the patient. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The distal clevis pin of the instrument was found to be dislodged, allowing looseness of the grip tips. In addition, the grip cable was found to be loose and derailed at the distal pulley. Engineering evaluation also found the instrument clamping pulley with corrosion and damage. Engineering evaluation concluded that the damage was likely due to mishandling in cleaning. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. On (b)(4) 2013, intuitive surgical inc. (isi) contacted the initial reporter of this complaint and obtained additional information regarding the reported event. The initial reporter indicated that the prograsp forceps instrument was inspected prior to use. The initial reporter denied that the instrument collided with any other instruments during the surgical procedure. The instrument fragment that fell into the patient was retrieved using a forceps instrument used in conjunction with the da vinci s system. The initial reporter also indicated that the patient recovered as planned and was in good condition. The initial reporter denied that any post-op tests or secondary surgical procedures were performed as a result of the reported issue. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3066419 | 3039323 | 2935837 | 2955842201301145 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-07 | 2013-04-05 | Yes | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-07 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-02-25 | Malfunction |  | It was reported that during reprocessing that insulation was damaged on the monopolar cautery instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the insulation was damaged. The main tube insulation exhibited damage along the entire shaft. Insulation was found with several gouge and scratch marks. The marks were not axially aligned with the tube. Material was missing and appeared to be scraped off. Engineering concluded damage may be due to misuse. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CAUTERY INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 3066426 | 3039354 | 2935799 | 2955842201301143 | 1 | NAY | Manufacturer |  | Other | Initial | 2013-03-07 | 2013-04-05 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-21 | Malfunction |  | It was reported that during a da vinci si mitral valve repair procedure, as soon as the instrument was set into the instrument drape and locked on, the pa pushed the instrument into the cannula, we noticed the cable were broken. It looked fine before insertion into the instrument arm and cannula. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument cable was broken. One grip close cable was broken at the distal idler pulley. The idler pulley spun freely but had damage on the surface. The cable segment stuck out at the instrument wrist. The other cables at the wrist were not damaged. Engineering concluded the damage was most likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3066428 | 3039357 | 2935871 | 2955842201301146 | 1 | NAY | Manufacturer |  | Other | Initial | 2013-03-06 | 2013-04-05 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-03-08 | Malfunction |  | It was reported that prior to starting a da vinci si procedure a broken cable on the permanent cautery hook instrument was found. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken cable. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instrument wrist. The other cables at the wrist were not damaged. Electrical continuity test passed. An additional finding was one of the conductor caps on one side was missing. The missing part measured roughly about 0. 222 x 0. 237. Also both bosses were broken off. This was the part that held the yaw pulley in place. Engineering concluded that damage was most likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3066444 | 3039359 | 2935873 | 2955842201301147 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-08 | 2013-04-05 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-03-04 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy, the cable broke on the grasping protractor instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the broken cable. The pitch cable was broken at the distal clevis hub. Cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3066489 | 3039384 | 2935898 | 2955842201301148 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-08 | 2013-04-05 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No | Reuse | 2013
40009308 | 2013-03-07 | Malfunction |  | It was reported that during a da vinci total benign hysterectomy the prograsp forceps instrument did not open properly. There were no missing pieces or fallen pieces reported the planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument did not open properly. One grip close cable was derailed from the force amp pulley. The grips could not properly straighten and the grip/yaw motion was not smooth. The cable derailment was likely due to the cable losing contact with pulley during wrist articulation. An additional observation not reported by site was tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 060 -. 530 in length and not aligned with the tube axis. Damage may have been caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3066491 | 3039385 | 2935899 | 2955842201301149 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2013-03-18 | 2013-04-05 | No | Yes | Health Professional | 40009308 | Device Returned To Manufacturer | No | 2013-03-18 | Device Not Retu | No | Reuse | 2013
400272 | 2013-03-07 | Injury | Required Intervention | It was reported that during a da vinci si myomectomy procedure, the blade on the harmonic ace curved shears insert accessory installed on the harmonic ace curved shears instrument broke off and fell into the patient's pelvis. The accessory fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Conclusions - the harmonic ace curved shears insert accessory has not been returned for evaluation, as the site reported will not be returning the accessory. At this time, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned (post engineering evaluation) of if additional information is received. On (b)(4) 2013, isi contacted the isi clinical sales representative (csr) who reported this event. The csr indicated that the harmonic ace curved shears instrument and harmonic ace curved shears insert accessory were inspected by the site prior to use, and there were no issues noted and there were no issues noted by the site during installation of the accessory onto the instrument. The csr indicated that the surgeon was using the harmonic ace curved shears instrument with the harmonic ace curved shears insert accessory installed to remove fibroids from the patient's uterus. The csr indicated that after approximately 20 minutes of use, a piece from the insert accessory broke off and fell into the patient and that the broken fragment was removed laparoscopically through an assist port. The csr indicated that the patient's fibroids may have had some calcification and that he noted that while the surgeon was utilizing the instrument to remove the patient's fibroid, it appeared that the surgeon was applying some force to the instrument. The csr indicated that he observed that the surgeon was manipulating the instrument as if it was a wristed instrument. The csr indicated that the surgical procedure was completed with an instrument of a different type and there was no harm to the patient. The csr indicated that the surgeon has not reported to him that the patient has returned to the hospital due to any post-surgical complications as a result of the reported event. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3083052 | 3052959 | 2949473 | 2955842201301142 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-03-07 | 2013-04-05 | Yes | Yes | Health Professional | 400272 | No | No | 2013-03-07 | Device Not Retu | Yes | Initial | 2013
IS3000 A60P7 | 2013-03-02 | Injury | HospitalizationLife ThreateningRequired Intervention | It was reported that after a da vinci st benign hysterectomy procedure was performed, the patient came back to the hospital with post-operative symptoms of pressure and bloating. The surgeon opened the patient and reportedly noticed a small bowel injury. The surgeon proceeded to perform a bowel resection procedure. During the bowel resection procedure, the patient was placed on a ventilator. The patient reportedly almost passed away during the bowel resection procedure for an unknown reason. At the time the initial reporter reported the event to intuitive surgical inc. (isi), the patient was in the icu and was doing well. | Conclusions - on (b)(4) 2013, isi contacted the isi clinical sales representative (csr) assigned to the site and obtained additional information regarding the reported event. The csr indicated that the da vinci si benign hysterectomy procedure was performed on (b)(6) 2013 and lasted approximately 30 minutes. The csr indicated that the surgical procedure went smoothly without any patient complications reported. On (b)(6) 2013, the patient returned to the hospital with symptoms of pressure and bloating. A ct scan with contrast was performed and a perforation of the bowel was found. The surgeon performed as bowel resection via open surgery to repair the bowel injury. According to the csr, the patient did not do well the first night post-op and was put on a ventilator. By the second day post-op, the patient was doing well. The patient was reportedly in the icu for two weeks before being discharged from the hospital. The csr indicated that the general surgeon at the site determined that the bowel injury was due to cannula entry. There was no allegation that a malfunction of the da vinci si system, an instrument, or an accessory occurred or contributed to the patient's injury. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3083092 | 3052979 | 2949493 | 2955842201301151 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-03-06 | 2013-04-05 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2013-03-06 | No | No | Reuse | 2013
42017910 | 2013-03-08 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the surgical staff believed that they noticed a spark at the point where the monopolar curved scissors instrument connects to the instrument cord exterior to the pts body The site was reportedly able to power off the generator and the reported issue stopped The site was able to continue using the arm associated with the reported event and the surgical procedure was completed There was no allegation of any harm injury or adverse outcome to a pt or surgical staff member | The instrument has not been returned for eval The root cause of the customer reported failure mode cannot be determined If additional info is received a fu mdr will be submitted On b4 2013 the isi field service engineer fse visited the site for troubleshooting of the reported issue The fse found the pt side cart psc with all arms draped and the monopolar curved scissors instrument used in this procedure was still installed on pt side manipulator psm 1 The cable and cannula used during the procedure were not available for inspection The fse observed foreign black materialsubstance around the instruments cautery cable stud and the area of the cautery cable connection The cautery cable involved with the event was not available for inspection or troubleshooting The fse performed full system verification on the da vinci si system and found the system to be operating to original equipment mfrs specifications The fse instructed the site to return the mcs instrument to isi for eval The fse also advised the site to locate the cautery cable involved in this complaint and if not found to discard all cables in all their trays or have them tested by biomed for safety and functionality prior to their next surgical procedure Intuitive surgical has made several attempts to contact the site to obtain additional info concerning the reported event however no additional info has been provided as of the date of this report As of b4 2013 there have been no reported reoccurrences of the issue at the hosp The customer reported complaint dose not itself constitute a mdr reportable event however the reported malfunction if to recur could cause or contribute to an adverse event | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3103702 | 3055663 | 2952177 | 2955842201301150 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2013-03-08 | 2013-04-05 | No | Yes | Health Professional | 42017910 | No | No | 2013-03-08 | Device Not Retu | No | Unkown | 2013
42029602 | 2013-01-22 | Malfunction |  | During a da vinci si prostatectomy procedure, the user facility reportedly identified a broken wire on the large suturecut needle driver instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that one of the grip cables was broken at the distal idlers and was sticking out at the instrument's wrist. The idler pulley spun freely and was not damaged. No other cables at the instrument's wrist were damaged. Additional observation not initially reported by the site was indentation on the pulley and main tube damage. There were indentations on the edge of the distal pulley and scratches on the surface. The main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3063031 | 3035718 | 2932232 | 2955842201301107 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-06 | 2013-04-04 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2013-03-05 | Other | Other | It was reported that during a da vinci si surgical procedure, the site experienced system error code 23008. With the assistance of an isi technical support engineer (tse) the site performed an emergency power off (epo) of the system and restarted the system. However, the error code recurred. With the assistance of the tse, the site performed additional trouble-shooting techniques; however, the issue persisted. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The investigation conducted by field service engineering (fse) found that the system error code 23008 experienced by the customer was associated with the left master tool manipulator (mtml). The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The system was repaired by replacing the affected mtml. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 23008 appears when the da vinci si safety systems determines that the angular position of one or more robotic joints on the specified manipulator, as measured by the joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety systems put da vinci in a recoverable safe state. The mtml was returned to isi for evaluation and engineering was able to replicate the system error code 23008 experienced by the customer. Performance testing of the mtml in-house found that upon initialization of the mtml, system error code 23008 would trip due to faulty motor/encoder on axis 3. The axis 3 motor and encoder were replaced to repair the mtml. As of (b)(4) 2013 there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3063050 | 3035720 | 2932234 | 2955842201301108 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-05 | 2013-04-04 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-02-13 | Malfunction |  | It was reported that during a da vinci si partial nephrectomy procedure the tenaculum forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a frayed pitch cable. The frayed pitch cable was found at the wrist. No damage was found at the clevis. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3063306 | 3035997 | 2932509 | 2955842201301109 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-08 | 2013-04-04 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | Yes | 2013-03-08 | Device Not Retu | No | Initial | 2013
42017910 | 2013-02-25 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci si surgical procedure the monopolar curved scissors (mcs) instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to replicate the customer reported complaint, instrument not recognized. The instrument was installed on an in-house system and was successfully recognized. Engineering found that the tube extension was broken and missing a piece at the proximal clevis, which was likely retained by the tip cover. The clevis was dislodged from the tube extension. Engineering concluded that the tube extension likely broke due to excessive side loading or other type of mishandling. Engineering also found a scratch mark at the distal end of the main tube with light material removal. The scratch is short in length and not aligned with the tube axis, suggesting instrument collisions or rough handling may have caused the damage. The instrument passed electrical continuity testing. 62 - the instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3063308 | 3036000 | 2932512 | 2955842201301110 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-03-07 | 2013-04-04 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42017207 | 2013-02-15 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical maryland bipolar forceps instrument was noted to have bent wires. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the instrument had a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. There were also various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3063341 | 3036005 | 2932517 | 2955842201301112 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-07 | 2013-04-04 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2013-03-07 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-15 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical fenestrated bipolar forceps instrument was noted to have misaligned tips. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the instrument grips were not aligned. The one grip tip was bent, causing side-to-side misalignment of the grips. There was a. 034 offset at the tips. Engineering concluded that the damage was likely due to mishandling. Engineering also found a deep scratch at the distal end on the main tube with light material removal. The scratch was short in length and not aligned with the tube axis, suggesting instrument collisions or rough handling may have caused the damage. The electrical continuity testing passed. No other damage was found. 62 - the instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which found during failure analysis evaluation could likely cause or contribute to an adverse event were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3063339 | 3036008 | 2932522 | 2955842201301111 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-07 | 2013-04-04 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-03-07 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-03-06 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si total benign hysterectomy procedure, while the nurse was attempting to put the tip cover accessory on the monopolar curved scissors (mcs) instrument, the instrument was noted to have a tear. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the tube extension was cracked at the proximal clevis interface. The clevis was dislodged from the tube extension. The tube extension was likely cracked due to excessive side loading. The instrument passed electrical continuity testing. Engineering also found various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which was found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3063322 | 3036013 | 2932527 | 2955842201301113 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2013-03-08 | 2013-04-04 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-27 | Device Not Retu | No |  | 2013
42022703 | 2013-02-13 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure, the pk dissecting forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the instrument with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. The instrument passed electrical continuity testing. Engineering also found one grip was bent, causing side-to-side misalignment of grips. There was a. 041 offset at the tips, indicating overloading at the tip. Engineering concluded that damage was likely due to mishandling. There were also various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3063350 | 3036016 | 2932530 | 2955842201301115 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-08 | 2013-04-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-08 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the large needle driver instrument track wires were noted to be broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found indentations on the edge of the distal pulley. There were also various scratch marks on the distal end of main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3063356 | 3036022 | 2932536 | 2955842201301117 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-08 | 2013-04-04 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-20 | Malfunction |  | It was reported during a da vinci si prostatectomy procedure, the large needle driver instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found a damaged pulley. The distal idler pulley on which the broken cable was seated had a section that was bent inwards. Engineering concluded the pulley damage was likely due to mishandling and the damage likely contributed to the cable breakage. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3063378 | 3036023 | 2932537 | 2955842201301118 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-21 | 2013-04-04 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No | Reuse | 2013
42011704 | 2013-03-06 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the needle driver instrument was noted to have a broken tip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one of the grip tips broken at the grip base. Engineering concluded that damage was likely due to excess force applied to the grip tips, likely mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3063323 | 3036033 | 2932545 | 2955842201301114 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-08 | 2013-04-04 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-03-08 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the small grasping retractor instrument had a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering also found scratches on the surface of the distal pulley. There were also various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3063329 | 3036036 | 2932548 | 2955842201301116 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-08 | 2013-04-04 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No | Reuse | 2013
40018012 | 2013-03-07 | Malfunction |  | It was reported that during a da vinci si surgical procedure the mcs 8mm tip cover accessory installed on the monopolar curved scissors instrument was noted to be torn. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the insulated tip cover accessory was torn. Engineering concluded that the location of tears and damage was likely due to mishandling. There were two areas with damage. The one was in the middle of the tube and the orientation of the cut showed the damage occurred from the outside. The second included tearing, in addition a missing part at the tip of the tip cover accessory. The missing part measured roughly about 0. 157 x 0. 054. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3063764 | 3036314 | 2932828 | 2955842201301119 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-08 | 2013-04-04 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2013-03-08 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-03-01 | Malfunction |  | During a da vinci si cholecystectomy procedure, a wire reportedly broke on the mega needle driver instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that one of the grip cables was broken at the distal idlers pulley and was sticking out at the instrument's wrist. The idler pulley spun freely but had damages on the edge and surface. No other cables at the instrument's wrist were damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3063908 | 3036458 | 2932972 | 2955842201301121 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-06 | 2013-04-04 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2013-03-06 | Malfunction |  | It was reported that the glass part fell off the tip of the optic. There was no allegation of any harm, injury or adverse outcome to a patient. | The endoscope was returned and evaluated. Engineering evaluation found that the endoscope was received with damage to the distal hermetic seals, resulting in fluid invasion and subsequent major damage to optical components. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3063959 | 3036482 | 2932994 | 2955842201301120 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-06 | 2013-04-04 | No | Yes | Health Professional | IS2000 A51P8 | No | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42807103 | 2013-02-08 | Malfunction |  | Prior to starting a da vinci si single site surgical procedure, the user site realized that the cannula was broken. The reporter claimed that the bowl of the cannula was moving freely around the cannula's shaft. The user site converted the da vinci si surgical procedure to a laparoscopic procedure. There was no allegation of harm or injury to a patient or fragments falling into a patient. | The instrument was returned and evaluated. Engineering inspected the cannula and confirmed that the cannula's shaft can move with respect to the bowl feature. The weld joining the cannula's shaft and bowl was cracked around the circumference. Since the orientation of the cannula shaft relative to the cannula bow is critical to the function of the item, the cannula cannot be used. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | SINGLESITE 5 X 250 MM CURVEDCANNULA ARM 1 | ENDOSCOPIC ACCESSORY | 3064017 | 3036543 | 2933057 | 2955842201301122 | 1 | NAY | Manufacturer | DistributorUser facility | Other | InitialFollowup | 2013-03-06 | 2013-04-04 | No | Yes | Health Professional | 42807103 | Device Returned To Manufacturer | No | 2013-03-06 | No | No | Reuse | 2013
42022703 | 2013-02-22 | Malfunction |  | It was reported that prior a da vinci s surgical procedure a wire frayed on the pk dissecting forceps. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the reported issue. The grip cable was frayed at the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at the wrist were not damaged. Additionally, scratches on the surface of the distal pulley were found. Electrical continuity test passed. Engineering also found various scratch marks showing light material removal on the distal end of the main tube. The scratches were short in length and not axially aligned with the tube. Damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3064021 | 3036572 | 2933189 | 2955842201301123 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-05 | 2013-04-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-22 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure a broken wire was found between the cutters on the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken wire. The grip cable was broken at the distal idler. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Additional finding were scratches on the maintube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3064025 | 3036586 | 2933100 | 2955842201301124 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-05 | 2013-04-04 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-22 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure a wire on the pk dissecting forceps instrument was frayed. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a frayed wire. The instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis or other cables; however, scratches on the surface of the distal pulley were found. Electrical continuity test passed. Engineering also found the instrument tip was bent. One grip was bent, causing side to side misalignment of grips. There was a. 066 offset at the tips, indicating overloading at the tip. Damage likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3064124 | 3036675 | 2933189 | 2955842201301125 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-05 | 2013-04-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-03-04 | Malfunction |  | The user facility reportedly identified a broken cable from the grasping retractor instrument during central processing. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was broken at the distal clevis hub. The cable segment containing the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3061421 | 3033750 | 2930264 | 2955842201301097 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-04 | 2013-04-03 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42020705 | 2013-03-04 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the user facility identified broken wires on the tenaculum forceps instrument. There was no allegation of patient harm or fragments falling into a patient. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was broken at the distal clevis hub. The cable segment containing the crimp was still installed in the clevis and did not exhibit any damage or wear marks. There was no other damage found with the instrument. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3061425 | 3033751 | 2928236 | 2955842201301098 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-03 | No | Yes | Health Professional | 42020705 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-05 | Malfunction |  | The user facility reportedly identified a broken wire with the mega suturecut needle driver instrument prior to starting a da vinci si surgical procedure. The instrument was not used on a patient after the reported issue was identified. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment was sticking out at the instrument's wrist. No other cables at the instrument's wrist were damaged. Additional observation not initially reported by the site was pulley damage. The pulley had scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3061427 | 3033753 | 2930267 | 2955842201301099 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-03 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-14 | Malfunction |  | The user facility reportedly identified a broken cable on the monopolar curved scissors (mcs) instrument during central processing. The instrument reportedly was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the grip cable was frayed at the distal idler pulley and noted that the frayed strands were sticking out at the instrument's wrist. No other cables at the instrument's wrist were damaged. Electrical continuity testing of the instrument passed. The clinical risk evaluation performed on the health hazard assessment for the monopolar curve scissors states: in a fragment-causing failure (such as a cable failure, a crimp separating from a cable, etc) for the mcs instrument the tip cover accessory will hold such fragments inside the tip cover, preventing loss into the patient during the removal of the instrument. Because the tip cover acts as a barrier, these failures, which might be reportable in other instruments, are not reportable events for the mcs instrument. Additional observations not initially reported by the site was pulley damage and main tube damage. The pulley had scratches on the surface of the distal pulley. In addition, there was a crack found on the distal end of the main tube by the proximal side. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3061441 | 3033775 | 2903465 | 2955842201301100 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-06 | 2013-04-03 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the prograsp forceps instrument was not recognized. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. The instrument was checked for recognition on an in-house system which successfully recognized the instrument. The pogo pins did not stick and were not contaminated. Dcp verification of the instrument passed. An additional finding not reported was a loose pitch cable. The pitch cable was found loose at the distal clevis hub. Both cable crimps remained in the clevis and no damage was found on the cable. Upon removing the housing a pitch cable was found loose at the clamping pulley but no loose clamping pulley screw was found. Engineering also found scratches on the main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3061511 | 3033847 | 2923053 | 2955842201301101 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-03 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-04 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, a broken cable was identified on the fenestrated bipolar forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the broken cable. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. An electrical continuity test passed. Additional finding were scratches on the pulley. There were scratches on the surface of the distal pulley. Scratches on main tube were also found. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. Engineering also found the instrument tip was bent. One grip was bent, causing side to side misalignment of grips. There was a. 065 offset at the tips, indicating overloading at the tip. Engineering concluded the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3061618 | 3033958 | 2930472 | 2955842201301102 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42018906 | 2013-02-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the double fenestrated grasper instrument was not recognized by the da vinci si system. Reportedly the instrument was not used on a patient and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. The instrument was checked for recognition on an in-house system system, which successfully recognized the instrument. The instrument's pogo pins did not stick out and were not contaminated. The dallas chip programmer verification of the instrument passed. Engineering evaluation also found cable and main tube issues that were not reported by the site. One grip close cable was loose at the distal idlers. The idler pulley spun freely and was not damaged. The wire was also loose inside the housing by the clamping the pulley but the screws were found to be secured. The idler block showed no sign of damage. The distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the main tube damage may due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3061631 | 3033959 | 2930473 | 2955842201301103 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-03 | No | Yes | Health Professional | 42018906 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-13 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the jaws were misaligned on the maryland bipolar forceps instrument. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the reported issue. The instrument was returned with the grips misaligned. One grip tip was bent, causing side to side misalignment of the grips. There was a. 017 offset at the tips. Engineering concluded that damage was likely due to mishandling. An additional observation not reported by the site were deep scratches on the main tube. The distal end of the main tube had various scratches that exhibited light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. An electrical continuity test passed. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3061733 | 3034071 | 2930585 | 2955842201301104 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-05 | 2013-04-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-05 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, a wire was frayed and off the track on the pk dissecting forceps instrument. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported a frayed wire, however, for clarification the damaged instrument component was a broken pitch cable. Based on this clarification, the customer reported complaint was confirmed. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks however scratches on the surface of the distal pulley were found. Electrical continuity test passed. Additionally, the grip cable was derailed. One grip close cable was derailed at the distal idler pulley. The grips could still open and close, but movement may not have been precise. The cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. The fleet angle between the proximal and distal pulleys may have contributed to the derailment. Engineering found scratches on the main tube. The distal end of the main tube had various scratch marks showing light material removal. Scratches were short in length and not axially aligned with the tube. Engineering concluded the damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3061767 | 3034082 | 2928158 | 2955842201301105 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-05 | 2013-04-03 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-01 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, a cable on the large needle driver instrument snapped. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the reported issue. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the instruments wrist were not damaged. Engineering also found scratches on the on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3061768 | 3034084 | 2930598 | 2955842201301106 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-05 | 2013-04-03 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-04 | Malfunction |  | It was reported during processing,the cables from the prograsp forceps instrument were found to be frayed. Reportedly the instrument was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the grip cable was frayed and was sticking out at the instrument's wrist. No other damage was found with the other instrument's cables. Additional observation not initially reported by the site was a derailed grip cable associated with the same reported frayed grip cable. The grip cable was derailed at the distal idler pulley. The grips can still open and close but movement may not be precise. The cable derailment is likely due to cable losing contact with the pulley during wrist articulation. The fleet angle between the proximal and distal pulleys may have also contributed to the derailment. The last observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3058415 | 3030718 | 2927232 | 2955842201301072 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-02 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-04 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the wire from the maryland bipolar forceps instrument was broken. Reportedly the instrument was not used on a patient and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable at the distal idler was frayed. The pitch cable was also found to be sticking out at the bottom of the instrument. There was no other damage found on the pulley or clevis. Electrical continuity tests were performed and passed. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3058404 | 3030729 | 2927243 | 2955842201301073 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-02 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-22 | Malfunction |  | It was reported that prior to a da vinci s procedure and during processing, the user facility identified a broken cable from the large suturecut needle driver instrument. Reportedly the instrument was not used on a patient after the reported issue was identified. There was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the grip cable was broken at the distal idlers and was sticking out at the instrument's wrist. The idler pulley spun freely and was not damaged. No other cables at the instrument's wrist were damaged. An additional observation noted by engineering that was not reported by the site was indentation on the pulley. There was an indentation on the inner core of the pulley as well as scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE SUTURE CUT NEEDLEDRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3058438 | 3030769 | 2927283 | 2955842201301074 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-26 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure, the wires from the large suturecut needle driver instrument became unravelled. No missing or fallen pieces were reported to have fallen into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that the grip cable was broken and was sticking out at the instrument's wrist. The idler pulley spun freely and exhibited pulley rim damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE SUTURE CUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3058443 | 3030806 | 2927320 | 2955842201301075 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-12 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure, the monopolar curved scissors (mcs) instrument was not closing while the surgeon was performing a bladder neck resection. The mcs instrument with the mcs tip cover accessory installed was removed and upon inspection of the mcs instrument, the cable on the instrument was noted to be broken. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The electrical continuity testing passed. Engineering also found the blades were corroded both sides with visible missing materials eaten up. The damage was most likely due to improper cleaning. Engineering also found various scratch marks on the distal end of the main tube showing light material removal. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the damage to the instrument's main tube, which found during failure analysis evaluation could likely cause or contribute to an adverse event if the failure mode were to recur. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3059269 | 3031611 | 2928125 | 2955842201301076 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-03-06 | 2013-04-02 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42020505 | 2013-02-22 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the fenestrated bipolar forceps instrument tips did not align. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of the grips. There was a. 127 offset at the tips, indicating overloading at the tip. The electrical continuity testing passed. Engineering also found a frayed and loose pitch cable at the distal clevis hub. No other cable damage was found. Engineering also found scratches on the distal pulley. There were various scratch marks on the distal end of the main tube showing light material and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3059297 | 3031614 | 2908456 | 2955842201301077 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42003306 | 2013-02-22 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the black diamond micro forceps instrument was noted to have a bent tip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the main tube was broken by the proximal area. The clevis was dislodged from the main tube. There were missing pieces from the main tube but engineering was unable to measure due to how it was damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | BLACK DIAMOND MICRO FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3059298 | 3031615 | 2928129 | 2955842201301078 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42003306 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-05 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, the large suturecut needle driver instrument was noted to have a separated cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist are not damaged. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3059342 | 3031624 | 2928169 | 2955842201301084 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-31 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the fenestrated bipolar forceps instrument suddenly stopped responding to the system. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to duplicate the customer reported complaint. Engineering found the instrument was checked for recognition on an in-house system; the system successfully recognized the instrument. The pogo pins did not stick and were not contaminated. The dcp (dallas chip programmer) verification of instrument passed. The instrument was driven and moved intuitively and the grips opened and closed. Engineering found various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. Engineering also found scratches on the pulley. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. There are scratches on the surface of the distal pulley. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3059315 | 3031630 | 2928144 | 2955842201301079 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42805512 | 2013-03-05 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure the cadiere forceps instrument wrist snapped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the distal clevis was broken on both sides at the proximal clevis. No additional damage was found on the distal end. Engineering concluded that damage was likely due to excessive side loading at the distal tip during use or other type of mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS SINGLESITE INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3059302 | 3031631 | 2928145 | 2955842201301080 | 1 | GCJ | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42805512 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-05 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the pk dissecting forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one conductor wire was broken at the yaw pulley exit. The wire was detached from the connection at the grip. The electrical continuity testing failed. The yaw pulley showed no signs of arcing. The yaw pulley was missing a. 220 x. 047 piece opposite of the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. It was unclear how the piece of yaw pulley broke. Engineering also found the one grip was bent, causing side-to-side misalignment of the grips. There was a. 059 offset at the tips, indicating overloading at the tip. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3059335 | 3031644 | 2928158 | 2955842201301081 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-05 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the large suturecut needle driver instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the grip cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at wrist were not damaged. Engineering also found an indentation on the inner core of the pulley plus scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3059336 | 3031646 | 2928160 | 2955842201301082 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-05 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure the large suturecut needle driver instrument was noted to have a separated cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damages on the surface. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found slight white residues on the clamping pulleys and the cables inside the housing. Engineering concluded that the damage was likely due to mishandling during cleaning. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3059348 | 3031655 | 2928169 | 2955842201301085 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-03-05 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical large suturecut needle driver instrument was noted to have a separated cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument grip close cable was found frayed and derailed at the distal idler pulley. The grips can still be opened and closed, but movement may not be precise. Physical damages were found on distal idler pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3059344 | 3031666 | 2928180 | 2955842201301083 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-05 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure and during central processing, the user facility identified frayed cables from the pk dissecting forceps instrument. Reportedly the instrument was not used on a patient after the reported issue was identified. There was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was broken at the distal end. The clevis segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear mark. The other cables at the instrument's wrist were not damaged. Electrical continuity test were performed and passed. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 3059374 | 3031678 | 2928158 | 2955842201301086 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42019007 | 2013-02-22 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the cable from the cobra grasper instrument was identified to be frayed but not completely separated. No missing or fallen pieces were reported to have fallen into the patient. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed that the grip cable was frayed at the instrument's wrist. The frayed segments were various in length. No damage was found on the pulley or distal clevis hub. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3059362 | 3031686 | 2928200 | 2955842201301087 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42019007 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42018404 | 2013-03-05 | Malfunction |  | It was reported that prior to a da vinci s surgical procedure and during reprocessing, the user facility identified a broken wire from the permanent cautery spatula instrument. Reportedly the instrument was not used on a patient after the reported issue was identified. There was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was broken at the distal end. The clevis segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear mark. The other cables at the instrument's wrist were not damaged. Electrical continuity testing of the instrument passed. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3059385 | 3031709 | 2928223 | 2955842201301088 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-05 | 2013-04-02 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42020705 | 2013-03-04 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the user facility identified broken wires from the tenaculum forceps instrument. Reportedly, the instrument was not used on a patient and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was broken at the distal clevis hub. The cable containing the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Additional observation not initially reported by the site was main tube damage. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal. The scratches were short in length and were not axially aligned with the tube. Evidence not conclusive but the main tube damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3059433 | 3031722 | 2928236 | 2955842201301089 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-02 | No | Yes | Health Professional | 42020705 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42019007 | 2013-02-27 | Injury | Required Intervention | It was reported that during a da vinci s hysterectomy procedure, a piece from the cobra grasper instrument broke off and fell into the patient. The broken instrument piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that instrument was returned with the distal pin broken off along with the pulleys. The pin and pulleys were returned with the instrument. The cable remained intact. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3059458 | 3031773 | 2928287 | 2955842201301090 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-02 | Yes | Yes | Health Professional | 42019007 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-16 | Injury | Required Intervention | It was reported that during a da vinci si ureteral reimplantation procedure, one of the cables at the end of the clamp on the fenestrated bipolar forceps instrument broke and fell into the patient. The broken piece of cable was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering evaluation also found that the distal end of the instrument's main tube had various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that damage to the main tube was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, isi contacted the initial reporter of this event. The initial reporter indicated that the instrument was inspected prior to use and during the inspection no issues were noted. The initial reporter indicated that the surgeon was performing a da vinci si ureteral reimplantation procedure when the cable on the fenestrated bipolar forceps instrument broke. The initial reporter indicated that the instrument did not make contact with any other instrument during the surgical procedure and that it is unknown as to what caused the cable on the instrument to break. The initial reporter indicated that no additional surgical procedure was required to remove the cable fragment from the patient and that the fragment was recovered during the surgical procedure. The initial reporter has denied that the patient has returned to the hospital due to any post- surgical complications as a result of the reported event. No other information was provided. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 3059428 | 3031774 | 2928288 | 2955842201301095 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-05 | 2013-04-02 | Yes | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-05 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-14 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the fenestrated bipolar forceps instrument jaws did not meet as they should. The surgical staff swapped out the instrument for a different instrument to complete the procedure. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint that the instrument jaws do not meet. Engineering found a broken pitch cable. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. Engineering also found scratches on the surface of the distal pulley. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3059461 | 3031775 | 2928289 | 2955842201301096 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-06 | 2013-04-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-31 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the prograsp forceps instrument was not recognized. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to duplicate the customer reported complaint. The instrument was placed on an in-house system and successfully recognized. The pogo pins did not stick and were not contaminated. Dcp (dallas chip programmer) verification of the instrument passed. Engineering found 48 - engineering also found a broken pitch cable at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering also found various scratch marks on distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube, which found during failure analysis evaluation could likely cause or contribute to an adverse event were to recur. | PROGRASP FORCEPS ISNTRUMENT | ENDOSCOPIC INSTRUMENT | 3059457 | 3031789 | 2928303 | 2955842201301091 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-06 | 2013-04-02 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-06 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical fenestrated bipolar forceps instrument cables was noted be frayed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. The electrical continuity testing passed. Engineering also found scratches on the surface of the distal pulley. Engineering also found various scratch marks on the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3059427 | 3031791 | 2928305 | 2955842201301092 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-06 | 2013-04-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-05 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the fenestrated bipolar forceps instrument cable broke and would not allow the surgeon to rotate the instrument correctly. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was broken at the wrist. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3059460 | 3031792 | 2928306 | 2955842201301093 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-06 | 2013-04-02 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-03-04 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the large needle driver instrument was noted to have a wire completely separated and hanging out. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found indentations on the edge of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3059429 | 3031793 | 2928307 | 2955842201301094 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-06 | 2013-04-02 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
| 2013-03-27 | Malfunction |  | Da vinci vessel sealer was being used during surgery when an error code appeared on vision cart monitor: "blade exposed". Staff was unable to reuse the instrument so it was taken out and saved to be submitted to the manager. This is the third of these devices we have reported. What was the original intended procedure?robotic gyn surgery. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST ONE | INSTRUMENT COMPUTER CONTROLLED | 3076516 | 3049343 | 2945857 | 3049343 | 1 | NAY | User Facility |  |  | Initial | 2013-04-02 | 2013-04-02 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
42805706 | 2013-02-05 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the customer noted that the curved scissors single site instrument was not being recognized by the da vinci robot system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on the in-house da vinci robot system and driven with no problems. Instrument passed recognition and engagement tests. Additional observations not reported by site were broken main tube and deep scratches on main tube. Instrument main tube was found to be broken at the proximal end where tube enters the housing. Tube insulation also exhibited scratches and gouge marks near distal end. Engineering concluded that damage was likely due to mishandling. The single site instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Insertion and removal of instruments use care when inserting single-site instruments into the cannula to avoid catching the tip on the seal or cannula bowl and bending the shaft excessively. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | CURVED SCISSORS SINGLESITE INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3055603 | 3028285 | 2924799 | 2955842201301042 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-04-01 | No | Yes | Health Professional | 42805706 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-01 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the customer noted that the 'curb was unsticked' on the maryland bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent, causing side to side misalignment of grips. There was a. 0315 offset at the tips. Engineering concluded that damage was likely due to mishandling. Additional observation not reported by site was damage housing. The top left corner of the luer plate located at the back of the chassis was cracked. Engineering concluded that damage was likely due to mishandling. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3055531 | 3028298 | 2924807 | 2955842201301041 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-02-27 | 2013-04-01 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-26 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure, the customer noted a broken cable located at the distal end of the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely, but it exhibited some slight damages. The cable segment stuck out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3055541 | 3028306 | 2924815 | 2955842201301031 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-04-01 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
428065 | 2013-02-28 | Injury | Required Intervention | It was reported that during a da vinci si cholecystectomy procedure, the surgical staff noticed that a fragment from the single-site port accessory broke off and fell into the patient. There was no allegation of any harm, injury, or adverse outcome to a patient. | The single-site port has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, isi contacted the isi csr who was present during the surgical procedure. The csr indicated that the single-site port was inspected prior to use and that the surgeon did not experience any difficulty or issues inserting the single site port accessory into the patient. However, after the endoscopic camera manipulator (ecm) was docked to the patient, the surgeon observed that a fragment the size of a crumb from the single site port accessory had broken off and fell into the patient. Using two endowrist grasper instruments, the surgeon attempted to retrieve the broken fragment; however, the fragment was not located. The csr indicated that the surgeon attempted to locate the broken piece; however, after approximately 15 minutes the surgeon made the decision to discontinue his attempts to retrieve the fragment. The single-site port accessory was replaced and the planned surgical procedure was completed with no further issues. Per the csr, the surgeon indicated to her that the patient is recovering well and has not experienced any post-surgical complications as a result the reported event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | SINGLESITE PORT ACCESSORY | ENDOSCOPIC ACCESSORY | 3055595 | 3028337 | 2924846 | 2955842201301043 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-28 | 2013-04-01 | Yes | Yes | Health Professional | 428065 | No | No | 2013-02-28 | No | Yes | Initial | 2013
42017910 | 2013-02-22 | Malfunction |  | Event Description: --/-- It was reported that during reprocessing of a monopolar curved scissors instrument, the tips would not open. There were no missing pieces or fallen pieces reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. The complaint of the tips not opening could not be confirmed. The instrument was placed on a is3000 system and driven. Recognition and engagement test passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Functional performance testing did not find any issues. Additional observations not reported by site were a bent banana plug and removed pad printing. The banana plug was bent to one side. A monopolar cord could still be installed, but would not seat flush against the chassis. Electrical continuity still passes. The tube extension had various areas with missing pad printing. All areas with missing pad printing had toothlike gouge marks that were also present on the shoulder section below the pad printing and on the tube reinforcement ring. Both issues were likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3055712 | 3028449 | 2924963 | 2955842201301047 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-02-28 | 2013-04-01 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42011006 | 2013-02-14 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the surgeon noticed the feeling in the left hand was not normal. The surgical staff removed the precise bipolar instrument from arm number two and visualy inspected it. The nurse found that the jaw was not engaged completely and that the jaw shifted in the direction vertical to the open and close direction. The staff replaced the instrument and the planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the reported issue. The instrument was retuned with the grips not aligned with each other. One grip tip was bent,causing side to side misalignment of the grips. There was a. 0450 offset at the tips. Damage was likely due to mishandling. Additional observation not reported by site was deep scratches on main tube. The distal end of main tube had various scratches that exhibited light material removal and a rough surface finish. The scratches varied in size but were mostly small and thin. Damage is likely due to mishandling. Electrical continuity test passed. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3055717 | 3028473 | 2924982 | 2955842201301048 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-26 | Malfunction |  | It was reported that during a da vinci s surgical procedure a cable snapped on the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the cable snapped. One grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had slight damage on the surface. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3055718 | 3028474 | 2924983 | 2955842201301052 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-02-20 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tenaculum forceps instrument was not recognized by the da vinci si system. Reportedly the instrument was not used on a patient and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering was unable to replicate the alleged recognition issue. The returned instrument was successfully recognized by the in-house is3000 system. The pogo pins did not stick out and they were not contaminated. A dallas chip programmer verification of instrument was conducted and passed. Additional observations not reported by the site were frayed cable and tube damage. One grip close cable was frayed at the distal idler pulley. The frayed strands were sticking out at the instrument's wrist. Other cables at instrument's wrist were not damaged. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal. The scratches were 0. 070 - 0. 160 in length in length and were not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3055700 | 3028475 | 2924984 | 2955842201301044 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-04-01 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-02-26 | Malfunction |  | It was reported that during a da vinci s surgical procedure a wire broke while sewing with the mega needle driver instrument. A fragment was reported to have been retrieved. The instrument was replaced with a new one and the procedure was completed. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3055753 | 3028501 | 2925010 | 2955842201301046 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42019408 | No | No | 2013-03-01 | No | No | Unkown | 2013
42020505 | 2013-01-24 | Malfunction |  | It was reported that during a da vinci si colectomy procedure the fenestrated bipolar forceps would not grasp. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed with a replacement instrument of the same type. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The grasping functionality was reduced by bent grips. Both grips were bent and one was twisted so that there was a. 160 gap at the tips when grips were in the closed position. The bent grips had separation between yaw pulley and yaw pulley cover, indicating overloading at the tip. Additional observation not reported by site was tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 175 in length and not aligned with the tube axis. Grip and tube damage was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3055772 | 3028506 | 2892138 | 2955842201301049 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-11 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the jaws of the fenestrated bipolar forceps instrument were not aligned. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the jaws were not aligned. One grip was bent, causing side to side misalignment of the grips. There was a. 122 offset at the tips, indicating overloading at the tip. Electrical continuity test passed. Additional finding were scratches on the maintube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Damage is likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3055775 | 3028509 | 2925018 | 2955842201301050 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-03-01 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure part of the tips fell off the small clip applier instrument before it was inserted into the patient. There were no missing pieces or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported issue. Both clip applier grips were broken toward the tip point. The customer did not return the broken parts back with instrument. The damage is likely due to mishandling. Additional finding is scratches on pulley. There were scratches on the surface of the distal pulley. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3055780 | 3028514 | 2926245 | 2955842201301053 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-01 | Malfunction |  | It was reported that during reprocessing that a broken wire was noted on the maryland bipolar forceps instrument. There were no missing pieces or fallen pieces reported. | The instrument was returned and evaluated. The complaint of a broken wire was noted while cleaning was not confirmed. The instrument grip cable was found derailed at the distal end. One of the grip cables slipped off the idler pulley. The grips could still open and close but movement may not be precise. No other damages found on distal clevis. Additional damage found was deep scratches on the main tube. Distal end of main tube has several scratch marks exhibiting light material removal and a rough surface finish. Scratches were short in length and not axially aligned with the tube. Damage may be due to mishandling. Electrical continuity passed. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3055758 | 3028527 | 2925021 | 2955842201301051 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42001103 | 2013-02-11 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the trocar was dropped prior to be inserted in the patient side manipulator (psm) arm. No patient harm, adverse outcome or injury was reported. | The cannula was returned and evaluated. Per the customer reported complaint, engineering observed that the cannula was found with a physical damage. The cannula exhibited a flat surface at the shaft opening. Engineering concluded that damage was likely caused when the cannula was dropped by the customer. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: inspection instructions warning: cannula defects can be caused by dropping the cannula or handling it roughly. | 5MM CANNULA ACCESSORY | ENDOSCOPIC INSTRUMENT ACCESSOR | 3055923 | 3028641 | 2925150 | 2955842201301061 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42001103 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42019407 | 2013-03-01 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the wire of the mega needle driver instrument snapped. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at the wrist. Other cables at wrist were not damaged. Engineering also observed some scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3055895 | 3028646 | 2925160 | 2955842201301054 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42019407 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-19 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted bent prongs on the connection of the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the bottom bipolar pin was found to be bent sideways. A bipolar cord could not be properly installed due to the bent pin. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were 0. 085 -. 150 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3055905 | 3028664 | 2908477 | 2955842201301057 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-03-01 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the customer noted a broken cable on the tenaculum forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3055934 | 3028669 | 2925158 | 2955842201301056 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42020706 | No | No | 2013-03-01 | No | No | Reuse | 2013
42017910 | 2012-12-27 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si cholecystectomy procedure, the customer noted that the monopolar curved scissors instrument was dull. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, the instrument was placed on an in-house da vinci robot system to perform a latex cut test. The monopolar hot shears curved scissors instrument did not cut cleanly through 0. 006 latex. The latex got snagged at instrument's scissors tips. The blade edges were not damaged, but the edges exhibited wear at the tips, negatively affecting cut performance. Additional observation not reported by site pad printing damage. The orange pad printing was found to be removed from one side of the instrument's tube extension. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3055908 | 3028671 | 2925200 | 2955842201301060 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42009308 | 2013-02-18 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the prograsp forceps instrument was not being recognized by the da vinci robot system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was checked for recognition on an in-house da vinci si robot system. The instrument was successfully recognized by the da vinci robot system, showing 5 uses left. Pogo pins did not stick and were not contaminated. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were 0. 075 - 0. 275 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3055944 | 3028678 | 2925207 | 2955842201301058 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-01 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a frayed cable on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3055937 | 3028689 | 2925183 | 2955842201301055 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-27 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure, the customer noted that the tips of the maryland bipolar forceps instrument did not meet, and would not hold tissue. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent causing side to side misalignment of grips. There was a 0. 030 offset at the tips. Bent grip exhibited a separation of the yaw pulley and the pulley cover at the glue joint, indicating overloading at the tip. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were 0. 060 - 0. 105 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3055949 | 3028693 | 2925187 | 2955842201301059 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-04-01 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-03-04 | Malfunction |  | It was reported that during processing, the pulley wires from the prograsp forceps were found to be exposed at the end of the instrument. Reportedly the instrument was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. The instrument's pulley wires were confirmed to be exposed. The pitch cable at the distal end of the instrument was frayed. The cable segment that contains the crimp is still installed in the clevis. There was an indentation next to the groove where the cable was frayed. Evidence not conclusive, but the cable damage was most likely linked to the dent found at the proximal clevis. No other damaged cables were found. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length , suggesting it might be caused by instrument collisions or other type of mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3057021 | 3029598 | 2926106 | 2955842201301062 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-01 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-04 | Malfunction |  | It was reported that during processing, the mega suturecut instrument had pulley wires that were exposed at the end of the instrument. Reportedly the instrument was not used on a patient after the reported issue was identified and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable was frayed and was sticking out at the instrument's wrist. No other damage was found at the clevis or with other cables. Additional observation not initially reported by the site was insulation damage on the distal end of the main tube. Engineering noted that material was missing and that the surface of main tube appeared to be scraped off. Damage is approximately 14 away from the base of the housing. Engineering concluded that the damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3057036 | 3029603 | 2926111 | 2955842201301063 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-01 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-11 | Malfunction |  | It was reported during a da vinci si sacropolpexy procedure the pk dissecting forceps instrument wrist got stuck rotating in the middle of the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp remained in the clevis. The clevis did not exhibit excessive damage. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3057143 | 3029727 | 2926241 | 2955842201301064 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-04 | 2013-04-01 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-02-19 | Malfunction |  | It was reported that during reprocessing the da vinci s surgical small grasping retractor instrument was noted to have a loose wire at the distal end after it was cleaned and sterilized. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was found loose at distal the clevis hub. Both cable crimps remained in the clevis and no damage was found on the cable. Upon housing removal the pitch cable was found loose at the clamping pulley, but no loose clamping pulley screw was found. Engineering also found a frayed pitch cable. The instrument was found with a frayed pitch cable at the distal clevis hub. The frayed pitch cable was found in the same area as the loose wire. No damage was found at the clevis. Engineering also found various scratch marks on the main tube at the distal end showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. Engineering opened the housing and found corrosion on the clamping pulleys. Engineering concluded that damage was likely due to improper cleaning. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR 8MMINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3057160 | 3029729 | 2926243 | 2955842201301065 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-01 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-02-26 | Malfunction |  | It was reported during a da vinci si surgical procedure the small clip applier instrument was noted to have broken tips, separated from the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that both clip applier grips were broken off and broken pieces were not returned. The grips were fractured at midpoint of the tip. Engineering concluded that damage was likely due to mishandling. Engineering also found indentations on the distal idler pulley and deep scratch on the main tube. The distal idler pulleys had several indentations along the edges. The distal end of the main tube had a scratch mark exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3057168 | 3029731 | 2926245 | 2955842201301070 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-01 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-03-04 | Malfunction |  | It was reported during a da vinci s surgical procedure the fenestrated bipolar forceps instrument was noted to have a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one conductor wire was broken at the yaw pulley exit. The wire was detached from the connection at the grip. The electrical continuity testing failed. The yaw pulley showed no signs of arcing. Engineering also found a piece of the yaw pulley broken off the opposite side of the broken conductor wire. The missing piece measured roughly about 0. 045 x 0. 193. The broken piece was not returned with the instrument. Engineering found various scratch marks on the main tube at the distal end showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3057169 | 3029732 | 2926246 | 2955842201301071 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-01 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-03-04 | Malfunction |  | It was reported prior to starting a da vinci s surgical procedure the pk dissecting forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one conductor wire was broken at the yaw pulley exit. The wire was detached from the connection at the grip. The wire stuck out approximately. 2275 at the wrist. The electrical continuity testing failed. The yaw pulley showed no signs of arcing. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3057163 | 3029747 | 2926235 | 2955842201301067 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-01 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-26 | Malfunction |  | It was reported prior to starting a da vinci s surgical procedure the fenestrated bipolar forceps instrument tips was misaligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip was bent, causing side-to-side misalignment of grips. There was a. 087 offset at the tips, indicating overloading at the tip. The electrical continuity testing passed. 48 - engineering also found various scratch marks on the main tube at the distal end showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3057162 | 3029748 | 2926236 | 2955842201301066 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-01 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-03-04 | Malfunction |  | It was reported during a da vinci si surgical procedure the grasping retractor instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip cable was broken at the distal clevis. The other cables at the wrist were not damaged. The idler pulley spun freely and did not exhibit damage. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3057164 | 3029749 | 2926237 | 2955842201301068 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-01 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-19 | Malfunction |  | It was reported during a da vinci si total benign hysterectomy procedure the large suturecut needle driver instrument was noted to have a broken wire at the tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken grip cable. The broken grip cable stuck out at the wrist. The idler pulley spun freely and exhibited pulley face and rim damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3057165 | 3029750 | 2918672 | 2955842201301069 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-04 | 2013-04-01 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-03-04 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-29 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy w/hysterectomy procedure, the customer noted broken wires at the distal end of the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3053896 | 3027006 | 2923520 | 2955842201301012 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42011006 | 2013-02-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was noted that a wire was sticking out on the precise bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable located at the distal clevis hub. There was no damage found at the clevis. The frayed segment was approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PRECISE TM BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3053985 | 3027064 | 2923578 | 2955842201301013 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-27 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the wire of the fenestrated bipolar forceps instrument was sticking out. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054067 | 3027200 | 2923801 | 2955842201301014 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-27 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that wires of the fenestrated bipolar forceps instrument were sticking out. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054147 | 3027287 | 2923801 | 2955842201301015 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-25 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the customer noted broken wires on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at wrist. Other cables at wrist were not damaged. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 3054186 | 3027303 | 2924448 | 2955842201301016 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42019007 | 2013-02-08 | Malfunction |  | It was reported that during a davinci s gynecology procedure, the customer noted a frayed and snapped cable towards the tip of the cobra grasper instrument. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a broken grip cable sticking out at the wrist. There was no damage found on idler pulley. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was missing. The clevis did not exhibit any damage or wear marks. Additional observation not reported by site was abnormal wear on the main tube. The surface of the main tube was splintering off. The blue housing was removed and the clamping pulley showed signs of excessive residue residing around the pitch and grip cable. Engineering concluded that damage was likely due to improper cleaning process. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054279 | 3027414 | 2923928 | 2955842201301017 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42019007 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-19 | Malfunction |  | The jaw from the fenestrated bipolar forceps instrument was identified to be misaligned during a da vinci si myomectomy procedure. Nothing reportedly fell into the patient and there was no allegation of patient harm. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip of the instrument was bent, causing side to side misalignment of grips. There was a 0. 109 offset at the tips, indicating overloading at the tip. Electrical continuity passed. Recurrence of the alleged failure mode is not expected to be likely to cause or contribute to a death, serious injury or serious deterioration in the state of health of a patient. An additional observation was noted for the instrument's main tube during investigation. The distal end of main tube had various scratch marks showing light material removal and a rough surface finish. Scratches were short in length and not axially aligned with the tube. Evidence not conclusive, but damage may be due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054341 | 3027467 | 2923981 | 2955842201301018 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-03-01 | Malfunction |  | It was reported during a da vinci si total benign hysterectomy procedure the mega suturecut needle driver instrument cable popped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found indentations on the edge of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054345 | 3027474 | 2923988 | 2955842201301020 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-03-01 | 2013-03-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-03-01 | Malfunction |  | It was reported during a da vinci si surgical procedure the maryland bipolar forceps instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the damaged instrument component was a frayed pitch cable. The instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. The electrical continuity testing passed. Engineering also found various scratch marks on the main tube at the distal end showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054381 | 3027480 | 2923994 | 2955842201301021 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-03-29 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-26 | Malfunction |  | It was reported during a da vinci si surgical procedure, the maryland bipolar forceps instrument insulator had sparked. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that both yaw pulleys exhibited charring and localized melting at the grip base, between the grips. The charring was a sign of an arcing event. The evidence not conclusive, but charring may have been due to a breach in the insulation, carbonized tissue creating a conductive path, or high generator settings. Engineering also found various scratch marks on the main tube at the distal end showing light material removal. The scratches were. 095 -. 180 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054384 | 3027482 | 2923996 | 2955842201301023 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-01 | 2013-03-29 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-22 | Malfunction |  | The instrument was returned, no further information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found a bent grip and tube damage. The one grip was bent, causing side-to-side misalignment of grips. There was a. 045 offset at the tips. The bent grip did not exhibit separation of the yaw pulley and pulley cover at the glue joint, but the likely cause was still overloading at the tip. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 200 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054386 | 3027484 | 2923998 | 2955842201301022 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-03-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-07 | Malfunction |  | On (b)(6) 2013, it was reported that the fenestrated bipolar forceps instrument was not working. The issue was identified prior to use and was not used on a patient. There was no allegation of harm or injury to patients. | On (b)(4) 2013, intuitive surgical received the instrument involved with this complaint. Engineering investigated the returned product and noted that the alleged complaint was not confirmed. An additional finding was noted during investigation, which was not initially reported by the reporter. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054377 | 3027489 | 2924003 | 2955842201301019 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-28 | 2013-03-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-02-28 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-02-12 | Malfunction |  | It was reported during a da vinci si prostatectomy procedure the mega needle driver instrument clamp was noted to be dislocated at the base of the scissor. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to replicate the customer reported complaint. Engineering found the instrument with frayed cable at distal idler pulley. The frayed cable section was approximately 0. 03 in length. No damage was found on the pulley. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054412 | 3027507 | 2924021 | 2955842201301024 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-01 | 2013-03-29 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-22 | Malfunction |  | It was reported that after a da vinci si renal cyst decortication procedure, the customer noted a snapped cable on the large needle driver instrument. The customer indicated that nothing fell into a patient during a surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a broken grip cable sticking out at the wrist. The idler pulley was able to spin freely. Additional observation not reported by the site was indentations and ridges to the pulley. Engineering concluded that the damage was likely due to excessive force. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054515 | 3027562 | 2924076 | 2955842201301025 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-19 | Malfunction |  | Customer reported that the mega suture cut needle driver instrument was inspected as usual prior to being inserted into the patient. The surgeon found the instrument to be partially responsive during the da vinci s hysterectomy procedure. Upon inspection under the endoscope, it was noticed that a cable was broken. The instrument was replaced by a new one and the surgery was completed without any additional problem. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Engineering also found scratches on the surface of both distal pulleys. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURE CUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054595 | 3027625 | 2924139 | 2955842201301026 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-07 | Malfunction |  | Customer reported that the fenestrated bipolar forceps instrument was inspected as usual prior to being inserted into the patient. The surgeon found the instrument to be non-responsive during the da vinci si prostatectomy procedure. The instrument was replaced by a new one and the surgery was completed without any additional problem. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one conductor wire was found to be broken at the yaw pulley exit. The wire was found to be detached from its connection at the grip. Instrument passed electrical continuity test. The yaw pulley exhibited no signs of arcing. Engineering also found scratches on the surface of the distal pulley. 48 - additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054624 | 3027673 | 2924187 | 2955842201301027 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-02-26 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the cable on the tenaculum forceps instrument broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the alleged broken cable complaint. The pitch down cable was broken at the distal clevis hub. The cable segment that contains the crimp was returned detached from the distal end. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054707 | 3027730 | 2924244 | 2955842201301028 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-02-28 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tip on the mega needle driver instrument was observed broken. There was no allegation of patient harm or fragments falling into a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable at the instrument's tip was broken and stuck out at the instrument's wrist. The idler pulley spun freely and did not exhibit any visible damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054786 | 3027821 | 2924335 | 2955842201301030 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-29 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
42018907 | 2013-02-27 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted a frayed wire at the distal end of the double fenestrated grasper instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at wrist. Other cables at wrist were not damaged. Engineering also found scratches on the surface of the distal pulley. 48 - additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054919 | 3027902 | 2924416 | 2955842201301029 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42018907 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-25 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the prograsp forceps instrument was not recognized by the system. Reportedly the instrument was not used on a patient and there was no allegation of harm or injury to a patient. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument passed recognition and engagement testing on an in-house is3000 system. The pogo pins did not stick and did not show any signs of contamination. The instrument was driven and moved intuitively and the grips opened and closed. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054899 | 3027911 | 2924425 | 2955842201301032 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-29 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-28 | Malfunction |  | It was reported that the pk dissecting forceps instrument's cables broke. This reported issue was identified during a da vinci s hysterectomy procedure. The reporter indicated that nothing fell into the patient and there was no patient harm. | The instrument was returned and evaluated. Engineering confirmed that the pitch cable at the distal clevis hub was broken. The cable segment that contains the crimp was still installed in the clevis and did not exhibit any damage or wear marks. An electrical continuity test was performed with the instrument and passed. Additional observations found during investigations that were not reported by site was damage to the instrument's tip and main tube. One grip was bent causing side to side misalignment of grips. There was an 0. 053" offset at the tips, indicating overloading at the tip. The distal end of main tube had various scratch marks showing light material removal. Scratches were short in length and not axially aligned with the tube. Evidence not conclusive, but damage may be due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054934 | 3027916 | 2924430 | 2955842201301036 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-29 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-26 | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument's cable broke during a da vinci si hysterectomy procedure. The reporter indicated that nothing fell into the patient and there was no patient harm. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054923 | 3027927 | 2924441 | 2955842201301033 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-29 | No | Yes | Health Professional | 42020505 | No | No | 2013-02-28 | No | No | Reuse | 2013
42020706 | 2013-02-26 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure a broken cord on the tenaculum forceps instrument was found. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the broken cord with the following clarification; the cable was frayed. One grip close cable was frayed at the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at wrist were not damaged. Additional observation not reported by site was tube damage. The distal end of main tube had various scratch marks with light material removal. The scratches were. 080 -. 425 in length and not aligned with the tube axis. Damage may have been caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. This the customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054924 | 3027928 | 2924244 | 2955842201301034 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-26 | Malfunction |  | It was reported that the mega suturecut needle driver instrument's cable snapped. This reported issue was identified during a da vinci si hysterectomy procedure. The reporter indicated that nothing fell into the patient and there was no patient harm. | The instrument was returned and evaluated. Engineering confirmed that one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment also stuck out at the instrument's wrist. Additional observation not reported by site was damage to the pulley. There were couple of indentations on the edge of the distal pulley, suggestive of excessive force to the pulley on an event. Evidence not conclusive, but pulley damage most likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054937 | 3027934 | 2924448 | 2955842201301037 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-29 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-14 | Malfunction |  | It was reported that prior to starting a da vinci s procedure the large needle driver instrument did not seat in the robot correctly. The instrument was not used in the case. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the instrument did not seat in the robot correctly. The instrument was installed on an in-house is3000 system for a functional test. One grip axis input did not engage properly. Visual inspection showed the input disc had excessive lateral play. The housing was removed to find the grip axis bearing with a loose cage feature. The loose bearing feature resulted in lateral play of the input disk/shaft, causing possible engagement/motion issues. 48 - additional observations not reported by site were corroded clamping pulleys and tube damage. Clamping pulleys exhibited discoloration and signs of light pitting corrosion. The main tube exhibited wear marks with light material removal and a rough surface finish throughout the entire length of the tube (prox, mid, distal). Both failures may be caused by improper cleaning. No other damage was found. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3054942 | 3027941 | 2918519 | 2955842201301035 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-03-29 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 |  | Injury | Death | It was reported that after completion of a da vinci si hysterectomy procedure performed on (b)(6) 2013, the patient expired during recovery. No other details were provided. | On (b)(6) 2013, intuitive surgical inc. (isi) contacted the isi csr who initially reported the event to obtain additional information concerning the reported event. The csr indicated that he attempted to obtain additional information regarding the reported event from the surgeon and robotics coordinator at the site; however, the surgeon and robotics coordinator declined to provide any information regarding the reported event. The csr stated that he was present for approximately 30 minutes during the surgical procedure and during that time he did not observe that any complications had occurred. The csr indicated that the site has not made any allegations that the da vinci si system, instruments, and accessories caused or contributed to the patient's demise. Isi has also made several attempts to contact the risk management department to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. The site has declined to release any information regarding the reported event. A review of the da vinci si system logs showed no system malfunction occurred during the procedure performed on (b)(4) 2013. As of the date of this report, isi has not received any complaints of a malfunction of the da vinci si system, instruments, or accessories related to this event. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3054964 | 3027945 | 2924459 | 2955842201301039 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-03-01 | 2013-03-29 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-03-01 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-02-28 | Malfunction |  | It was reported that during processing, the cables from the mega needle driver instrument were exposed. The instrument reportedly was not used on a patient after the reported issue was identified. The initial reporter did not report any patient harm or fragments falling into a patient as a result of the reported issue. | Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. The instrument was returned and evaluated. Engineering confirmed that the instrument's grip cable was exposed. The grip cable was frayed at the distal idler pulley and distal clevis and the frayed strands were sticking out at the instrument's wrist. The other cables at the instrument's wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054965 | 3027946 | 2924460 | 2955842201301038 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-29 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-28 | Malfunction |  | It was reported that during processing, the wires from the fenestrated bipolar forceps were identified to be frayed. The instrument reportedly was not used on a patient after the reported issue was identified. The initial reporter did not report any patient harm or fragments falling into a patient as a result of the reported issue. | Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. The instrument was returned and evaluated. Engineering confirmed that the instrument's grip cable was frayed at the distal idler pulley and the frayed strands were sticking out at the instrument's wrist. The other cables at the instrument's wrist were not damaged. Electrical continuity tests were performed and passed. Engineering evaluation also found that the distal end of the instrument's main tube exhibited various scratch marks with light material removal and had a rough surface finish. The scratches were short in length and were not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3054944 | 3027966 | 2924480 | 2955842201301040 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
IS3000 A60P3 | 2013-02-27 | Other | Other | It was reported that prior to starting a da vinci si surgical procedure, the surgeon experienced a yellow and green image through the right eye of the high resolution stereo viewer on the surgeon side cart (ssc). The site installed their backup camera cable; however, the image through the right eye went blank, with the assistance of an isi technical support engineer (tse) the site checked the blue fiber cable connection to the ssc. The site found no issues. With the assistance of the tse the site replaced the backup camera cable with the original camera cable; however, the issue persisted. The patient had been administered anesthesia when the surgeon decided to abort the planned surgical procedure and reschedule it for a later date. | Conclusions: the investigation conducted by the field service engineer (fse) and review of the system error logs determined that the vision issue experienced by the customer was associated with the double camera controller (doco). The doco refers to the two camera controller units (ccu) that control the acquisition and processing of the image from the camera. One ccu is assigned to the surgeon's right sid image (right ccu) and one to his left (left ccu). The system was repaired by replacing the affected doco. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DAVINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3070649 | 3032641 | 2929155 | 2955842201301045 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2013-02-27 | 2013-03-29 | Yes | Yes | Health Professional | IS3000 A60P3 | Yes | Yes | 2013-02-27 | Yes | No | Reuse | 2013
| 2013-03-28 | No Answer Provided |  | While taking the da vinci s permanent cautery hook out of the trocar during a robotic cholecystectomy, the hook would not come out of the trocar. The hook came out with the trocar removal with no harm to the patient. The hook has a little nub on the side of the instrument that didn't allow it to come out. The hospital needs reimbursement from company for unused usages. Device #1is this a laboratory device or laboratory test?no. |  | DA VINCI | SYSTEM SURGICAL COMPUTER CONTR | 3076541 | 3049359 | 2945873 | 3049359 | 1 | NAY | User Facility |  |  | Initial | 2013-03-29 | 2013-03-29 | No |  | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
42030901 | 2013-02-22 | Malfunction |  | It was reported that during a da vinci si gastrostomy procedure, the customer noted that one of the cables broke on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Engineering also found indentations on the edge of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3053023 | 3026272 | 2922786 | 2955842201301004 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-28 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-20 | Malfunction |  | It was reported that during a da vinci si gastrectomy procedure a wire at the wrist of the prograsp forceps instrument broke. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Additional finding were scratches on the main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Damage may be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3053042 | 3026317 | 2922831 | 2955842201301006 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-28 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-21 | Malfunction |  | It was reported during a da vinci s prostatectomy procedure the fenestrated bipolar forceps instrument was noted to have a separated cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable is separated. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. Engineering also found scratches on the surface of the distal pulley. Additionally there were various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3053085 | 3026327 | 2922841 | 2955842201301005 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-28 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-25 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure a wire on the mega suturecut needle driver instrument was sticking out at the distal end. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported wire sticking out at distal end. However for clarification, the nonconformance was a broken grip cable. The broken grip cable stuck out at the wrist. The idler pulley spun freely and exhibited pulley face and rim damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3053111 | 3026352 | 2922866 | 2955842201301007 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-28 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-02-21 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the permanent cautery hook instrument exhibited a damaged and exposed wire at the hook end. Customer also reported discoloration to the top of the shaft (brown part) of the tip of the instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did exhibit some damage at the proximal clevis area. Engineering also observed indentations marks around the edges and some pieces of plastic breaking off. Engineering concluded that damage was likely due to mishandling due to visible signs of damage found around the proximal clevis. Site also made reference to some instrument discoloration located at the distal end of the instrument. For clarification, engineering observed moisture inside the instrument. Usually after the instruments are autoclaved, it is possible that pressure or moisture can be trapped inside the proximal clevis area. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3053235 | 3026459 | 2922973 | 2955842201301008 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-02-26 | 2013-03-28 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-02-22 | Malfunction |  | It was reported that during a da vinci s colorectal procedure a wire was coming out of the small grasping retractor instrument. There were no missing pieces or fallen pieces reported. The instrument was changed and the planned surgical procedure was completed. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The pitch down cable was frayed at the distal clevis hub. The frayed strands stuck out at the wrist. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3053319 | 3026536 | 2923050 | 2955842201301009 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-26 | 2013-03-28 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure wires were sticking out at the end of the prograsp forceps instrument. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. The instrument was found with a frayed cable at the distal idler pulley. The frayed cable section was approximately 0. 07 in length. No damage was found on the pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3053323 | 3026539 | 2923053 | 2955842201301011 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-28 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure a broken wire on the mega suturecut needle driver instrument was identified during set up and therefore, was not used in the case. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at wrist were not damaged. Additional findings were scratches on the pulley. There were scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | INSTRUMENT | 3053320 | 3026551 | 2923065 | 2955842201301010 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-28 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42004906 | 2013-01-18 | Malfunction |  | It was reported that during a da vinci si right middle lobectomy procedure, the surgical staff noticed that the graphite housing of the cadiere forceps instrument split. The surgical staff stated that the instrument was stuck laterally and the trocar had to be removed from the patient to get the instrument out of the trocar. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was not returned to intuitive surgical inc. (isi) for evaluation. However, the site provided photographic images of the reported instrument involved in this complaint to isi for evaluation. Based on the photographic images provided by the site, engineering evaluation noted that the proximal clevis of the instrument had broken off from the main tube. The main tube was shown with fibers sticking out at the distal end. Engineering evaluation concluded that the damage was likely due to excess loading force applied perpendicular to the clevis and the main tube. Engineering evaluation also noted that the instrument was likely misused or mishandled. On (b)(4) 2013, isi contacted the initial reporter of this complaint and obtained additional information regarding the reported event. The reporter stated that no fragments from the instrument fell into the patient. The reporter also explained that the housing of the instrument split which prevented the surgical staff from removing the instrument from the patient. The reporter stated that the cannula and instrument were removed simultaneously to allow for removal of the instrument from the patient. According to the reporter, the planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3049677 | 3023422 | 2919936 | 2955842201300970 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-03-04 | 2013-03-27 | No | Yes | Health Professional | 42004906 | No | Yes | 2013-03-04 | No | No | Reuse | 2013
42017910 | 2013-02-25 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, during set up it was noted that a piece of the tip hot shears monopolar curved scissors instrument was broken off, the instrument was not used on the case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with one of the blade tips with a small piece broken off. Engineering concluded that damage was likely due to mishandling. No other damage was found. Instrument showed 0 life remaining. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | HOT SHEARS TM MONOPOLAR CURRVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3049887 | 3023549 | 2920063 | 2955842201300971 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-25 | 2013-03-27 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-18 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, while finishing suturing mesh, the customer noticed a wire protruding from distal tip of the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional finding not reported by the site was surface scratches on the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050074 | 3023651 | 2905594 | 2955842201300972 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-25 | 2013-03-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-11 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, while suturing vaginal cuff, the customer noticed wires protruding from distal tip of the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment sticks out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050154 | 3023702 | 2920216 | 2955842201300973 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-25 | 2013-03-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
420003 | 2013-02-25 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the clips kept dropping from the small clip applier instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; a follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. Isi contacted initial reporter clinical sales representative (csr). She indicated that she visited the account and she noted that during insertion of the clip, the masters were squeezed prematurely. The csr was present during the following case; she confirmed that the surgeon was able to use the small clip applier instrument with no further issues. The endowrist instruments instructions for use (ifu) specifically states: caution: if excessive motion of the wrist or grips occurs, the clip should be removed and a new clip inserted. This is because if the wrist moves significantly, the grip action may be affected and the clip could become slightly compressed. This could result in the clip falling out of the grip and into the patient during insertion or engagement. If clip is not flush with the tip of the clip applier, the clip should be removed, grips opened and a new clip inserted. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050681 | 3024143 | 2920657 | 2955842201300974 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-02-25 | 2013-03-27 | No | Yes | Health Professional | 420003 | No | No | 2013-02-25 | No | No | Unkown | 2013
IS3000 A60P6 | 2013-02-21 | Injury | Required Intervention | It was reported that during a da vinci si partial nephrectomy procedure, the patient side manipulator arm 1 behaved in a sluggish manner. The site contacted isi technical support engineering and indicated that they troubleshot the issue by installing different instruments onto to the affected psm; however, the issue persisted. The or nurse indicated to the tse that they were not very deep into the patient and that they could not observe any interference with the instrument or psm. The site declined further trouble-shooting assistance and indicated that they were going to proceed with the procedure. On (b)(4) 2013, the isi clinical sales representative followed up with the surgeon concerning the reported event. The csr indicated that the surgeon indicated to him that the patient's vena cava was damaged earlier in the procedure and during repair of the patient's anatomy, he experienced articulation issues with the instruments installed on the 1 psm and the 2 psm. The csr indicated that the surgical staff troubleshot the issue by reseating the instruments and the sterile adapters. However, the installed instruments intermittently failed to articulate properly. The csr indicated that the surgeon reported that due to the patient's bleeding he was unable to troubleshoot the issue further. The csr indicated that the surgeon indicated to him that damage to the patient's vena cava was repaired using the da vinci si surgical system and that the planned surgical procedure was also completed with the system. The csr indicated that the site reported that they observed that one of the inner axis insertion posts located on the 1 psm was protruding out and was not flush. No other details concerning this observation were provided to the csr by the site. | The investigation conducted by the field service engineer (fse) determined that the articulation issue experienced by the site was related to the patient side manipulator arm 1. The psm is an instrument arm located on the patient side cart, which provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found that the wrist output assembly post heights of axis 4 through axis 7 on the psm were out of specification, thus causing the issue experienced by the site. The psm was repaired by replacing the affected output assembly posts. Engineering evaluation also found that during functional testing of the psm, the psm failed an in-house brake drop test and axis 5 and axis 6 potentiometers had jitter. The main brakes and all of the potentiometers were also replaced in the psm. On (b)(4) 2013, isi contacted the isi clinical sales representative who initially reported the patient injury. The csr indicated that the surgeon indicated to him that damage to the patient's vena cava was not caused by a malfunction of the da vinci si surgical system, instruments and/or accessories; but rather he (surgeon) accidentally nicked the patient's vena cava. Per the csr, the surgeon stated i accidentally got into the patient's vena cava. Per the csr, the surgeon indicated that the articulation issues with the psms occurred while attempting to repair the patient's anatomy. The csr indicated that the surgeon also reported to him that the patient is recovering well and that the patient has not returned to the hospital due to experiencing any post surgical complications as a result of the reported event. As of (b)(4) 2013, there have been no reported recurrences of this failure mode at this hospital. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3050728 | 3024191 | 2920705 | 2955842201300975 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-25 | 2013-03-27 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-26 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that wires were hanging out at side of the fenestrated bipolar forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3050850 | 3024293 | 2920807 | 2955842201300976 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-26 | 2013-03-27 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
420179 | 2013-02-11 | Injury | Other | It was reported that during a da vinci si prostatectomy procedure, when the surgeon activated bipolar cautery energy from the footpedal on the surgeon side cart, cautery energy from the monopolar curved scissors (mcs) instrument was activated causing a burn to the patient's bowel. | On (b)(4) 2013 an inspection of the site's da vinci si system conducted by an isi field service engineer (fse) was unable to replicate the cautery issue experienced by the site. Functional testing performed on the site's system by the fse found that the site's da vinci si surgical system functioned within specification. On (b)(4) 2013 isi contacted the isi clinical sales representative (csr) who reported this event. The csr indicated that she met with the operating room director and operating room scrub technician at the site on (b)(4) 2013 and the operating room scrub technician indicated to her that automatic activation of cautery energy was due to the surgical staff had incorrectly connected the electrical surgical unit (esu) cables into the improper receptacles on the esu. The scrub technician indicated that the esu connections were corrected, and the issue was resolved. The planned surgical procedure was completed. A general surgeon at the hospital was consulted to access the damage to the patient's bowel burn. The general surgeon's assesment of the burn to the patient's bowel determined that no repair of the affected area was required. On (b)(4) 2013, isi contacted the surgeon who performed the surgical procedure. The surgeon indicated that while he was retracting the patient's bowel using the maryland bipolar forceps instrument, he heard the alarm from the esu and immediately stopped using the maryland bipolar forceps instrument. The surgeon indicated that he had not touched the footpedal and that the instrument activated when he had not pressed the footpedal. The surgeon indicated that the patient sustained a small burn injury to the sigmoid colon and that he attempted to consult with a general surgeon concerning the patient's injury; however, there was no general surgeon available at that time. The surgeon indicated that he assessed the damage to the patient's bowel and he determined that the damage did not require any repair. The surgeon indicated that the patient remained in the hospital an additional day for monitoring and that the patient was discharged from the hospital. The surgeon indicated that the patient is recovering well and has not returned to the hospital due to experiencing any post-surgical complications as a result of the reported event. Intuitive surgical's instruments and accessories user manual precautions and warnings indicate: caution: please refer to the individual esu manufacturer's user manual for operating instructions. Warning: do not use monopolar instruments with a bipolar source output as this may cause damage to the instrument and harm to the patient or medical personnel. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | MARYLAND BIPLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3050863 | 3024317 | 2920831 | 2955842201300977 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-25 | 2013-03-27 | Yes | No | Health Professional | 420179 | No | No | 2013-02-25 | No | No | Reuse | 2013
42030901 | 2013-02-26 | Malfunction |  | It was reported during a da vinci s surgical procedure the mega suturecut needle driver instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was broken. The one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050926 | 3024362 | 2920876 | 2955842201300980 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-25 | Malfunction |  | It was reported that after a da vinci si surgical procedure, a loose wire at the wrist of the fenestrated bipolar forceps instrument was found. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the customer reported loose wire, however, for clarification the damaged instrument component was a frayed pitch cable. The instrument was found with a frayed pitch cable at distal clevis hub. No damage was found at the clevis. No other cable damage was found. Electrical continuity test passed. An additional finding was a scratch on the pulley. There were scratches on the surface of the distal pulley. Scratches on the maintube were also found. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Damage is likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3050924 | 3024374 | 2920888 | 2955842201300979 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-25 | 2013-03-27 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-26 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted a broken cable on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was found to be broken and sticks out at the wrist. The idler pulley exhibited pulley face and rim damage. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050925 | 3024375 | 2920889 | 2955842201300978 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-26 | 2013-03-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-26 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure the mega suturecut needle driver instrument was noted to have broken wires. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to duplicate the customer reported complaint. Engineering found the one grip cable was frayed at the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at the wrist were not damaged. The idler pulley was not damaged and still rotated freely. Engineering also found damage on the instrument blade. There were several indentations the midpoint of the blade. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050929 | 3024376 | 2920890 | 2955842201300983 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-26 | 2013-03-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-02-27 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure the monopolar cautery instrument insulation was noted to being scraped and the bottom of the instrument was cracked. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3050936 | 3024380 | 2920905 | 2955842201300987 | 1 | NAY | Manufacturer | User facility | BIOMEDICAL ENGINEER | Initial | 2013-02-27 | 2013-03-27 | No | Yes | Health Professional | 42014204 | No | Yes | 2013-02-27 | No | No | Unkown | 2013
42029602 | 2013-02-20 | Malfunction |  | It was reported during a da vinci si total benign hysterectomy procedure the large suturecut needle driver instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050945 | 3024384 | 2920898 | 2955842201300994 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2013-05-30 | 2013-03-27 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | Yes | 2013-05-30 | Yes | No | Reuse | 2013
42009308 | 2013-02-27 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure the prograsp forceps instrument was noted to have a broken cable when it was removed from the sterile pack; it was not used in the case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken cable. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering also found scratches on the surface of the distal pulley. Engineering also found various scratch marks on the main tube at the distal end showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 3050947 | 3024385 | 2920899 | 2955842201300995 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-03-27 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-25 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, a wire was loose at the wrist of the prograsp forceps instrument. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a loose wire. One grip close cable was loose at the distal idlers. The idler pulley spun freely and was not damaged. The clamping pulley screws were found to be secured but a grip cable was broken by the clamping pulley which caused the loose wire to show at the idler. Additional findings were corroded clamping pulleys. Multiple clamping pulleys exhibited corrosion around the bottom of the pulley. The failure may be cleaning related. Lastly, scratches on maintube were found. The distal end of main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Damage is likely due to mishandling. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050928 | 3024386 | 2920900 | 2955842201300982 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-25 | 2013-03-27 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-02-26 | Malfunction |  | It was reported during a da vinci s surgical procedure the grasping retractor instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was broken. The one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR 8MM INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050930 | 3024387 | 2920901 | 2955842201300981 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-27 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-01-25 | Malfunction |  | It was reported during a da vinci si benign hysterectomy procedure the surgeon had difficulty getting the tenaculum forceps instrument to grasp and hold tissue. Another tenaculum forceps was used to complete the case. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050932 | 3024389 | 2920903 | 2955842201300985 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-26 | 2013-03-27 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | Yes | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-02-27 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure the monopolar cautery instrument it was noted that the insulation was being scraped as the instrument was being inserted into the da vinci system. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3050904 | 3024391 | 2920905 | 2955842201300986 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2013-02-27 | 2013-03-27 | No | Yes | Health Professional | 42014204 | No | Yes | 2013-02-27 | No | No | Unkown | 2013
42018304 | 2013-02-11 | Malfunction |  | It was reported that during reprocessing that the hook on the permanent cautery hook instrument was broken. There were no missing pieces or fallen pieces reported. | The instrument was returned and evaluated. The instrument was found with one side of the distal clevis broken off at the distal end and a missing conductor wire cap. The cables were not damaged but the damage found at distal end could be linked to the broken distal clevis and the missing conductor wire cap. Damage was likely due to excessive side loading at the distal tip or other type of mishandling. Additionally there was discoloration at distal end, main tube, and housing. All three components of instruments look faded. The discoloration is likely due to improper cleaning. Electrical continuity test passed. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3050906 | 3024394 | 2920908 | 2955842201300989 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-25 | 2013-03-27 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42014504 | 2013-02-27 | Malfunction |  | It was reported during a da vinci si surgical procedure the debakey forceps instrument insulation was noted to be scraping as it was being inserted into the da vinci system. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | DEBAKEY FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050908 | 3024396 | 2920960 | 2955842201300990 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2013-02-27 | 2013-03-27 | No | Yes | Health Professional | 42014504 | No | Yes | 2013-02-27 | No | No | Unkown | 2013
42009308 | 2013-02-27 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure the prograsp forceps instrument was noted to have broken cables at the tip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken cable. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering also found various scratch marks on the main tube at the distal end showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050940 | 3024398 | 2920912 | 2955842201300991 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-03-27 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-25 | Malfunction |  | It was reported during a da vinci si total benign hysterectomy procedure the large suturecut needle driver instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050943 | 3024400 | 2920914 | 2955842201300993 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2013-05-30 | 2013-03-27 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | Yes | 2013-05-30 | Yes | No | Reuse | 2013
42020505 | 2013-02-20 | Malfunction |  | It was reported during a da vinci si total benign hysterectomy procedure, the fenestrated bipolar forceps instrument was noted to have broken cables at the tip of the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was broken. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. Engineering also found scratches on the surface of the distal pulley. Additionally, engineering found various scratch marks on the main tube at the distal end showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3050950 | 3024402 | 2920916 | 2955842201300992 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-03-27 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-28 | Malfunction |  | It was reported during a da vinci si total benign hysterectomy procedure the prograsp forceps instrument was noted to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a frayed cable at the distal idler pulley. The frayed cable section was approximately 0. 09 in length. No damage was found on the pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050949 | 3024406 | 2920920 | 2955842201300998 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-27 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-26 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, it was noted that cable of the mega suture cut needle driver instrument snapped during set-up. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The cable segment sticks out at wrist. Outer surface of the distal idler pulley also exhibited light scratches. Additional observation not reported by the site was derailed cable. The cable derailment suggested that the grip close cable could have been derailed prior to the cable breakage. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050972 | 3024407 | 2920921 | 2955842201300996 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-28 | Malfunction |  | It was reported during a da vinci si surgical procedure the mega suturecut instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was broken. The one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found scratches on the surface of the distal pulley. Engineering also found various scratch marks on the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050970 | 3024426 | 2920940 | 2955842201300997 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-22 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical fenestrated bipolar forceps instrument jaws were not opening. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. The grips opened and closed fine when input discs were manually rotated. The drive cables were not damaged. Engineering also found various scratch marks at the distal end showing light material removal. The scratches were. 070 -. 120 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3050952 | 3024427 | 2908477 | 2955842201300999 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-27 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-26 | Malfunction |  | It was reported that during reprocessing the da vinci si surgical prograsp forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering found the pitch cable was found to be frayed at the distal end. No other damage was found at the distal end. Engineering also found various scratch marks on the main tube a the distal end showing light material removal. The scratches were short in length and not aligned with the tube axis, suggesting they may have been caused by instrument collisions or rough handling. No other damage was found. Instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050971 | 3024428 | 2920942 | 2955842201301000 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-27 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-26 | Malfunction |  | It was reported that during reprocessing the da vinci si fenestrated bipolar forceps instrument was noted to have a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken wire. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. Engineering also found and indentation on the edge of the distal pulley. There were various scratch marks found on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3050974 | 3024429 | 2920943 | 2955842201301002 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-27 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-26 | Malfunction |  | It was reported that during reprocessing the da vinci si prograsp forceps instrument was noted to have a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken wire. The pitch cable was broken at the distal end. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at wrist were not damaged. Engineering also found the back idler pulley was corroded. The set of four back idler pulleys on the right inside of the housing exhibited some rust. The idler pulleys were not seized and still rotated. The evidence was not conclusive, but damage was likely due to improper cleaning. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050975 | 3024430 | 2920944 | 2955842201301003 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-27 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-02-25 | Malfunction |  | It was reported that prior to starting a da vinci si procedure a wire broke near the tip of the small clip applier instrument. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken wire. However for clarification, the nonconformance was a broken grip cable. The broken grip cable stuck out at the wrist. The idler pulley spun freely and exhibited pulley face and rim damage. An additional observation not reported was abnormal wear on the main tube. The surface of the main tube was discolored and splintering off. The housing was removed and the clamping pulley showed signs of excessive residue residing around the grip cable. The damage was likely due to improper cleaning process. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse ev ent. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050956 | 3024436 | 2920950 | 2955842201300984 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-25 | 2013-03-27 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42014504 | 2013-02-27 | Malfunction |  | It was reported during a da vinci si surgical procedure the debakey forceps instrument insulation was noted to be scraping as it was being inserted into the da vinci system. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | DEBAKEY FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050991 | 3024446 | 2920960 | 2955842201300988 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | Initial | 2013-02-27 | 2013-03-27 | No | Yes | Health Professional | 42014504 | No | Yes | 2013-02-27 | No | No | Unkown | 2013
42000606 | 2013-02-28 | Malfunction |  | The instrument was returned, no further information was provided. No patient harm, adverse outcome or injury was reported. | The instrument was returned, no information was provided. Engineering evaluated the instrument and found the grip cable was broken at the wrist. The idler pulley spun freely and exhibited pulley face and rim damage. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3050963 | 3024459 | 2920973 | 2955842201301001 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-28 | 2013-03-27 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-28 | Device Not Retu | No | Reuse | 2013
| 2013-03-15 | Malfunction |  | Vessel sealer handpiece malfunctioned with error code. Blade exposed on vision cart during procedure. Two vessel sealers were opened with the same error codes and were unable to be used. What was the original intended procedure?robotic assisted laparoscopic supracervical hysterectomy andbilateral salpingo-oophorectomy. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST | INSTRUMENT SURGICAL COMPUTER C | 3062052 | 3034411 | 2930925 | 3034411 | 1 | NAY | User Facility |  |  | Initial | 2013-03-27 | 2013-03-27 | No | Yes | Invalid Data |  | Yes | No Answer Provided |  |  |  |  | 2013
42018306 | 2012-12-26 | Malfunction |  | It was reported that prior to starting a davinci si prostatectomy procedure, the customer noted that the wire was exposed on the permanent cautery hook instrument. No pt harm, adverse outcome or injury was reported. | The instrument has not been returned for eval; therefore, the root cause of the customer reported failure mode cannot be determined. A f/u mdr will be submitted if the instrument is returned (post engineering eval) or if additional info is received. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3109130 | 3068595 | 2965129 | 2955842201300714 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2012-12-28 | 2013-03-27 | No | Yes | Health Professional | 42018306 | No | No | 2012-12-28 | Device Not Retu | No | Unkown | 2013
42019408 | 2013-02-14 | Malfunction |  | It was reported that during a da vinci surgical procedure a cable on the mega needle driver broke. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported broken cable; however for clarification, the nonconformance was a frayed grip cable. The frayed cable sections were in various lengths sticking out at the instruments wrist. The idler pulley exhibited rim and face damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3047511 | 3021342 | 2917856 | 2955842201300965 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-26 | 2013-03-26 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-14 | Malfunction |  | It was observed during central processing that the tips of the large needle driver instrument were bent. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results did not confirm the issue of bent tips. The instrument was found with one grip tip broken at the distal clevis at grip base. The broken grip tip was missing. Engineering evaluation also found the distal idler pulley to be bent on the edge. The pulley was able to rotate freely. No additional damage was found on the distal clevis or on any cables. Engineering evaluation concluded that the damage was likely due to excess force applied normal to the jaw. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3048389 | 3022005 | 2918519 | 2955842201300967 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-27 | 2013-03-26 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-27 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-18 | Malfunction |  | It was reported that halfway through a da vinci si procedure, the surgeon lost control of the prograsp forcep instrument. Upon inspection it was noticed that one of the wires has snapped. No adverse event was reported as a result of the alleged issue. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3048473 | 3022057 | 2918571 | 2955842201300966 | 1 | NAY | Manufacturer | DistributorHealth Professional | Nurse | Initial | 2013-02-25 | 2013-03-26 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a cable snapped at distal end of the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a cable snapped. One grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the instruments wrist. The other cables at the wrist were not damaged. An additional finding were scratches on the pulley. There were some scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3048495 | 3022094 | 2912616 | 2955842201300968 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-25 | 2013-03-26 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-19 | Malfunction |  | It was reported that during a da vinci total benign hysterectomy surgical procedure, a wire broke at the tip of the large suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed one grip cable was broken at the distal clevis. The broken segment stuck out approximately. 1500 at the instrument wrist. The idler pulley spun freely but also showed some wear. The other cables at the wrist were not damaged. An additional observation found were corroded back idler pulleys. The back idler pulleys exhibited corrosion in between each pulley. The idler pulleys were not seized and still rotated. Damage was likely due to improper cleaning. No other damage was found. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3048552 | 3022158 | 2918672 | 2955842201300969 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-25 | 2013-03-26 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
41029902 | 2013-02-25 | Malfunction |  | It was reported that during a da vinci si colon resection procedure, the surgical staff noticed that the tip of the suction/irrigator instrument fell off and also fell into the patient. The tip of the instrument was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue that the tip of the instrument fell off. The instrument was found with snake wrist dislodged from the distal end. The missing tip that fell into patient was retrieved by site but was not returned with instrument. Engineering evaluation concluded that the damage was likely be due to excessive side loading or other type of mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, intuitive surgical inc. (isi) contacted the initial reporter of this complaint and verified that no patients were harmed or injured because of the reported issue. The initial reporter stated that no x-ray or ct scan was performed as result of the reported issue. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | SUCTIONIRRIGATOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3045679 | 3019663 | 2916177 | 2955842201300963 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-03-01 | 2013-03-25 | No | Yes | Health Professional | 41029902 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | Yes | Initial | 2013
42009308 | 2012-04-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the prograsp forceps instrument did not work and had no power. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as a visual inspection of the instrument found no broken components on the exterior of the instrument. The instrument is not an energy instrument; therefore, there is no power that runs through the instrument. Engineering evaluation also found that the distal end of main tube has various scratch marks with light material removal. The scratches are. 075 -. 130 in length and were not aligned with the tube axis. Engineering concluded that the damage to the instrument's main tube was likely due to mishandling/misuse. No other was damage found. The instruments and accessories user manual specifically states: general precautions and warnings, handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The damage to the instrument's main tube found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3046426 | 3020323 | 2916837 | 2955842201300964 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-17 | 2013-03-25 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
| 2012-10-19 | Malfunction |  | Upon closing the patient, it was reported that a cap from the monopolar scissors was missing. Patient required a return to the operating room (or) for removal after tip was located by a computed tomography (ct) scan. What was the original intended procedure?total laparscopic hysterectomy with davinci assist. |  | ENDOWRIST HOT SHEARS | SYSTEMSURGICAL | 3049042 | 3022625 | 2987380 | 3022625 | 1 | NAY | User Facility |  |  | Initial | 2013-03-25 | 2013-03-25 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2013
| 2012-02-14 | Injury | Required InterventionHospitalizationDisability | On (b)(6) 2012, i had a da vinci assisted lavh, being released the following day on (b)(6) 2012. I felt as though i was recovering well for the first few days. Began having abdominal pain on (b)(6) 2012. Returned to surgeon on (b)(6) 2012 because i was not feeling well. Put on antibiotic and sent home. On (b)(6) 2012, went to er and was admitted because it was discovered that my right ureter had suffered a thermal injury during my hysterectomy and my continued decline the days prior were due to the injured ureter leaking urine into my abdominal cavity. After two failed attempts at treating the injury conservatively, i had to be opened up, reoperated on, and my bladder reconstructed to reach my now shorter ureter. |  | DA VINCI | DA VINCI ROBOT | 3055366 | 3027517 | 3015645 | MW5029520 | 1 | NAY | Voluntary |  | Patient | Initial | 2013-03-25 | 2013-03-25 | Yes | No | Service Personnel |  | No | No |  |  |  |  | 2013
| 2013-03-22 | Malfunction |  | Da vinci fenestrated bipolar cautery stopped working during the procedure. Despite having 8 uses left. Cautery cord and instrument were changed and cauterization worked. What was the original intended procedure?total hysterectomy. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST | SYSTEM SURGICAL COMPUTER CONTR | 3062046 | 3034433 | 2930947 | 3034433 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-03-25 | 2013-03-25 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
42020303 | 2012-11-20 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the large clip applier instrument would not open or close. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the grip cables inside the housing were broken, thus causing the issue experienced by the customer. Engineering evaluation also found that prior to removal of the housing, the luer plate with the flush tube still intact was dislodged at the back of the housing. The surface of main tube was also no longer smooth. Engineering concluded that damage to the luer plate and main tube was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings; handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the damage to the instrument's main tube found during failure analysis could likely cause or contribute to an adverse event if the failure mode were to recur. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | LARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3044510 | 3018652 | 2915166 | 2955842201300932 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-12 | 2013-03-23 | No | Yes | Health Professional | 42020303 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-11-28 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tips on the pk dissecting forceps instrument did not meet. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was bent, causing side to side misalignment of the grips. There was a 0. 022 offset at the grip tips. The grip tips did not meet together once the grips were closed. Engineering concluded that the damage was likely due to overloading at the tip or excessive force applied to the jaw. The instrument passed electrical continuity testing. Engineering evaluation also found that the pitch cable at the distal clevis hub was frayed. No other damage was found at the clevis. The instruments and accessories user manual specifically states: general precautions and warnings,, o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself does not constitute a mdr reportable event; however; damage to the instrument's pitch cable found during failure analysis could likely cause or contribute to an adverse event if the failure mode were to recur. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3044511 | 3018686 | 2915200 | 2955842201300933 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-12 | 2013-03-23 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42020906 | 2012-11-28 | Malfunction |  | It was reported that the wiring on the on the megasuturecut needle driver instrument was loose. No missing or fallen pieces were reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode with the following clarification that the loose wire observed by the customer was a broken grip cable. One grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the instrument's broken grip cable found during failure analysis could likely cause or contribute to an adverse event if the failure mode were to recur. | SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3044512 | 3018687 | 2826715 | 2955842201300962 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-27 | 2013-03-23 | No | Yes | Health Professional | 42020906 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-20 | Malfunction |  | It was reported that during a da vinci s ovarian cystectomy procedure, the customer noted a separated wire at the tip of the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that both pitch cables were found to be broken at the proximal clevis at the distal end. Both crimps were still installed in the clevis, but the cables stick out at wrist. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3042979 | 3017398 | 2913912 | 2955842201300940 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-22 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-02-13 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff noticed that a cable was sticking out of the tenaculum forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of a cable sticking out. The instrument was found to have one grip cable broken at the distal clevis. The broken cable segment was observed to be sticking out approximately. 7985 at the instrument's wrist. The idler pulley was able to spin freely and did not exhibit any damage. The other cables of the instrument's wrist did not exhibit any damage. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken hook tip and main tube scratch issues if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3043053 | 3017454 | 2913968 | 2955842201300941 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-22 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-20 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a frayed wire on the large needle driver instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The frayed segment was approximately 0. 04 in length. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3043049 | 3017461 | 2913975 | 2955842201300942 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-22 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42018305 | 2013-02-20 | Injury | Required Intervention | It was reported that during a da vinci si low anterior resection procedure, the hook at the tip of the permanent cautery hook instrument was pulled off and fell into the patient. The fragment was reportedly retrieved and the planned surgical procedure was completed. | \the instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a pulled off hook tip. The instrument was returned to isi with the hook detached from the distal end. The hook was found to be fractured within the yaw pulley, at the cross section of one of the holes in the shank. \engineering evaluation also found a broken ceramic sleeve and tube damage. The entire ceramic sleeve was missing from the yaw pulley. In addition, the distal end of the instrument's main tube was observed to have various scratch marks with light material removal. The scratches were. 070 -. 160 in length and not aligned with the instrument's tube axis. Engineering evaluation concluded that the damages were likely due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, isi contacted the initial reporter of this complaint and obtained additional information. The initial reporter indicated that a c-arm was used to help locate the tip fragment. The fragment was reportedly retrieved through one of the current port incisions. The initial reported stated that no extra surgical procedures were performed. According to the initial reporter, the planned lar procedure was completed as planned and no post-operative complications occurred. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3043066 | 3017496 | 2914010 | 2955842201300943 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-20 | 2013-03-22 | Yes | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-02-20 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was noted that black housing was torn, it was identified during set up, the thoracic grasper instrument was not used in case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the main insulation tube exhibited damage along the distal end. The distal end of the main tube exhibited various gauges marks with some material removed. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. Engineering also observed that the main tube surface was still smooth, but it was discolored with pale white marks. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3043101 | 3017527 | 2908097 | 2955842201300944 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-22 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-02-20 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was noted that black housing was torn, it was identified during set up, the thoracic grasper instrument was not used in case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the main insulation tube appeared to have scratches and gouge marks, with materials removed at the distal end. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3043166 | 3017590 | 2908097 | 2955842201300945 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-22 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42021504 | 2013-02-21 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the surgical staff noticed that the tip of the cardiac probe grasper instrument was broken off. There were no missing or fallen pieces reported. The reported issue was noticed while the instrument was in the sterile tray. The instrument was reportedly not used in the planned surgical procedure. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of that the tip of the instrument was broken off. The tip of one of the instrument's grips was found to be broken close to midpoint. The broken piece was not returned by the customer. The broken plane of the grip was straight with a smooth edge. No other damage found around the distal area of the instrument. Engineering evaluation was unable to determine how the instrument tip broke off. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken hook tip and main tube scratch issues if to recur could cause or contribute to an adverse event. | CARDIAC PROBE GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3043466 | 3017803 | 2914317 | 2955842201300947 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-21 | 2013-03-22 | No | Yes | Health Professional | 42021504 | Device Returned To Manufacturer | Yes | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-02-20 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, during set up, it was noted that the tips of the small clip applier instrument were bent, the instrument was not used in case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated Per the customer reported complaint engineering observed that the instrument was found with one grip tip bent and the other grip was broken at distal clevis at grip base The grip tip was bent at approximately a 123 angle from midpoint The other grip tip was missing a 0 1460 piece off the grip base No other damage found on the distal clevis or the cables Engineering concluded that damage was likely due to mishandling The customer reported complaint does not itself constitute a reportable event however the reported malfunction if to reoccur could cause or contribute to an adverse event This report does not admit that the report or information submitted under this report constitutes an admission that the device intuitive surgical or intuitive surgical employees caused or contributed to the reportable event The endowrist instruments instructions for use ifu specifically states general precautions and warnings handle instruments with care Avoid mechanical shock or stress that can cause damage to the instruments Do not use an instrument to clean debris from another instrument intraoperatively This may result in damage to the instruments or other unintended consequences such as disconnection of the instrument tip | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3043468 | 3017805 | 2914319 | 2955842201300946 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-22 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-21 | Malfunction |  | It was observed during central processing that the fenestrated bipolar forceps instrument had a broken cable. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of a broken cable. The instrument was found to have a broken pitch down cable at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables of the instrument's wrist did not exhibit any damage. Engineering evaluation also found various scratch marks with light material removal on the instrument's main tube. The scratches were. 125 -. 275 in length and not aligned with the tube axis. Engineering evaluation concluded that the scratches were likely caused by mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken hook tip and main tube scratch issues if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3043540 | 3017825 | 2914339 | 2955842201300948 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-22 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-20 | Malfunction |  | The prograsp forceps instrument was returned without a reported failure mode. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that main insulation tube was damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3043660 | 3017880 | 2914394 | 2955842201300949 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-02-20 | 2013-03-22 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-02-22 | Malfunction |  | It was reported that during reprocessing a wire was sticking out at the distil tip of the tenaculum forceps instrument. There were no missing pieces or fallen pieces reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported complaint. Based on visual inspection, there were no broken wires. Additional observation not reported by site were scratches on the maintube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. The damage may be due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3043855 | 3018020 | 2914534 | 2955842201300950 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-22 | 2013-03-22 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-02-22 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-18 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that the prograsp forceps instrument had a broken wire. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of a broken wire. The instrument was found to have a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed cable segment was approximately 0. 04 in length. The other cables of the instrument's wrist did not exhibit any damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken hook tip and main tube scratch issues if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3044052 | 3018184 | 2914698 | 2955842201300951 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-21 | 2013-03-22 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42011006 | 2013-02-06 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the physical damage on tips of the precise bipolar forceps instrument. | The instrument was returned and evaluated. Engineering was unable to confirmed the reported failure mode. Visual inspection of the instrument showed that the grips were not bent or damaged. No char marks were observed on the yaw pulleys. Additional observations not reported by site were main insulation tube damage and recognition failure. Distal end of main tube exhibited various scratch marks with light material removal. The scratches were 0. 120 - 0. 215 in length and they were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Instrument was checked for recognition using an in-house da vinci si robot system. The da vinci robot system failed to recognize the instrument on multiple installation attempts. The pogo pins did not stick and were not contaminated. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3044130 | 3018265 | 2914779 | 2955842201300952 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-22 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
420183 | 2013-02-20 | Injury | Required Intervention | Event Description: --/-- It was reported that the permanent cautery hook instrument broke while in use during a lower anterior colon resection by the surgeon. The surgeon was able to remove the broken instrument tip from the patient. The instrument was replaced and the planned surgical procedure was completed. The surgeon performed an x-ray and verified that no fragments were left inside the patient. Site completed 2 cases after with no errors or issues. System is operating per intuitive surgical specifications. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- Upon additional review, it was determined that medwatch report number 2955842-2013-00953 submitted on (b)(4) 2013 was a duplicate of medwatch report number 2955842-2013-00943 also submitted on (b)(4) 2013. Please accept this record as request to retract medwatch report number 2955842-2013-00953 as the original medwatch report 2955842-2013-00943 was sent to report this event. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PERMANENT CAUTERY HOOK INSTRUMENT | Unkown | 3044157 | 3018311 | 2914010 | 2955842201300953 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2013-02-20 | 2013-03-22 | Yes | Yes | Health Professional | 420183 | No | No | 2013-02-20 | Device Not Retu | No |  | 2013
42017207 | 2013-02-15 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure, the surgical staff noticed that the maryland bipolar forceps instrument had a broken cable. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of a broken cable. The instrument was found to have one frayed grip open cable at the distal idler pulley. Frayed strands were observed to be sticking out at the instrument's wrist. The other cables of the instrument's wrist did not exhibit any damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken hook tip and main tube scratch issues if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3044183 | 3018334 | 2914848 | 2955842201300954 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-22 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-20 | Malfunction |  | It was reported that after a successfully during a da vinci si surgical procedure, the retaining pin in the hinge of the maryland bipolar forceps instrument tip was loose and detached from the intended postilion. No fragments in patients were identified. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3044197 | 3018351 | 2914865 | 2955842201300955 | 1 | NAY | Manufacturer | ForeignHealth Professional | Other | Initial | 2013-02-21 | 2013-03-22 | No | Yes | Health Professional | 42017207 | No | No | 2013-02-21 | No | No | Unkown | 2013
42805206 | 2013-02-22 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted that at the tip of the monopolar cautery single site instrument wires were coming out of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with bent prongs at the cuatery adapter, causing difficulty to install the accessory tip. Engineering concluded that the damage was likely due to mishandling while installing the accessory tip. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The single site instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MONOPOLAR CAUTERY SINGLESITEINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3044240 | 3018372 | 2914886 | 2955842201300957 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-22 | 2013-03-22 | No | Yes | Health Professional | 42805206 | Device Returned To Manufacturer | Yes | 2013-02-22 | Device Not Retu | No | Reuse | 2013
42000207 | 2013-02-21 | Malfunction |  | It was reported that during the da vinci s prostatectomy procedure, the surgical staff felt an abnormal stuck feeling during exchanging the instrument. The staff was unable to insert another instrument, they inspected that cannula, and they found a broken portion on the cannula. The staff exchanged it into a backup cannula, and the planned procedure was completed without problem. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that cannula was found with the shaft cracked at the tip, with two hairline cracks approximately 1. 4' in length. Engineering concluded that damage was likely due to excess force applied normal to the cannula tip. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. The endowrist instruments instructions for use (ifu) specifically states: inspection instructions warning: do not use a damaged cannula or reducer. Damaged cannulae and reducers may abrade the instrument shaft and generate particles that may fall inside the patient. Warning: cannula defects can be caused by dropping the cannula or handling it roughly. | CANNULA INSTRUMENT ACCESSORY | ENDOSCOPIC INSTRUMENT ACCESSOR | 3044263 | 3018383 | 2914897 | 2955842201300959 | 1 | NAY | Manufacturer | DistributorHealth Professional | Physician | Initial | 2013-02-22 | 2013-03-22 | No | Yes | Health Professional | 42000207 | Device Returned To Manufacturer | Yes | 2013-02-22 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-21 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the customer observed exposed cables on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was derailed at the distal idler pulley. The grips were still able to open and close, but their movement could not be precise. The cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. Fleet angle between proximal and distal pulleys could have contributed to the derailment. Additional observation not reported by the site was frayed grip cable. The derailed cable was also found to be frayed. The grip cable was frayed at the distal idler pulley. Frayed strands stick out at the wrist. Other cables at wrist were not damaged. Engineering also found scratches on the pulley's surface, where grip cable was found to be derailed. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3044232 | 3018388 | 2914902 | 2955842201300956 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-22 | 2013-03-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-22 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-21 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted a broken cable on the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3044252 | 3018405 | 2920916 | 2955842201300958 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-22 | 2013-03-22 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-22 | Device Not Retu | No | Reuse | 2013
420205 | 2013-02-20 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, while performing dissection to take down the patient's bladder, the surgeon punctured the patient's bladder. It was reported that the patient had an anatomical issue related to a redundant bladder. | On (b)(4) 2013, isi contacted the clinical sales representative (csr) who was present during the surgical procedure. The csr indicated that the surgeon was using a permanent cautery spatula instrument and a fenestrated bipolar forceps instrument to take down the patient's bladder when damage to the patient's bladder occurred; however, it is unknown which instrument caused the damage to the patient's bladder. The csr indicated that there was no malfunction of the site's da vinci si surgical system, instruments or accessories that caused or contributed to the injury sustained by the patient. The csr indicated that the patient had a challenging anatomy and due to the thickness of the patient's bladder, the surgeon experienced some difficulty delineating the patient's bladder planes, causing the surgeon to inadvertently puncture the patient's bladder. The csr indicated that the surgeon repaired the patient's bladder using the da vinci si surgical system and the affected area was repaired with sutures. The csr indicated that the planned surgical was completed and the surgeon has not reported that the patient suffered any post-surgical complications as a result of the reported event. Based on the information provided, it was determined that the da vinci surgical system, instruments and/or accessories likely did not malfunction in a way that caused or contributed to the patient's injury, but, rather due to the patient's challenging anatomy, which caused the surgeon to inadvertently puncture the patient's bladder. Medwatch report 2955842-2013-00961 has been submitted for the permanent cautery spatula instrument that was also being used when the reported event occurred. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | FENESTRATED BIPOLAR FORCEPS INRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3044325 | 3018441 | 2914955 | 2955842201300960 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-20 | 2013-03-22 | Yes | No | Health Professional | 420205 | No | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
420184 | 2013-02-20 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, while performing dissection to take down the patient's bladder, the surgeon punctured the patient's bladder. It was reported that the patient had an anatomical issue related to a redundant bladder. | On (b)(6) 2013, isi contacted the clinical sales representative (csr) who was present during the surgical procedure. The csr indicated that the surgeon was using a permanent cautery spatula instrument and a fenestrated bipolar forceps instrument to take down the patient's bladder when damage to the patient's bladder occurred; however, it is unknown which instrument caused the damage to the patient's bladder. The csr indicated that there was no malfunction of the site's da vinci si surgical system, instruments or accessories that caused or contributed to the injury sustained by the patient. The csr indicated that the patient had a challenging anatomy and due to the thickness of the patient's bladder, the surgeon experienced some difficulty delineating the patient's bladder planes, causing the surgeon to inadvertently puncture the patient's bladder. The csr indicated that the surgeon repaired the patient's bladder using the da vinci si surgical system and the affected area was repaired with sutures. The csr indicated that the planned surgical was completed and the surgeon has not reported that the patient suffered any post-surgical complications as a result of the reported event. Based on the information provided, it was determined that the da vinci surgical system, instruments and/or accessories likely did not malfunction in a way that caused or contributed to the patient's injury, but, rather due to the patient's challenging anatomy, which caused the surgeon to inadvertently puncture the patient's bladder. Medwatch report 2955842-2013-00960 has been submitted for the fenestrated bipolar forceps instrument that was also being used when reported event occurred. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3044324 | 3018486 | 2915000 | 2955842201300961 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-20 | 2013-03-22 | Yes | No | Health Professional | 420184 | No | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
420205 | 2013-02-27 | Malfunction |  | During a robotic assisted hysterectomy the physician stated that the bipolar forceps was "floppy" and he could not control the movement of the forcep. The forcep instrument was removed, cleaned and re-attached to the manipulator and re-test with no success. The physician also reported that the needle driver instrument was not manipulating correctly. The manufacturer representative replaced both robotic instruments and had the system rebooted. The surgeon was able to complete the procedure without additional failures. ======================manufacturer response for robotic fenestrated bipolar forceps, da vinci si surgical system (per site reporter). ======================manufacturer to retrieve and evaluate. ======================manufacturer response for robotic mega suturecut needle driver, da vinci s (per site reporter). ======================manufacture to retrieve. What was the original intended procedure?robotic assisted hysterectomy. |  | DA VINCI S ENDOWRIST | SYSTEMSURGICAL | 3047307 | 3021241 | 2917755 | 3021241 | 1 | NAY | User Facility |  |  | Initial | 2013-03-22 | 2013-03-22 | No | Yes | Physician | 420205 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
420205 | 2013-02-27 | Malfunction |  | During a robotic assisted hysterectomy the physician stated that the bipolar forceps was "floppy" and he could not control the movement of the forcep. The forcep instrument was removed, cleaned and re-attached to the manipulator and re-test with no success. The physician also reported that the needle driver instrument was not manipulating correctly. The manufacturer representative replaced both robotic instruments and had the system rebooted. The surgeon was able to complete the procedure without additional failures. ======================manufacturer response for robotic fenestrated bipolar forceps, da vinci si surgical system (per site reporter). ======================manufacturer to retrieve and evaluate. ======================manufacturer response for robotic mega suturecut needle driver, da vinci s (per site reporter). ======================manufacture to retrieve. What was the original intended procedure?robotic assisted hysterectomy. |  | DA VINCI S ENDOWRIST | SYSTEMSURGICAL | 3047307 | 3021241 | 2917755 | 3021241 | 1 | NAY | User Facility |  |  | Initial | 2013-03-22 | 2013-03-22 | No | Yes | Physician | 420205 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
| 2013-03-11 | Malfunction |  | When the surgeon tried to clean up the tip of the pk forcep during the case, a robotic-assisted laparoscopic supracervical hysterectomy, a wire of the pk forcep peeled off. Took x-ray of pt's abdomen, which was negative. |  | INTUITIVE SURGICAL | PK DISSECTING FORCEPS | 3059619 | 3031079 | 2927593 | 3031079 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-03-19 | 2013-03-22 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2013
42011006 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgical staff noticed that the precise bipolar forceps instrument had been inserted with a torn string. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a torn string. The instrument was found to have a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis reportedly did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. Engineering evaluation also found scratches on the surface of the distal pulley. | PRECISE BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3040909 | 3015027 | 2911541 | 2955842201300924 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-01-07 | 2013-03-21 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-28 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the surgical staff noticed that the fenestrated bipolar forceps instrument had an exposed wire that seemed to be torn apart. There were no missing or fallen pieces reported. The planned surgical procedure was completed using a back-up instrument and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of an exposed wire. The instrument was found to have a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. Engineering evaluation also found scratches on the surface of the distal pulley. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3041162 | 3015353 | 2911867 | 2955842201300925 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-02-20 | 2013-03-21 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-21 | Malfunction |  | It was reported during a da vinci si surgical procedure the large needle driver instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3041265 | 3015404 | 2911918 | 2955842201300927 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-21 | No | Yes | Health Professional | 42000606 | No | No | 2013-02-21 | No | No | Unkown | 2013
42018306 | 2013-02-22 | Malfunction |  | It was reported during a da vinci si cholecystectomy procedure the permanent cautery hook instrument was noted to have broken off and fallen into the patient. All fragments were retrieved, there was no harm or injury to the patient. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument hook was broken off. Engineering performed visual inspection, found broken ceramic sleeve, and determined that was the result of the hook broken off. The customer returned the hook, but not the broken ceramic piece. Electrical continuity testing passed. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3041286 | 3015405 | 2911919 | 2955842201300929 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-02-22 | 2013-03-21 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-02-22 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-11-07 | Malfunction |  | It was reported during a da vinci si surgical procedure, the maryland bipolar forceps instrument jaws did not close properly. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to duplicate the customer reported complaint. Engineering found the instrument jaws did not close properly. One grip was bent, causing side-to-side misalignment of grips. There was a. 092 offset at the tips, indicating overloading at the tip. Electrical continuity testing passed. Engineering also found an indentation and scratches on the pulley. There was an indentation on the edge of the pulley and scratches on the surface of the distal pulley. There were also various scratches on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3041264 | 3015438 | 2911952 | 2955842201300926 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-03-07 | 2013-03-21 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-07 | Device Not Retu | No | Reuse | 2013
42020504 | 2013-02-11 | Malfunction |  | Event Description: --/-- It was reported during a da vinci s total benign hysterectomy procedure, the fenestrated bipolar forceps instrument was noted to have a broken wire at the bottom of the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the bipolar pins were bent and pushed in at the back end of the housing, which has no contact with the patient. The top bipolar pin was pushed into the back end and the bottom pin was bent. The chassis feature that acts as a hard stop for the pins was broken. The instrument passed electrical continuity testing. Recurrence of the alleged failure mode is not expected to cause or contribute to a death, serious injury or serious deterioration in the state of health of a patient. Based on this additional failure analysis investigation information, this mdr report is being retracted. --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | Reuse | 3041292 | 3015450 | 2911964 | 2955842201300928 | 1 | NAY | Manufacturer | OtherUser facility | Other | FollowupInitial | 2013-04-08 | 2013-03-21 | No | Yes | Health Professional | 42020504 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No |  | 2013
42022703 | 2013-02-21 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the pk dissecting forceps had a broken cable. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken cable. The instrument was found to have a broken conductor wire at the yaw pulley exit. The wire was found to be detached from the connection at the grip. Electrical continuity testing of the instrument failed. The yaw pulley did not show signs of arcing. Engineering evaluation noted that the yaw pulley was missing a. 128 x. 037 piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. Engineering evaluation was unable to determine how the piece of yaw pulley broke. An additional finding was scratches on the surface of the distal pulley. Also, a rough surface finish and various scratch marks showing light material removal were found on the distal end of the instrument's main tube. The scratches were short in length and not axially aligned with the tube. Engineering evaluation concluded that the tube damages were likely due to mishandling of the instrument. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3041354 | 3015518 | 2912032 | 2955842201300930 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-21 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-21 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the large needle driver instrument had a broken cable. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken cable. The instrument was found to have one broken grip cable at the distal clevis. The broken cable segment was sticking out approximately. 3875 at the instrument's wrist. The idler pulley was able to spin. Engineering evaluation also found the distal idler pulley with rough edges and one deep indentation. Engineering evaluation concluded that the grip cable likely broke from the damaged pulley. The other cables of the instrument's wrist did not exhibit any damage. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3041622 | 3015705 | 2911918 | 2955842201300931 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-21 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2013-02-19 | Other | Required Intervention | It was reported that during a da vinci s prostatectomy procedure, the customer experienced a problem with blurry vision on the right eye of the surgeon side console. With the assistance of isi technical support engineer (tse), the customer exchanged the scope and found it to be damaged. According to the customer, the gasket seal was broken. They installed a new endoscope, but the right eye was too cloudy. They replaced the camera head and the blue vision cable, but they had no image at all. There were no errors reported from the da vinci robot system. The surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned procedure. | The field service engineer visited the site to inspect the endoscope, the blue vision cable, and the camera head. The fse could not find any issues with the camera head or with the blue vision cable. The system was repaired by replacing the defective endoscope. The system endoscope projects light from the illuminator onto the surgical site and the video image of the surgical site captured by the endoscope is sent back through left and right channels to the camera head. Upon replacement of the endoscope, the fse tested the system and found it to be working to manufacturer's specifications with no defect or malfunction. The endoscope was returned to the original equipment manufacturer (oem) for evaluation. The oem confirmed that the endoscope was received with mechanical damage to distal window assembly hermetic seal resulting in fluid invasion and subsequent minor damage to optical components. On (b)(4) 2013 the da vinci robot coordinator provided additional information regarding the reported event. She indicated that the planned surgical procedure was converted to open due to the vision issue. She also indicated that the patient didn't experience any complications, and the patient was released from the hospital. She also indicated that since the fse repaired the da vinci robot system, they have not experienced any re-occurrence of this issue. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3041977 | 3015980 | 2912494 | 2955842201300935 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-19 | 2013-03-21 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2013-02-19 | Device Not Retu | No | Reuse | 2013
40000606 | 2013-02-21 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the large needle driver instrument had a broken cable. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken cable. The instrument was found to have one grip close cable broken at the distal idlers. The idler pulley was able to spin freely and did not exhibit damage. The cable segment was observed to be sticking out at the instrument's wrist. Engineering evaluation concluded that the cable broke under tensile loading. The other cables of the instrument's wrist did not exhibit any damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | LARGE NEEDLE DRIVER INSTRUMENT | Reuse | 3041950 | 3015990 | 2911918 | 2955842201300934 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-02-21 | 2013-03-21 | No | Yes | Health Professional | 40000606 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No |  | 2013
42031802 | 2013-01-18 | Malfunction |  | Event Description: --/-- It was reported during a da vinci si sigmoid colectomy procedure the small grasping retractor instrument wire fell off after the instrument was removed from the patient. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp fell out of the distal clevis hub. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | Reuse | 3042061 | 3016075 | 2912589 | 2955842201300938 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2013-03-26 | 2013-03-21 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | Yes | 2013-03-26 | Device Not Retu | No |  | 2013
42004906 | 2013-02-13 | Malfunction |  | It was reported that after a da vinci si sacropolpexy procedure the cadiere forceps instrument collapsed down while cleaning the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the unit collapsed down the top. The luer plate was dislodged from the chassis and pushed into the backend. A section of the chassis inner wall feature that held the luer plate was broken off. Engineering concluded the wall feature likely broke due to improper cleaning, which allowed the luer plate to dislodge. Engineering also found various scratch marks on the main tube at the distal end showing light material removal. The scratches were. 080 -. 125 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3042062 | 3016077 | 2912591 | 2955842201300939 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-25 | 2013-03-21 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-02-25 | No | No | Reuse | 2013
42030901 | 2013-02-12 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the mega suturecut needle driver instrument tip was not moving. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument tip was not moving. The one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the edge and on the surface. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3042023 | 3016100 | 2912614 | 2955842201300936 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-22 | 2013-03-21 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-22 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-20 | Malfunction |  | It was reported during a da vinci si surgical procedure, the mega suturecut needle driver instrument was noted to have frayed cables at the distal end. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken grip cable. The one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found some scratches on the surface of the distal pulley. There were various scratch marks on the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3042024 | 3016102 | 2912616 | 2955842201300937 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-25 | 2013-03-21 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No | Reuse | 2013
| 2013-03-12 | Malfunction |  | During the robotic procedure, the surgeon wanted to use the vessel sealer. The tech loaded the vessel sealer into arm 1 of the patient cart of the robot. It was witnessed that the vessel sealer will burn out and cut on the first try, afterwards the surgeon tried to use it again and it was witnessed that the vessel sealer would burn only. Another device was needed. There was no harm to the patient. |  | ENDOWRIST ONE VESSEL SEALER | SYSTEMSURGICALCOMPUTER CONTROL | 3047292 | 3021230 | 2917744 | 3021230 | 1 | NAY | User Facility |  |  | Initial | 2013-03-19 | 2013-03-21 | No | Yes | UNKNOWN |  | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
42009308 | 2013-01-29 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the prograsp forceps instrument had an exposed frayed wire at the distal end. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the nonconformance was a broken pitch cable. The pitch cable was broken at the wrist. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037874 | 3011940 | 2875490 | 2955842201300907 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-19 | 2013-03-20 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-19 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-20 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the pk dissecting forceps instrument was noted to have a broken wire at the distal end. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the nonconformance was a broken conductor wire. The one conductor wire was broken at the yaw pulley exit. The wire was detached from its connection at the grip. The electrical continuity testing failed. The yaw pulley was missing a piece at the same area of the conductor break, allowing the grip to move within the pulley and have a larger range of motion in the yaw. It was unclear how the yaw pulley broke. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3037877 | 3011941 | 2908455 | 2955842201300910 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-20 | Malfunction |  | It was reported that prior to a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have a frayed cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was broken at the distal end. The cable segment that contained the crimp was still installed in clevis and stuck out at the wrist. The clevis did not exhibit any damage or wear marks. The other cables at wrist were not damaged. Engineering also found a bent bipolar pin. The bottom pin was not aligned with the top pin it was bent up. Engineering concluded that damage was likely due to mishandling. Electrical continuity testing passed. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3037878 | 3011942 | 2908456 | 2955842201300911 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-20 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-20 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure, the fenestrated bipolar forceps instrument was noted to have frayed cables. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was frayed. The conductor wires were intact and undamaged, but the drive cable was frayed. The pitch down cable was frayed at the distal clevis hub. The frayed strands stuck out at the wrist. The other cables at wrist were not damaged. Engineering also found various scratch marks on the distal end of the main tube showing light material removal. The scratches were. 105 -. 200 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3037891 | 3011943 | 2908456 | 2955842201300912 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-20 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-07 | Malfunction |  | Event Description: --/-- It was reported during a da vinci si surgical procedure the monopolar curved scissors instrument was noted to have frayed and broken cables and would not allow the surgeon to move the wrist. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to duplicate the customer reported complaint. Engineering found that visual inspection shows no damage to the drive cables at the wrist. The scissors opened and closed fine when the input discs were manually rotated. Engineering found a broken tube extension and main tube damage. The tube extension was broken and missing. A 110 x. 110 piece at the proximal clevis interface. The clevis was dislodged from the tube extension. The tube extension fractured next to one of the keys that mate with the proximal clevis. The electrical continuity testing passed. There were various scratch marks on the distal end of the main tube showing light material removal. The scratches were. 070 -. 450 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 3037880 | 3011944 | 2908458 | 2955842201300915 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-02-20 | 2013-03-20 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42017207 | 2013-02-15 | Malfunction |  | It was reported during a da vinci si surgical procedure, the maryland bipolar forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken cable. The pitch cable was broken at the wrist. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3037894 | 3011945 | 2908459 | 2955842201300917 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-21 | 2013-03-20 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure, the cable on the maryland bipolar forceps instrument snapped during dissection on the left side of the uterus, making the instrument defective. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at distal end. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at wrist were not damaged. Engineering also found bent grip tips, bent bipolar pins, a broken conductor wire, and corroded clamping pulleys. The instrument came back with the grips not aligned. The one grip tip was bent at an angle, causing the tip to form an x. There was a. 0176 offset at the tips. The bipolar pin at the back end of the housing was bent. The bottom pin was not aligned with the top pin. Engineering concluded that damage was likely due to mishandling. The conductor wire connected to the top bipolar pin was broken from inside the housing. The wire no longer had a connection to the pin. Multiple clamping pulleys exhibited contamination that possibly caused increased friction, which could have been the cause of the cable breakage. Engineering concluded the damage was likely due to improper cleaning. The electrical continuity testing failed. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3037876 | 3011960 | 2908474 | 2955842201300908 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-20 | 2013-03-20 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-18 | Malfunction |  | It was reported that during a da vinci si sigmoid colectomy procedure the mega suture cut needle driver instrument lines were noted to be frayed. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument lines were frayed. The pitch cable was broken at the distal end. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. Engineering also found the back idler pulley was corroded. A few back idler pulleys exhibited some rust. The idler pulleys were not seized and still rotated freely. Engineering concluded the damage was likely due to improper cleaning. No other damage was found. The cleaning and sterilization user manual specifically states: o when scrubbing the tip of cautery instruments, take care not to damage the insulation. O do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. O prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037875 | 3011961 | 2908475 | 2955842201300909 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-20 | 2013-03-20 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-02-07 | Malfunction |  | It was reported during a da vinci si surgical procedure the thoracic grasper 5mm instrument jaws would not close. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument jaws would not close. The grips did not open/close properly when input discs were manually rotated. Visual inspection showed one grip link was broken. The link was broken at the edge of the hole that attaches to link pin. The cross sectional area of the link was at it's minimum at the location of breakage. The other distal end components were not damaged. No other damage was found. | THORACIC GRASPER 5MM INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037892 | 3011962 | 2908476 | 2955842201300914 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-20 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-20 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have a broken wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken wire. The pitch up cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. Engineering also found various scratch marks at the distal end of the main tube showing light material removal. The scratches were. 085 -. 185 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3037879 | 3011963 | 2908477 | 2955842201300913 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-20 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-11 | Malfunction |  | Event Description: --/-- It was reported during a da vinci s surgical procedure the pk dissecting forceps instrument was noted to have a broken wire at the tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp remained in the clevis. The clevis did not exhibit excessive damage. Engineering also found the conductor wire was broken at the yaw pulley exit, with damages found on the yaw pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | PK DISSECTING FORCEPS INSTRUMENT | Reuse | 3037893 | 3011964 | 2904346 | 2955842201300916 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-04-08 | 2013-03-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-04-08 | Device Not Retu | No |  | 2013
42030901 | 2012-11-30 | Malfunction |  | It was reported during a da vinci si total benign hysterectomy procedure the megasuturecut needle driver instrument wire came out of the pulley and malfunctioned. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the nonconformance was a derailed grip cable. The grips cannot open and close properly. The movement was hindered. Engineering also found the derailed grip cable was frayed. The frayed cable sections are in various lengths sticking out at the wrist. The idler pulley rim exhibited damage. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037881 | 3011965 | 2908479 | 2955842201300918 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-20 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42805912 | 2013-02-14 | Malfunction |  | It was reported during a da vinci si surgical procedure, the crocodile grasper instrument was noted to have a broken tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument tip was broken. The clevis was cracked by the proximal area and dislodged. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | CROCODILE GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037882 | 3011966 | 2908480 | 2955842201300919 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-22 | 2013-03-20 | No | Yes | Health Professional | 42805912 | Device Returned To Manufacturer | No | 2013-02-22 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-06 | Malfunction |  | It was reported during a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have wires sticking out and would not coagulate. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument wires were sticking out. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. Engineering also found various scratches on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3037895 | 3011967 | 2908481 | 2955842201300920 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-22 | 2013-03-20 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-22 | Device Not Retu | No | Reuse | 2013
| 2012-12-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy (bso) procedure, the surgeon was having issues with arm 1 control. The surgeon reportedly experienced resistance rotating and inserting an instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2012, the intuitive surgical technical service engineer (tse) had the customer reseat the sterile adapter but the reported issue was not resolved. In addition, the customer performed an emergency power off (epo) and a power cycle system but the issue was still not resolved. The customer claimed that the instruments were really difficult to insert and remove. The instrument cannula was inspected and found to be cracked and bent out of shape. The customer replaced the cannula and was able to complete the surgical procedure. On (b)(6) 2012, isi contacted the tse and obtained additional information. According to the tse, the cannula was found to be cracked upon removal of the cannula. According to the tse, the customer installed a new cannula and the issue was resolved. The tse did not know when the cannula was damaged. He also was unable to provide any other details regarding the cannula. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. | CANNULA | ENDOSCOPIC ACCESSORY | 3039166 | 3013279 | 2909793 | 2955842201300921 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-03-20 | No | Yes | Health Professional |  | No | No | 2012-12-10 | No | No | Unkown | 2013
42022703 | 2013-02-20 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the pk dissecting forceps instrument had a broken wire at the distal end. The reported issue was found during set up and the instrument was not used in the planned procedure. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken wire. One conductor wire was found to be broken at the yaw pulley exit. The wire was detached from its connection at the grip. The instrument failed electrical continuity testing. The yaw pulley did not show any signs of arcing. Engineering evaluation also noted that the yaw pulley was missing a. 179 x. 056 piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. Engineering evaluation was unable to determine how the piece of yaw pulley broke. On (b)(4) 2013, isi contacted the initial reporter of this complaint and verified that no fragments were reported to have fallen. In addition, the initial reporter denied that any patients were harmed or injured because of the reported issue. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3039623 | 3013638 | 2908455 | 2955842201300922 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
40027202 | 2013-02-06 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the surgical staff allegedly noticed that the harmonic ace curved shears insert was not working. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a that the ace insert was not working. The instrument insert accessory was installed on an in-house harmonic ace curved shears instrument without issues. The insert was placed on a system and driven. The instrument blade was observed to move sideways when the clamp arm opened and closed. The clamp arm appeared to be pushing blade. Repeated opening and closing of the blade on system resulted in a fracture of the curved blade during in-house testing. The blade was found to have fractured at the clamp arm pin. Engineering evaluation disassembled the insert and a wear mark was found on the blade where it meets the pin. Engineering evaluation concluded that the blade likely made contact with the clamp arm pin during energy activation, causing a crack to develop and energy transmission to fail at the tip. Engineering evaluation discovered a malfunction of the instrument that if to recur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 3039765 | 3013773 | 2910287 | 2955842201300923 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-20 | 2013-03-20 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | Yes | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-11-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the strings of the maryland bipolar forceps instrument snapped off on both sides. The procedure went well and the patient was transferred to the recovery room. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that both pitch cables were found to be broken at the distal clevis hub. Both cable segments that contain the crimp were still installed in the clevis. The clevis did not exhibit any damage or wear marks. Additional observation not reported by the site was a missing piece. The yaw pulley below the base was missing a piece on one of the grips. The yaw pulley also exhibited localized charring. Internal laboratory testing has determined that the arc track failure mode is most likely to occur when a bipolar instrument is energized with no tissue between the grips. Internal arcing in bipolar instruments results in a potential loss of product function, but no transfer of energy to the patient. Intuitive has concluded that the likelihood of injury due to this failure is remote. User are instructed to retain back-up instruments in case that an instrument fails to perform, and are also instructed in proper technique to minimize the likelihood of internal arcing. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3036742 | 3010833 | 2907347 | 2955842201300892 | 1 | NAY | Manufacturer | DistributorHealth Professional | Nurse | Initial | 2013-02-20 | 2013-03-19 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that the prograsp forceps instrument had cut wires. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a cut wire. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering evaluation also found scratches on the distal pulley. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3036992 | 3011196 | 2907710 | 2955842201300893 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-19 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-14 | Injury | Required Intervention | Event Description: --/-- Received medwatch form from the customer. According to the mdr report, pk dissector inspected prior to use in pt. Pk dissector notice by surgeon that it has frayed an hour after initial usage. Instrument malfunction by not being able to be removed from the port. The port with instrument in place removed from patient. Angle of dissector malfunction. When removed from port (outside of body) it was noticed that a cable had broke. Small cable piece found inside the trocar. X-ray taken per policy. X- ray taken = negative. | Manufacturer Narrative: --/-- --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was missing. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were 0. 075 - 0. 170 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | Reuse | 3037005 | 3011213 | 2907727 | 2955842201300894 | 1 | NAY | Manufacturer | OtherUser facility | Other | InitialFollowup | 2013-02-20 | 2013-03-19 | Yes | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No |  | 2013
42030901 | 2013-01-07 | Injury | Required Intervention | Received medwatch report from the customer, according to the mdr report, 'robotic instrument was noted to have broken during laparoscopic robotic assisted surgery. Wires from instrument were seen frayed inside patient and a small piece of metal wire was removed by m. D. An x-ray was performed at the end of the case to confirm no other foreign material was left inside. The x-ray performed to rule out any retained foreign items. X-ray took found zero foreign items seen. '. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely, but it exhibited some damage on the edge and on the surface of the pulley. The cable segment sticks out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037088 | 3011270 | 2907784 | 2955842201300895 | 1 | NAY | Manufacturer | OtherHealth Professional | Nurse | Initial | 2013-02-20 | 2013-03-19 | Yes | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-02-20 | Device Not Retu | No | Reuse | 2013
| 2012-11-27 | Injury | Required Intervention | Received medwatch report from the customer, according to the mdr report, 'surgeon uses product and was working without incident. When surgeon used the needle driver instrument with needle on it, she noticed the wire was frayed. The needle driver instrument was replaced. The needle driver instrument didn't appear to have any broken pieces. X-ray taken = negative. '. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037321 | 3011415 | 2907929 | 2955842201300896 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-20 | 2013-03-19 | Yes | Yes | Health Professional |  | No | Yes | 2013-02-20 | No | No | Reuse | 2013
42000606 | 2013-02-20 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff noticed that the large needle driver instrument had a wire that snapped. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a snapped wire. One grip close cable was found to be broken at the distal idlers pulley. The cable segment was sticking out at the instrument's wrist. The idler pulley was able to spin freely. Engineering evaluation also observed damages on the edge and on the surface of the idler pulley. The other cables of the instrument's wrist did not exhibit any damage. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037329 | 3011424 | 2907938 | 2955842201300897 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-19 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-11-20 | Injury | Required Intervention | Received medwatch report from the customer, according to the mdr report, 'during a robot-assisted laparoscopic myomectomy, uterine reconstruction, the pk dissecting forceps instrument stopped working. The wire controlling the articulating jaw had frayed. Surgeon notified. The instrument sequestered. X-ray taken, radiologist found nothing in the abdomen. '. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that both pitch cables were found to be broken at the distal clevis hub. The cable segments that contain the crimp were still installed in the clevis. The clevis did not exhibit any damage or wear marks. Additional observation not reported by site was a derailed cable. One grip close cable was found to be derailed at the distal idler pulley. The grips were still able to open and close, but their movement could not be precise. Evidence not conclusive, but cable derailment was likely due to cable losing contact with pulley during wrist articulation. No other damage was found. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3037391 | 3011461 | 2907975 | 2955842201300898 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-20 | 2013-03-19 | Yes | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-11-20 | Malfunction |  | Received medwatch report from the customer, according to the mdr report, 'instrument noted to have a tip damaged; during pre-operational check, the scrub noted a frayed edge. No harm to patient was reported, instrument was not used. '. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent, causing side to side misalignment of the grips. There was a 0. 097" offset at the tips, indicating overloading at the tip. Instrument passed electrical continuity test. Additional findings not reported by the site were frayed pitch cable and pulley's scratches. Instrument was found with frayed pitch cable located at the distal idler. The distal pulley also exhibited some scratches. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3037474 | 3011502 | 2908017 | 2955842201300899 | 1 | NAY | Manufacturer | OtherUser facility | Other | Initial | 2013-02-20 | 2013-03-19 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42019408 | 2012-10-16 | Injury | Required Intervention | Received medwatch report from the customer, according to the mdr report, 'during the operation, a wire detached from the tip of the needle driver instrument while it was being used in a laparoscopic surgery. It was unclear during surgery whether any wire fragments came loose/broke off in the patient's body cavity. The defect of this product resulted in the need for two intraoperative x-rays to verify that no wire fragments were left behind. '. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037491 | 3011572 | 2908086 | 2955842201300900 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-20 | 2013-03-19 | Yes | Yes | Health Professional | 42019408 | No | Yes | 2013-02-20 | No | No | Unkown | 2013
42034301 | 2013-02-20 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the thoracic grasper instrument had a damaged tube. The reported issue was found during set up and the instrument was not used in the planned procedure. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a torn black housing. The insulation coating of the instrument's main tube was found to be peeled off near the distal end. There also scratches observed along the main tube. On (b)(4) 2013, isi contacted the initial reporter and verified that no fragments fell into a patient. He also denied that any patients were harmed or injured because of the reported issue. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported main tube damage issue if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037500 | 3011583 | 2908097 | 2955842201300901 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-20 | 2013-03-19 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-10-16 | Injury | Required Intervention | Received medwatch report from the customer, according to the mdr report, 'during operation, a wire detached from the tip of the needle driver instrument while it was being used laparoscopically. It was unclear during surgery whether any wire fragments came loose in the patient's body cavity. The defect of this product resulted in the need for two intraoperative x-rays to verify no wire fragments were left behind. '. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely, but it exhibited some damages on the edge and on the pulley's surface. The cable segment sticks out at wrist. Other cables at wrist were not damaged. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037516 | 3011601 | 2908115 | 2955842201300902 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-20 | 2013-03-19 | Yes | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-02-20 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the thoracic grasper instrument had a damaged tube. The reported issue was found during set up and the instrument was not used in the planned procedure. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a torn black housing. The insulation of the instrument's main tube was found to have scratches and gouge marks with material removed at the distal end. Engineering evaluation concluded that the damage was likely caused by excess force contact on the main tube surface. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, isi contacted the initial reporter of this complaint and obtained additional information. The initial reporter verified that no patient's were harmed or injured because of the reported issue. He also denied that any fragments fell into a patient. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037531 | 3011606 | 2908097 | 2955842201300903 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-20 | 2013-03-19 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42019408 | 2012-09-13 | Injury | Required Intervention | Received medwatch report from the customer, according to the mdr report, 'under direct visualization during a robotic laparoscopic procedure, articulating wire snapped while the mega needle driver instrument was in use. No fragment of wire was noted in surgical field. The device was removed and examined. One wire noted to be broken. No patient harm or surgery delay. X-ray took post operation negative for any retained items. '. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable located at distal clevis hub. There was no damage found at the clevis. The frayed segment was approximately 0. 04 in length. No other cable damage was found. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037612 | 3011688 | 2908202 | 2955842201300904 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-20 | 2013-03-19 | Yes | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-09-12 | Malfunction |  | Received medwatch report from the customer, according to the mdr report, 'robotic mega suture cut needle driver instrument had a wire snap intra-operatively. No pieces observed to release into abdomen. '. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely but it exhibited some damage on the edge and on the surface. The cable segment sticks out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3037688 | 3011745 | 2908258 | 2955842201300905 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-20 | 2013-03-19 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-02-20 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-08-30 | Malfunction |  | Received medwatch report from the customer, according to the mdr report, 'fenestrated bipolar instrument had a misaligned jaw from a defective freyed cable. Not obvious patient harm. '. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3037705 | 3011786 | 2651627 | 2955842201300906 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-20 | 2013-03-19 | No | Yes | Health Professional | 42020505 | No | Yes | 2013-02-20 | No | No | Unkown | 2013
42031802 | 2013-02-19 | Malfunction |  | It was reported that during a da vinci si total proctocolectomy procedure, the surgical staff noticed that a cable in the small grasping retractor instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken cable. The pitch cable was found to be broken at the distal end of the instrument. The cable segment that contains the crimp was still installed in the clevis and was sticking out at the instrument's wrist. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Engineering evaluation also found deep scratches on the instrument's main tube. The distal end of the main tube had various scratches that exhibited light material removal and a rough surface finish. The scratches were approximately 2 long at distal end. Engineering evaluation concluded that the scratches were likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable and main tube scratch issues if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3034404 | 3008930 | 2905444 | 2955842201300879 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | Yes | 2013-02-19 | Device Not Retu | No | Reuse | 2013
42000304 | 2013-02-07 | Malfunction |  | It was reported that during a da vinci si thoracic sympathectomy procedure, the surgical staff noticed that tips of the small clip applier instrument broke off. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of the instrument's tips breaking off. The instrument was returned with both clip applier grips broken. The broken pieces were not returned. Both grips fractured near the base of the clip groove. Intact sections of the grips were in different lengths. One grip exhibited a small bend just below the fractured surface. Engineering evaluation concluded that the damage may be due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken tip issue if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3034460 | 3009030 | 2905544 | 2955842201300880 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 42000304 | Device Returned To Manufacturer | No | 2013-02-19 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-19 | Malfunction |  | It was reported that during a davinci si diaphragmatic hernia repair procedure, the customer noted a broken cable on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were 0. 105 - 0. 450 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3034552 | 3009080 | 2905594 | 2955842201300881 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-19 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-02-12 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the customer noted a broken cable on the mega needle driver instrument. The instrument was replaced and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3034616 | 3009102 | 2905616 | 2955842201300882 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | Yes | 2013-02-19 | Device Not Retu | No | Reuse | 2013
40030901 | 2013-02-18 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, the surgical staff noticed that the wire at the tip of the mega suturecut needle driver instrument was broken and exposed. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken wire. The instrument was found to have a broken grip cable at the wrist. The idler pulley was able to spin freely. Engineering evaluation also noted that the instrument exhibited pulley face and rim damage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken wire issue if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3034641 | 3009139 | 2905653 | 2955842201300883 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 40030901 | Device Returned To Manufacturer | No | 2013-02-19 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-18 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that one of the wires of the mega suturecut needle driver instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken wire. The instrument was found to have one grip close cable broken at the distal idlers. The idler pulley was able to spin freely and was not damaged. The cable segment was sticking out at the instrument's wrist. The other cables of the instrument's wrist did not exhibit any damage. Engineering evaluation also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3034796 | 3009207 | 2905721 | 2955842201300885 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-19 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-31 | Malfunction |  | It was reported that prior to starting a da vinci si sacrocolpopexy procedure, it was noted that the tips of the fenestrated bipolar forceps instrument were misaligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent, causing side to side misalignment of grips. There was a 0. 078" offset at the tips, indicating overloading at the tip. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3034875 | 3009233 | 2883261 | 2955842201300886 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-19 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-13 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si myomectomy procedure, the surgical staff noticed that the fenestrated bipolar forceps instrument had a broken cable. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation confirmed the alleged issue that the instrument had a broken cable. The instrument's pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. Engineering evaluation observed that the clevis did not exhibit excessive damage. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | Reuse | 3035027 | 3009312 | 2905826 | 2955842201300889 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No |  | 2013
42022703 | 2013-02-01 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that one of the wires of the pk dissecting forceps instrument was exposed. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of an exposed wire. The grip cable was found to be frayed at the distal idler pulley. Frayed strands were observed to be sticking out at the instrument's wrist. The other cables of the instrument's wrist did not exhibit any damage. Electrical continuity testing of the instrument passed. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3035149 | 3009418 | 2884912 | 2955842201300888 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-19 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-15 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that one of the wires of the mega suturecut needle driver instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken wire. One grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely. The cable segment was sticking out at the instrument's wrist. The other cables of the instrument's wrist did not exhibit any damage. Engineering evaluation also found scratches on the surface and edges of the idler pulley. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3035208 | 3009468 | 2905721 | 2955842201300887 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-19 | Device Not Retu | No | Reuse | 2013
40022702 | 2012-07-16 | Malfunction |  | It was reported during a da vinci s surgical procedure that the pk dissecting forceps instrument were faulty. No other information was provided. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to duplicate the customer reported complaint. Engineering found no problem at the tip area. Engineering found broken housing. All four housing snap tabs were broken. The housing could be removed as a result of the damage. Engineering concluded the instrument may have been mishandled. The electrical continuity testing passed. Engineering also found various scratch marks at the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage to the main tube was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3035721 | 3009953 | 2906467 | 2955842201300890 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-03-06 | 2013-03-18 | No | Yes | Health Professional | 40022702 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-12 | Malfunction |  | It was reported that during a davinci si surgical procedure, the customer noted a broken cable on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment sticks out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA SUTURE CUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3035699 | 3009967 | 2904345 | 2955842201300884 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-19 | 2013-03-18 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-02-19 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2013-02-11 | Death | Death | It was alleged that the patient had a multi-port cholecystectomy on (b)(6) 2013 that resulted in an injury to the bowel from the trocar during insertion at the beginning of the case. The patient complained about pain two days after the surgery and returned to the hospital. The patient was readmitted, allegedly septic and an emergency surgery was performed to repair the bowel. The patient expired on (b)(6) 2013. | On (b)(4) 2013, the isi company representative was contacted to request additional information regarding the reported complaint. It was indicated that no malfunctions occurred with the da vinci si system, instruments or accessories; no complications were noted during the surgery. It was reported that the initial surgery was uneventful and there were no adhesions. The manufacturer of the trocar is unknown at the time of this report. The patient's prior medial history is not known. A review of system logs for the system allegedly used in the procedure did not produce any data for the given system at the time period indicated. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | DA VINCI SI SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3035700 | 3009969 | 2906483 | 2955842201300891 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-02-18 | 2013-03-18 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-02-18 | No | No | Reuse | 2013
420227 | 2013-03-15 | Malfunction |  | During a robotic hysterectomy case, the surgeon found the pk dissecting forceps endowrist instrument to be faulty. On further inspection, the wires that control the instrument broke. The abdomen was surveyed by surgeon and pa; it was confirmed no foreign bodies were left in patient abdomen. It was also noted that out of the 10 uses of the pk dissecting forceps, there were 2 uses left. Manufacturer response for pk endowrist dissecting forceps, pk endowrist dissecting forceps (per site reporter). Device was returned for their quality review. They will respond with their results. What was the original intended procedure?robotic assisted hysterectomy. |  | ENDOWRIST PK DISSECTING FORCEPS | SYSTEMSURGICALCOMPUTER CONTROL | 3053720 | 3026842 | 2923356 | 3026842 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-03-18 | 2013-03-18 | No | Yes | Physician | 420227 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
42030901 | 2013-02-21 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the cable on the megasuturecut needledriver instrument was noted to be broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip cable was found frayed at the distal and proximal clevis. The frayed strands were sticking out at the wrist and the distal idler pulley was damage. A small dent found on the edge of the pulley may have likely caused the cable to fray. The other cables at the wrist were not damaged. Engineering evaluation also found several indentations on both blades of the instrument. Engineering concluded that the damage to the blades was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3033358 | 3007935 | 2904449 | 2955842201300873 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42018304 | 2013-02-11 | Malfunction |  | It was reported that during a davinci si surgical procedure, the jaws did not align on the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Customer reported that the jaws didn't align; however, the permanent cautery hook instrument does not have any grips. Upon visual inspection the tip of the permanent cautery hook instrument didn't appear to be bent. Additional observations not reported by site were charred components, missing conductor cap, cracked proximal clevis, and missing luer plate. The yaw pulley and one distal clevis ear exhibited char marks and localized melting. No damage to the conductor wire was observed. The conductor cap was missing from the distal end. The yaw pulley boss feature that mates with cap was broken. The proximal clevis exhibited a crack in the axial direction located halfway between ears. The orientation of the crack was consistent with overloading the tip with the wrist pitched. The luer plate and the flush tube were missing from the backend housing. The chassis outer wall feature that mates with the luer plate was also found to be broken. Engineering concluded that the damage found on the luer plate was likely due to improper cleaning. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033361 | 3007939 | 2904453 | 2955842201300878 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-15 | 2013-03-17 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2013-02-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-11 | Malfunction |  | It was reported that during reprocessing of the pk dissecting forceps instrument the cable on the instrument was found to be loose. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable frayed at both the back idler pulley and the distal end. The same cable was also derailed at the back idler pulley. No other cables were damaged at the wrist. There was no damage or corrosion found on back idler pulley. The instrument also passed electrical continuity testing. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 3033372 | 3007956 | 2904470 | 2955842201300872 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-17 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-21 | Malfunction |  | It was reported that during reprocessing of the megasuturecut needledriver instrument, the cable on the instrument was noted to be broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely, but has damages on the edge and on the surface. Engineering concluded that damage to the pulley was likely due to excess force contact and mishandling. The cable segment was sticking out at the wrist. The other cables at the wrist were not damaged. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3033373 | 3007957 | 2904472 | 2955842201300874 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-21 | Malfunction |  | It was reported that during reprocessing of the megasuturecut needledriver instrument, the cable on the instrument was noted to be broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment was sticking out at the wrist. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3033374 | 3007958 | 2904472 | 2955842201300875 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-21 | 2013-03-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-05 | Malfunction |  | It was reported that after successfully da vinci s prostatectomy procedure on (b)(6) the surgical staff found a jagged wire of the large needle driver instrument during the irrigation process. The planned procedure was completed without problem at that time. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip cable was found to be frayed at the distal the idler pulley and the proximal clevis. The frayed strands stick out at the wrist. Other cables at wrist were not damaged. There was not damage found on the idler pulley, the idler pulley was able to rotate freely. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3033375 | 3007959 | 2904473 | 2955842201300876 | 1 | NAY | Manufacturer | DistributorHealth Professional | Physician | Initial | 2013-02-15 | 2013-03-17 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2013-02-15 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-11 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted that the jaws of the maryland bipolar forceps instrument were not aligning. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent, causing side to side misalignment of grips. There is a 0. 125 offset at the tips. The bent grip had a separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were 0. 090 - 0. 165 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033376 | 3007960 | 2904474 | 2955842201300877 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-15 | 2013-03-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-14 | Malfunction |  | It was reported that during a davinci s surgical procedure, the customer noted a broken wire on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033350 | 3007926 | 2904447 | 2955842201300864 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-03-16 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-11-29 | Malfunction |  | Event Description: --/-- It was reported that during a davinci si surgical procedure, the customer noted a broken wire on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. --/-- Additional Event Description: --/-- On (b)(6) 2013 we received a user facility medwatch from the fda with the following information: 'during a robotic assisted hysterectomy it was noted by the surgeon that the wire was frayed at the tip of the instrument at the pulley system. The mega suturecut needle driver was removed from the patient and use. '. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment sticks out at wrist. Other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2013 we received the user facility medwatch form from the fda,. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | Reuse | 3033353 | 3007928 | 2904442 | 2955842201300865 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-03-19 | 2013-03-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No |  | 2013
42022703 | 2013-02-04 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgeon noticed that the instrument was not grasping good on tissue. The surgeon did not notice any loose cables under magnification. The surgical staff replaced the instrument for another pk dissecting forceps instrument to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on the in-house da vinci system robot and driven. Recognition and engagement test passed. Instrument moved intuitively with full range of motion in all directions. The grips opened, closed and grasped properly. Instrument was found to be fully functional. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033357 | 3007932 | 2904446 | 2955842201300869 | 1 | NAY | Manufacturer |  | Other | Initial | 2013-02-14 | 2013-03-16 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-15 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a broken cable on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was a bent grip. One grip was found to be bent, causing side to side misalignment of grips. There was a 0. 055 offset at the tips. The bent grip exhibited separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. Instrument passed electrical continuity test. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033328 | 3007933 | 2904447 | 2955842201300871 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-15 | 2013-03-16 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-15 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-02-13 | Malfunction |  | It was reported that during a davinci si surgical procedure, the customer noted a broken cable on the tenaculum forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3033323 | 3007949 | 2904463 | 2955842201300866 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-03-16 | No | Yes | Health Professional | 42020706 | No | No | 2013-02-14 | No | No | Unkown | 2013
42030901 | 2013-02-08 | Malfunction |  | It was reported that during a davinci si ovarian cystectomy procedure, the customer noted a broken cable at the tip of the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers and grip hub. The idler pulley was able to spin freely and did not exhibit any damage. The cable segment sticks out at wrist. Grip hub exhibits light scratches near the grip hub breakage location. Other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3033325 | 3007951 | 2904465 | 2955842201300867 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-03-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-05 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgeon noticed that the bipolar black cord was sticking out further than usual once the pk dissecting forceps instrument was inserted in the cannula. The surgical staff exchanged to another pk dissecting forceps instrument to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the conductor wire was found to be dislodged at the yaw pulley exit. The wire was loose from its connection at the grip. Instrument passed electrical continuity test. The yaw pulley shows no signs of arcing. Additional observation not reported by the site was a broken yaw pulley. The yaw pulley was missing a 0. 078 x 0. 110 piece. It is unclear how piece of yaw pulley broke. In addition, engineering also observed a slipped crimp cable. One of the cable crimp slipped out of the yaw pulled where the broken pulley occurred. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033326 | 3007952 | 2904311 | 2955842201300868 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-03-16 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-06 | Malfunction |  | It was reported that during cleaning and sterilization, the tips of the maryland bipolar forceps were found to be damaged while cleaning. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033327 | 3007953 | 2904467 | 2955842201300870 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-03-16 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-04 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff noticed that the mega suturecut needle driver instrument had a broken wire at the hinge. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(6) 2013, isi contacted the initial reporter of this complaint and verified that no fragments fell into the patient. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | DOUBLE FENESTRATED GRASPERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3031959 | 3006593 | 2832909 | 2955842201300833 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-04 | 2013-03-15 | No | Yes | Health Professional | 42030901 | No | Yes | 2013-01-04 | No | No | Reuse | 2013
42000306 | 2013-02-15 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure the small clip applier instrument tips broke when the clips were loaded. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that both clip applier grips are broken at the midpoint. The damage was likely due to excess force applied normal to the grip while trying to install a clip. Engineering concluded that damage was likely due to mishandling/user error. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3031938 | 3006606 | 2903120 | 2955842201300834 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-15 | 2013-03-15 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-02-15 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-14 | Malfunction |  | The monopolar curved scissors instrument was returned without any allegations of malfunction. No further information was received. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. The instrument was placed on a system and driven. Recognition and engagement testing passed. There was no evidence of blade damage found and pogo pins were not stuck or contaminated. Engineering also found several scratch marks at the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. Electrical continuity testing passed. No other damage was found. 62 - the instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3032308 | 3006893 | 2903407 | 2955842201300837 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-15 | 2013-03-15 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-02-15 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-13 | Malfunction |  | It was reported that during a da vinci s surgical procedure the surgical staff noted the large needle driver instrument had a small dip on the wire at the wrist during irrigation. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a small dip on the wire of the large needle driver instrument wrist. Visual inspection of the wrist under magnification showed a small fray in the one grip close cable. The cable was frayed at the grip hub. The frayed strands were stuck out from the cable. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3032338 | 3006898 | 2903412 | 2955842201300838 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the fenestrated bipolar forceps instrument was noted to have a wire was sticking out of the jaws of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the wire that sticks out of the jaws of the instrument. The pitch cable was found frayed at the proximal clevis. The frayed strands stuck out at the wrist. No other damage was found at the distal end. The electrical continuity testing passed. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3032332 | 3006921 | 2879513 | 2955842201300836 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-14 | Malfunction |  | It was reported that prior to starting a da vinci si prostatectomy procedure the prograsp forceps instrument was noted to have a frayed cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument was found with frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment is approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3032342 | 3006937 | 2903451 | 2955842201300839 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-14 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the monopolar curved scissors instrument was noted to be broken at the tip. No missing or fallen pieces were reported. No patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3032349 | 3006951 | 2903465 | 2955842201300840 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 42017910 | No | Yes | 2013-02-14 | No | No | Unkown | 2013
42009308 | 2012-12-14 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the prograsp forceps instrument was noted to have a wire broken at the hinge. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3032371 | 3006978 | 2856382 | 2955842201300841 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 42009308 | No | Yes | 2013-02-14 | No | No | Unkown | 2013
42030901 | 2012-12-20 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure that the megasuturecut needle driver instrument was noted to have a broken wire at the hinge. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3032373 | 3006984 | 2813514 | 2955842201300842 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 42030901 | No | Yes | 2013-02-14 | No | No | Unkown | 2013
42009308 | 2012-12-17 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the surgical staff noticed that the prograsp forceps instrument had a broken cable at the hinge. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3032754 | 3007367 | 2856382 | 2955842201300843 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-12-17 | 2013-03-15 | No | Yes | Health Professional | 42009308 | No | No | 2013-12-17 | No | No | Unkown | 2013
42030901 | 2013-01-08 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, the surgical staff allegedly noticed that the mega suturecut needle driver instrument had a broken wire. No patient harm or adverse outcome was reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken wire. One grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and was not damaged. The cable segment was sticking out at the instrument's wrist. Other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3032795 | 3007384 | 2894974 | 2955842201300844 | 1 | NAY | Manufacturer | User facility | Nurse | Initial | 2013-02-15 | 2013-03-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-02-15 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-15 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the large suturecut needle driver instrument was not recognized by the system. The reported issue was found during set up and was not used in the planned procedure. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a frayed wire. Engineering evaluation found a frayed wire at the distal end. One grip close cable was observed to be broken at the distal idlers pulley. The idler pulley was able to spin freely but slight damages on the edge and on the surface were found. The cable segment was sticking out at the instrument's wrist. Other cables at instrument's wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- Correction to description from the initial report: it was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the large suturecut needle driver instrument had a frayed wire at the distal end. The reported issue was found during set up and the instrument was not used in the planned procedure. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | Reuse | 3032833 | 3007434 | 2903948 | 2955842201300845 | 1 | NAY | Manufacturer | User facility | Other | FollowupInitial | 2013-02-15 | 2013-03-15 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-02-15 | Device Not Retu | No |  | 2013
42009308 | 2013-02-15 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the prograsp forceps instrument had a broken cable. The reported issue was found during set up and the instrument was not used in the planned procedure. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken cable. The instrument was found to have a frayed cable at distal idler pulley. The frayed cable section was approximately 0. 07 in length. No damage was found on the pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3032911 | 3007499 | 2904013 | 2955842201300846 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-15 | 2013-03-15 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-15 | Device Not Retu | No | Reuse | 2013
40027202 | 2012-10-30 | Malfunction |  | It was reported that during a da vinci si surgical procedure the harmonic ace curved shears insert tip melted. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the nonconformance was a melted teflon pad. The teflon pad was found burnt and melted away in the garage track. The material was removed. The clamp arm was still able to open/close, but the insert was not fully functional without the teflon pad. The clamp arm was found with black marks. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC ACE CURVED SHEARS INSTRUMENT ACCESSORY | ENDOSCOPIC ACCESSORY | 3032929 | 3007530 | 2904068 | 2955842201300847 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | Yes | 2013-02-14 | Device Not Retu | No | Reuse | 2013
40027202 | 2012-10-30 | Malfunction |  | It was reported that during a da vinci si surgical procedure the harmonic ace curved shears insert tip melted. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the nonconformance was a burnt clamp arm with a missing teflon pad. The teflon pad was found missing. The clamp arm was still able to open/close, but insert was not fully functional without the teflon pad. The clamp arm was found with black marks. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC ACE CURVED SHEARS INSERT INSTRUMENT ACCESSO | ENDOSCOPIC ACCESSORY | 3032962 | 3007554 | 2904068 | 2955842201300848 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-01-31 | Injury | Required Intervention | It was reported to isi that a few days after a da vinci si pyloroplasty procedure, the patient experienced problems with high fever. After consulting with his surgeon, he was brought back to the emergency room for an exploratory surgery. It was alleged that during the open traditional surgery a 3cm hole was found on the patient's bowel. It was reported that the surgeon removed 6 of the patient's colon and a colostomy bag was installed. The patient was later transferred from the intensive care unit to a private room. | Isi contacted initial reporter clinical sales representative, he indicated that the patient was released from the hospital. Intuitive surgical contacted the hospital to request more information concerning the reported event; however, as of the date of this report no response has been received. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3032967 | 3007560 | 2904074 | 2955842201300849 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-13 | 2013-03-15 | Yes | Yes | Health Professional | IS3000 A60P8 | No | No | 2013-02-13 | No | No | Reuse | 2013
42030901 | 2013-02-15 | Malfunction |  | It was observed during central processing that the mega suturecut needle driver instrument had a broken cable. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken cable. One grip close cable was found to be broken near the proximal pulleys and prox clevis cable hole. The idler pulley was able to spin freely and did not exhibit damage. The cable segment was observed to be sticking out at the instrument's wrist. Engineering evaluation also found prox clevis wearing. The prox clevis cable hole exhibited wear on one edge. Engineering evaluation concluded that the wear on the hole may have contributed to the cable breakage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3032995 | 3007567 | 2904081 | 2955842201300850 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-15 | 2013-03-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-15 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-02-07 | Malfunction |  | It was reported that during a da vinci si surgical procedure that the maryland bipolar forceps instrument knuckle could not be moved probably because a wire was broken. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument knuckle couldn't be moved. Both of the yaw pulleys exhibited charring and localized melting at the grip base. The charring may have been due to a breach in the insulation. No broken cables were found. Electrical continuity testing passed. Internal laboratory testing has determined that the arc track failure mode is most likely to occur when a bipolar instrument is energized with no tissue between the grips. Internal arcing in bipolar instruments results in a potential loss of product function, but no transfer of energy to the patient. Intuitive has concluded that the likelihood of injury due to this failure is remote. User are instructed to retain back-up instruments in case that an instrument fails to perform, and are also instructed in proper technique t o minimize the likelihood of internal arcing. Engineering also found various scratch marks on the distal end of the maintube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3032976 | 3007574 | 2904088 | 2955842201300835 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-13 | 2013-03-15 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2013-02-13 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-14 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si myomectomy procedure, the surgical staff noticed that the fenestrated bipolar forceps instrument stopped working and that the wire at the wrist of the instrument fractured, causing a piece of the fractured wire to fall into the patient's pelvis. The broken fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the pitch up cable was broken at the distal clevis hub. In addition, the cable segment that contains the crimp was also observed to be missing. The clevis did not exhibit any damage or wear marks. The other cables at instrument's wrist were not found to be damaged. Engineering evaluation also found tube damage on the instrument. The distal end of the instrument's main tube had various scratch marks with light material removal. The scratches were. 075 -. 215 in length and not aligned with the tube axis. Engineering evaluation concluded that the scratches were likely caused by mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable and main tube scratch issues if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- On (b)(4)2013, the initial reporter contacted isi and provided the following additional information regarding the reported event: the instrument was inspected prior to use and no problems were noted. She could not recall if any external collisions occurred during the procedure with the instrument. She also did not know if the instrument was removed at any time during the procedure. The initial reporter mentioned that the time of the procedure was extended since a fragment had to be retrieved. However, she could not recall how much longer the procedure lasted. The initial reporter denied that any x-rays or other imaging was taken as a result of the reported issue. According to the initial reporter, the patient recovered uneventfully. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | Reuse | 3033062 | 3007641 | 2904155 | 2955842201300851 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2013-02-15 | 2013-03-15 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-02-15 | Device Not Retu | No |  | 2013
42018907 | 2013-02-14 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si myomectomy procedure, the surgical staff noticed that the wire at the wrist on the double fenestrated grasper instrument fractured, causing a piece of the wire from the instrument to fall into the patient. The broken fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable at the instrument's wrist was broken. The clevis did not exhibit any damage or wear marks. The cable segment that contains the crimp was also observed to be missing. Intuitive surgical has made several attempts to contact the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2013, the initial reporter contacted isi and provided the following additional information regarding the reported event: the instrument was inspected prior to use and no problems were noted. She could not recall if any external collisions occurred during the procedure with the instrument. She also did not know if the instrument was removed at any time during the procedure. The initial reporter mentioned that the time of the procedure was extended since a fragment had to be retrieved. However, she could not recall how much longer the procedure lasted. The initial reporter denied that any x-rays or other imaging was taken as a result of the reported issue. According to the initial reporter, the patient recovered uneventfully. | DOUBLE FENESTRATED GRASPERINSTRUMENT | Reuse | 3033061 | 3007648 | 2904162 | 2955842201300853 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2013-02-15 | 2013-03-15 | No | Yes | Health Professional | 42018907 | Device Returned To Manufacturer | Yes | 2013-02-15 | Device Not Retu | No |  | 2013
IS3000 A60P7 | 2013-01-17 | Injury | Required InterventionLife Threatening | It was reported that approximately two weeks after successful completion of a da vinci si single site cholecystectomy procedure performed on (b)(6) 2013, the patient experienced post-surgical complications as a result of an alleged bowel perforation. Reportedly, the patient underwent a surgical procedure using open surgical techniques to repair the affected area. | On (b)(4) 2013, isi contacted the isi clinical sales representative (csr) who reported this event. The csr indicated that he was present during the entire surgical procedure and that he did not observe any malfunction of the site's da vinci si surgical system, instruments and/or accessories. The csr indicated that the procedure was successfully completed and there were no intra-operative complications experienced by the patient. The csr indicated that after completion of the surgical procedure, the surgeon indicated to him that he (surgeon) felt that the surgery went well. The csr indicated that he spoke to the hospital's house physician concerning this event and the house physician indicated to him that the damage found to the patient's bowel was a clean cut and did not have the appearance of damage caused by cautery energy. The csr indicated that he also spoke to the surgeon regarding this event. Per the csr, the surgeon indicated to him that it is unknown if damage to the patient's bowel occurred during creation of the port site incision or if the patient injury occurred during the surgical procedure. The csr indicated that the surgeon indicated to him that the patient has been discharged from the hospital and is recovering. No other details concerning the patient were provided by the surgeon. Isi has concluded that based on the information provided by the site, the root cause of the bowel injury sustained by the patient is indeterminable. If isi receives additional information concerning this event, a follow up medwatch report will be submitted to the fda. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3033063 | 3007649 | 2904163 | 2955842201300852 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-13 | 2013-03-15 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2013-02-13 | No | No | Reuse | 2013
42003306 | 2013-02-08 | Malfunction |  | It was reported that during a vasovasostomy procedure using the da vinci si surgical system, it was noted that a little piece of cable was observed at the top of the black diamond micro forceps instrument and that the instrument stopped working. No missing or fallen pieces were reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch down cable is lightly frayed at the distal clevis hub. Small frayed strands stick out at the wrist. The clevis hub exhibits scratch marks adjacent to the fraying. The other cables at the wrist were not damaged. Engineering evaluation also found that the distal end of the main tube had various scratch marks with light material removal. The scratches were. 075 -. 150 in length and were not aligned with the tube axis. Engineering concluded that the damage to the main tube were likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | BLACK DIAMOND MICRO FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3033092 | 3007659 | 2904173 | 2955842201300854 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-13 | 2013-03-15 | No | Yes | Health Professional | 42003306 | Device Returned To Manufacturer | Yes | 2013-02-13 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-18 | Malfunction |  | It was observed during central processing that the pk dissecting forceps instrument had a frayed wires. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033096 | 3007665 | 2904179 | 2955842201300855 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-18 | 2013-03-15 | No | Yes | Health Professional | 42022703 | No | No | 2013-02-18 | No | No | Unkown | 2013
IS3000 A60P8 | 2013-02-13 | Injury | Required Intervention | It was initially reported that during a da vinci si sigmoid colectomy procedure, the customer experienced a problem with one of the patient side manipulator (psm) arms. The customer reported that the psm arm 1 was not following the movements of the surgeon properly during the case. It was alleged that the bed side assistant was looking at the drapes when the arm moved forward and contacted the patient's aorta. The customer converted to an open procedure and a vascular surgeon was called to repair the damage. It was also reported that the site was unsure if the arm was touched or if the surgeon had his head in the viewer with his hands off the master. The surgeon reported to the clinical sales representative: it is clear to me that there was pent up energy stored on the arm of the robot. When this was released the instrument shot forward with a force and at a speed greater than humanly possible. It then appeared to set back half way between the starting point (where it was not functioning) and the point of maximal contact in the retroperitoneum. I am not sure whether the recoil after the stabbing was the robot returning to a rest or me instinctively pulling back on the controls. The instrument involved in the incident was the harmonic ace instrument in port 1 in the rlq (the port itself was pushed in beyond the remote center because of the patient's body habitus and the need for additional length while addressing the flexure). It had been working normally until about 30 seconds prior to the incident when it stopped responding to commands at which point isi technical support had the or staf check the field and she identified no physical barriers and noted there were no error messages. The other active instrument was in port 3 in the subxyphoid area and was functioning normally throughout. The third port (2) in the rlq was inactive with a prograsp forceps instrument gently retracting small bowel in the pelvis. The 2 active instruments were clearly within view the whole time. | The field service engineer (fse) visited the site to inspect the patient side manipulator (psm). He performed a test drive and sine cycle on the psm with no issues. The fse also inspected the instruments and the cannulae that were used during the procedure. The fse inspected the cannulae to ensure that all were within specification. The fse found one out of specification. Site confirmed that this cannula would be taken out of rotation. According to the surgical staff, all cannulae are inspected before they are returned back into circulation. The fse also inspected the instruments used during the procedure. The harmonic ace instrument exhibited long scratches and gouge marks on it consistent with use of the out of specification cannula. Surgical staff confirmed that this instrument would be also taken out of rotation. The fse inspected the psm cable tension, and adjusted to specification. The fse was unable to reproduce the exact customer reported failure. According to the information provided to the fse by the surgical staff, the patient condition was stable. No report of any further complications was made. A review of the system logs showed no system malfunction occurring during the procedure on (b)(6) 2013. As of(b)(6) 2013, there have been no reported recurrence of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3033175 | 3007725 | 2904239 | 2955842201300857 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-13 | 2013-03-15 | Yes | Yes | Health Professional | IS3000 A60P8 | No | No | 2013-02-13 | Device Not Retu | No | Reuse | 2013
42004906 | 2013-02-12 | Malfunction |  | Event Description: --/-- On (b)(4) 2013, we received a user facility medwatch from the fda with the following information: '8mm cardiere forceps instrument broke during use. Wire at pulley system frayed and exposed. '. --/-- Additional Event Description: --/-- It was reported that during a davinci s surgical procedure, the customer reported that the pulley broke on the cadiere forceps instrument. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp fell out of the clevis, and it was returned with the instrument. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- (b)(4). | CADIERE FORCEPS INSTRUMENT | Reuse | 3033161 | 3007749 | 2904263 | 2955842201300856 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-03-19 | 2013-03-15 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-03-19 | Device Not Retu | No |  | 2013
42022703 | 2013-02-06 | Malfunction |  | It was reported that during reprocessing of the pk dissecting forceps instrument it was noted that the tips on the instrument were not aligned. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was bent, causing side to side misalignment of the grips. There was a. 040 offset at the tips. The bent grip did not exhibit separation of the yaw pulley and pulley cover at the glue joint, however, engineering concluded that damage to the grip was likely due to overloading at the tip of the instrument. Engineering evaluation also found that the distal end of the main tube had various scratch marks with light material removal. The scratches were. 100 -. 175 in length and were not aligned with the tube axis. Engineering concluded that damage to the main tube was likely caused by mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033201 | 3007797 | 2904311 | 2955842201300858 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-13 | 2013-03-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-13 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-15 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the insulation on the monopolar curved scissors (mcs) instrument was damaged. No missing or fallen parts were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the tube extension was cracked at the distal end and there were no missing pieces. A piece of tube extension was sticking out at proximal clevis. Engineering concluded that the damage to the tube extension was likely due to excessive side loading or other type of mishandling. Engineering evaluation also found that the distal end of the main tube has various scratches that exhibited light material removal and had a rough surface finish. The scratches varied in size but were mostly small and thin. Engineering concluded that damage to the main tube was likely due to mishandling. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's main tube found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033220 | 3007809 | 2904323 | 2955842201300863 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-15 | 2013-03-15 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-02-15 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-10-17 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the tenaculum forceps instrument behaved in a floppy manner and did not respond. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable at the distal end was broken. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. There were also several wear marks on the grips. No other damage was found the customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3033233 | 3007829 | 2904343 | 2955842201300859 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-13 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, the surgeon observed that the wire on the prograsp forceps instrument broke and the instrument stopped working. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode; however, for clarification, the instrument malfunction was a frayed grip cable. The frayed cable sections were in various lengths and were sticking out at the wrist. The idler pulley exhibited rim and face damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENTS | ENDOSCOPIC INSTRUMENT | 3033234 | 3007830 | 2899442 | 2955842201300860 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-14 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the end of the megasuturecut needledriver instrument broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with a frayed cable at the distal idler pulley. The frayed cable section was approximately 0. 30 in length. The grip cable was also found derailed at the distal pulley, with excess damage on the edge of the distal pulley. Engineering concluded that damage to the pulley was likely due to excess force contact and mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3033235 | 3007831 | 2904345 | 2955842201300861 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-11 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the wire on the pk dissecting forceps instrument was sticking out. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp is remained installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. Rument has 4 uses remaining. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3033219 | 3007832 | 2904346 | 2955842201300862 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-02-14 | 2013-03-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-08-07 | No | No | Reuse | 2013
400077 | 2013-03-01 | Malfunction |  | Completed robotic assisted laparoscopic cholecystectomy and discovered (2) green tabs that were originally attached to green port seals became dislodged during case while inserting instruments through port. These were discovered and retrieved by surgeon at end of case. What was the original intended procedure?robotic assisted laparoscopic cholecystectomy. Device #1is this a laboratory device or laboratory test?no. |  |  | LAPAROSCOPE | 3065081 | 3038079 | 2934593 | 3038079 | 1 | GCJ | User Facility |  | Nurse | Initial | 2013-03-15 | 2013-03-15 | No | Yes | Nurse | 400077 | No | Yes |  |  |  |  | 2013
IS2000 S A51P8 | 2013-02-12 | Injury | Required InterventionLife Threatening | It was reported that post a da vinci s hysterectomy procedure, it was discovered that the patient sustained rectal/colon injuries, which required the patient to undergo additional surgical procedures. | On (b)(6) 2013, isi contacted the isi clinical sales representative (csr) assigned to the site. The csr indicated that he spoke to the (b)(6) at the hospital. Per the csr, the (b)(6) indicated to him that the da vinci s hysterectomy procedure was successfully completed and that there were no intra-operative complications experienced by the patient. The csr indicated that the or director indicated to him that the patient developed unspecified post-operative complications; however, the (b)(6) was unable to provide the date/time that the patient began to experience the post-operative complications. The csr indicated that the or director indicated to him that the date of the surgical procedure is unknown; however, no malfunction of the da vinci s surgical system, instruments, or accessories occurred during the surgical procedure. The csr indicated that the (b)(6) indicated to him that the patient underwent a laparoscopic procedure at the hospital and it was discovered that sutures were stuck to the patient's large colon and that upon removal of the sutures, a perforation of the patient's large colon was discovered and that the affected area was repaired. The csr indicated that the (b)(6) was unable to provide him with information regarding the current status of the patient. Based on the information provided by the site, it is indeterminable as to what caused or contributed to the injury sustained by the patient. Intuitive surgical has continued to request additional information from the site regarding this event; however, as of the date of this report no additional information or response has been provided. A follow-up mdr will be submitted to the fda if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3030612 | 3005215 | 2901749 | 2955842201300828 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-12 | 2013-03-14 | Yes | No | Health Professional | IS2000 S A51P8 | No | No | 2013-02-12 | No | No | Reuse | 2013
IS2000 A51P8 | 2013-01-18 | Injury | Required Intervention | It was reported that prior to starting a da vinci s prostatectomy procedure, during startup of the da vinci s system, the customer experienced a problem with error code 45341. The site contacted isi for technical support engineering (tse) assistance. Review of the site's system log by the tse found that the system error code 45341 was related to the video process assembly on the da vinci s robot system. The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm or injury was initially reported. | The field service engineer visited the site to inspect the da vinci robot system. He replaced the blue vision cable, and the camera head. The camera head produces three dimensional images to the da vinci standard vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system error code 45341 appears when the da vinci s safety system determines there is a loss of synchronization between the left and right video streams. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. Upon determining this condition, the safety system put da vinci in a recoverable safe state. The camera cable was inspected by the field service engineer; he was able to replicate the reported failure mode and observed that there was an internal short. The system was repaired by replacing the defective camera cable. Field service engineer tested the system and found it to be working to manufacturer's specifications with no defect or malfunction. The camera head was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported problem, however, the camera showed evidence of a mechanical shock. The camera was upgraded as a precaution. Fse reviewed the internal system error logs, and he also identified error code 23020 associated with the patient side manipulator (psm). System error code 23020 appears when the da vinci safety system determines one of the switches in a specific manipulator is showing inconsistent signals on two switch leads. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. Upon determining this condition, the safety system put da vinci in a recoverable safe state. The psm was replaced and returned to intuitive surgical for further inspection. The patient side manipulator is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The psm was returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported problem; however, the cannula switch, flat flex cable were replaced as a precaution. The cannula junction board was also upgraded. On (b)(6) 2012 the da vinci robot coordinator provided additional information regarding the reported event. She indicated that the planned surgical procedure was converted to open due to the vision issue. She also indicated that the patient required 3 units of blood; however, she didn't know the reason. She indicated that the patient was recovering well from the open procedure. As of (b)(6) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3030660 | 3005314 | 2901848 | 2955842201300829 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-12 | 2013-03-14 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-12 | Malfunction |  | It was reported that during a da vinci s cholecystectomy procedure, the cable on the large suturecut needledriver instrument broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3030748 | 3005346 | 2899507 | 2955842201300830 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-12 | 2013-03-14 | No | Yes | Health Professional | 42029602 | No | Yes | 2013-02-12 | No | No | Unkown | 2013
42022703 | 2013-02-11 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wires on the pk dissecting forceps instrument frayed and snapped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis does not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering evaluation also found that the distal pulley surface exhibited scratches. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENTS | ENDOSCOPIC ELECTROSURGICAL INS | 3030769 | 3005373 | 2904447 | 2955842201300831 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-13 | 2013-03-14 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-13 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-02-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tip of the small clip applier instrument was observed broken. No missing or fallen pieces were reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a. 125 x. 050 piece detached from one grip. The grip fractured towards the base of the clip groove and had a bend below the fracture surface. The intact grip was bent around the same location as the broken grip. The location of the fracture/bending aligns closely with the installation depth of the grips on clip rack. Engineering concluded that the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3030918 | 3005521 | 2902035 | 2955842201300832 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-13 | 2013-03-14 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-02-13 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2013-02-11 | Malfunction |  | It was initially reported that during a da vinci si prostatectomy procedure, the fenestrated bipolar instrument was firing without the surgeon's input. Isi technical support engineer (tse) advised the customer to check the valley lab generator settings to ensure that the auto fire button was not turned on. The da vinci coordinator nurse was not able to find the auto fire settings. The customer swapped to a backup generator; the prostatectomy procedure was successfully completed with no patient harm. | The clinical sales representative (csr) followed up with initial reporter da vinci coordinator nurse. She indicated that someone in the surgical staff accidentally turned on the auto fire switch on the valley lab generator. They were able to switch to a backup generator to complete the prostatectomy procedure. The procedure was successfully completed with no patient harm. The patient was discharged from the hospital. The da vinci si instructions for use (ifu) specifically states: caution: inadvertent electrosurgical energy may cause serious injury or surgical complications to the patient. It is important to ensure a full understanding of the da vinci surgical system energy user interface and use caution when working near critical anatomy. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3027330 | 3001799 | 2898333 | 2955842201300808 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-11 | 2013-03-13 | No | Yes | Health Professional | IS3000 A60P6 | No | Yes | 2013-02-11 | No | No | Reuse | 2013
42031802 | 2013-02-11 | Malfunction |  | It was reported that during a da vinci si low anterior resection procedure, the wire on the small grasping retractor instrument broke. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was broken at the distal idlers. The idler pulley spun freely and did not exhibit damage. The cable segment sticks out at the wrist. The other cables at the wrist were not damaged. Engineering evaluation also found that the distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 165 in length and were not aligned with the tube axis. Engineering concluded that the damage to the main tube was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3027590 | 3002146 | 2898680 | 2955842201300809 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-11 | 2013-03-13 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-02-11 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-01-15 | Malfunction |  | It was reported that during a davinci s surgical procedure, the customer noted that wires were protruding from the mega needle driver instrument. Nothing fell into the patient and no harm to the patient occurred. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument had a frayed grip cable. The frayed cable section was sticking out at the wrist. There was no damage found on the pulley. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3028012 | 3002461 | 2898995 | 2955842201300810 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-13 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
400158 | 2013-02-11 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the snap-fit paddle blade accessory installed on the snap-fit scalpel instrument fell into the patient. The blade accessory was retrieved and no other fallen pieces were reported. The planned surgical procedure was completed. No patient harm or adverse outcome was reported. | On (b)(4) 2016, isi contacted the isi clinical sales representative (csr) who initially reported this event for additional information. The csr indicated that he was present during the surgical procedure when the reported failure mode occurred. The csr indicated that the snap-fit scalpel instrument and snap-fit paddle blade were inspected by the surgical staff prior to use and there were no issue noted. Per the csr, it was the site's first time utilizing the snap-fit scalpel instrument for da vinci surgical procedures. The csr denied that the snap-fit scalpel instrument collided with any other instruments during the procedure. The csr indicated that an x-ray performed on the patient showed the location of the blade accessory and the blade accessory was retrieved, no additional surgical procedure was required to remove the blade accessory and the planned surgical procedure was completed. The csr indicated that after completion of the surgical procedure, he installed a new blade accessory of the same type onto the snap-fit scalpel instrument and that he was unable to replicate the reported issue. The csr indicated that the blade accessory fit properly onto the snap-fit scalpel instrument and that he did not note any issues when he installed the blade accessory onto the instrument. The csr also indicated that his inspection of the instrument and the initial blade accessory found no damages. The csr indicated that the issue experienced by the site may have caused by improper installation of the blade accessory onto the instrument. The csr indicated that the site has declined to return the affected blade accessory and instrument, since no malfunction of the devices were found. The csr indicated that the site reported to him that there was no harm to the patient and the patient has not experienced any post-surgical complications as a result of the reported event. | SNAPFIT PADDLE BLADE ACCESSORY | ENDOSCOPIC INSTRUMENT ACCESSOR | 3028152 | 3002546 | 2899080 | 2955842201300811 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-12 | 2013-03-13 | Yes | Yes | Health Professional | 400158 | No | No | 2013-02-12 | No | Yes | Initial | 2013
41029902 | 2013-02-11 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the gray part of the suction irrigator instrument's tip broke off fell into patient. The broken piece was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with the distal tip and inner lumen dislodged from snake wrist disk. The missing distal tip that fell into patient was returned, it was taped to suction irrigator instrument main tube. The instrument's snake wrist also exhibited a dislodged disk. Engineering concluded that damage was likely due to excessive side loading or other type of mishandling. Instrument has no remaining uses left. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | SUCTIONIRRIGATOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3028419 | 3002737 | 2899271 | 2955842201300814 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-13 | No | Yes | Health Professional | 41029902 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-04 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the wire on the megasuturecut needle driver instrument separated. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with one grip close cable broken near the proximal pulleys and proximal clevis cable hole. The idler pulley spun freely and did not exhibit damage. The cable segment stuck out at the wrist. Engineering evaluation also found, that the proximal clevis cable hole exhibited wear on one edge. The other cables at the wrist were not damaged. Engineering concluded that wear on the hole was likely due to the cable breakage. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3028347 | 3002767 | 2899300 | 2955842201300813 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-11 | 2013-03-13 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-10 | Malfunction |  | It was reported prior to starting a da vinci si hysterectomy procedure, a little piece from the top of the fenestrated bipolar forceps instrument fell off and the instrument had limited movement. There was no report that any piece(s) from the fenestrated bipolar forceps instrument fell into the patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that one yaw pulley was missing a. 160 x. 060 piece on one side, allowing the grip to move within the pulley and have a larger range of motion in yaw. Engineering concluded that it was indeterminable as to how the piece of yaw pulley broke. Engineering evaluation also found that one wire was detached from its connection at the grip. The wire was broken on the same yaw pulley that broke; however, the location of the wire breakage was located on the opposite side. The instrument also passed electrical continuity testing. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 3028466 | 3002855 | 2899388 | 2955842201300815 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-11 | 2013-03-13 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-11 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-11 | Malfunction |  | It was reported that during a davinci si cholecystectomy procedure, the prograsp forceps instrument's cable broke. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks 0. 095 - 0. 250 with light material removal. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3028507 | 3002909 | 2899442 | 2955842201300816 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-04 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the fenestrated bipolar forceps instrument was not being recognized by the da vinci robot system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was checked for recognition using the in-house is3000 da vinci robot system. The system successfully recognized the instrument, showing 7 uses left. The pogo pins did not stick and were not contaminated. Engineering was unable to replicate the recognition issue. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks 0. 065 - 0. 120 with light material removal. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3028563 | 3002945 | 2899478 | 2955842201300817 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-13 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-01-12 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy the monoploar curved scissors instrument tip burned out and or quit working. Another monopolar curved scissors instrument was used to complete the surgical procedure. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the tip was burned. The instrument was found with a char mark on the extension tube. The char mark measured roughly about. 132 x. 278. The mcs tip cover was not returned for failure analysis. Instruments passed electrical continuity test. Additional observation not reported by the site was pad printing removal. Next to char mark on extension tube, found pad print was removed. The missing print measured roughly. 081 x 0. 422. Damage to the print pad was most likely due to improper cleaning. 62 - the instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. Follow up attempts have been made to contact the site regarding the mcs tip cover; however, no additional information has been provided at this time. A follow-up mdr will be submitted if the instrument the mcs tip cover is returned (post engineering evaluation) or if additional information is received. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3028570 | 3002950 | 2899483 | 2955842201300818 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2013-02-11 | 2013-03-13 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-02-11 | Device Not Retu | No | Reuse | 2013
42805206 | 2013-02-07 | Malfunction |  | It was reported that during a da vinci si single site surgical procedure, the surgical technician experienced difficulty removing the a-sure cautery hook accessory installed on the monopolar cautery instrument, causing the tip of the monopolar cautery instrument to break. No missing or fallen pieces were reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3028578 | 3002961 | 2899494 | 2955842201300812 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-11 | 2013-03-13 | No | Yes | Health Professional | 42805206 | No | No | 2013-02-11 | Device Not Retu | No | Unkown | 2013
42009308 | 2012-11-13 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the prograsp forceps instrument had a broken wire at the tip of the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken wire at the tip. The one grip open cable was frayed near the force amp pulley and the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 3028589 | 3002970 | 2899503 | 2955842201300819 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-14 | 2013-03-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-02-12 | Malfunction |  | It was reported that during a da vinci s cholecystectomy procedure, the cords on the large suturecut needledriver instrument broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable at the distal idlers was broken. The idler pulley spun freely and was not damaged. The cable segment was sticking out at wrist. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3028592 | 3002974 | 2899507 | 2955842201300820 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-12 | 2013-03-13 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | Yes | 2013-02-12 | Device Not Retu | No | Initial | 2013
42022703 | 2013-02-13 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the pk dissecting forceps instrument wire broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument wire was broken. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. Engineering also found various scratch marks on the distal end of the maintube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3028593 | 3002976 | 2899509 | 2955842201300821 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-13 | 2013-03-13 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-02-13 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-04 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure that the 8mm fenestrated bipolar forceps instrument was not working. The surgical staff was unable to get the instrument seated on the robot arm and the light would not change from yellow. No missing or fallen pieces were reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. Engineering also found various scratch marks on the distal end of the main tube that showed light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. The last finding was scratches on the surface of both of the distal pulleys. Engineering concluded that the damage to the main tube was likely due to mishandling. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3028597 | 3002977 | 2899510 | 2955842201300822 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-13 | 2013-03-13 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-13 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-12 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure that the fenestrated bipolar forceps instrument cable was noted to be broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was broken. The pitch cable was broken at the distal end. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. The electrical continuity testing passed. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3028603 | 3002982 | 2899515 | 2955842201300823 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-13 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-11 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the prograsp forceps instrument was attached to the robot arm and after entering the patient's abdomen, the system alarmed. The instrument was removed and was noted to have frayed fibers at the articulating neck of the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was frayed. The one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3028518 | 3002984 | 2899517 | 2955842201300825 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-12 | 2013-03-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-12 | Malfunction |  | It was reported that during reprocessing the 8mm large needle driver instrument was noted to have broken wires. No missing or fallen pieces were reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found there were indentations and scratches on the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3028605 | 3002985 | 2899518 | 2955842201300824 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-12 | 2013-03-13 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-12 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the pk dissecting forceps instrument was not recognized by the system. Reportedly, during reprocessing of the instrument, the cable was noted to be broken. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as functional testing of the instrument found that it successfully passed recognition testing on an is3000 in-house system. Functional testing also found that the instrument had 3 uses remaining. The pogo pins did not stick and were not contaminated. The instrument passed electrical continuity testing. Engineering evaluation also found that the pitch cable at the distal clevis hub was broken. The cable segment that contains the crimp remained installed in the clevis and the clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3028610 | 3002992 | 2899525 | 2955842201300826 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-13 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-28 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the pk dissecting forceps instrument would not dock because of a bent prong. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue that the instrument would not dock properly because of a bent prong. The upper left pin on the 5 pin pk connector was found to be bent downwards. A pk cord could not be properly installed due to the bent pin. Electrical continuity testing of the instrument passed. Engineering evaluation also found tube damage. The distal end of the instrument's main tube had various scratch marks with light material removal. The scratches were. 070 -. 080 in length and not aligned with the tube axis. Engineering evaluation concluded that the scratches were likely due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Intuitive surgical has attempted to contact the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported main tube scratch issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3028620 | 3003002 | 2899535 | 2955842201300827 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-13 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-02-08 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wires on the prograsp forceps instrument were frayed and the instrument would not articulate. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the customer reported complaint; however, for clarification, the failure mode was a broken pitch cable. The pitch cable was broken at the wrist. The cable segment that contains the crimp remained installed in the clevis and the clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3025680 | 2999839 | 2896373 | 2955842201300798 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-11 | 2013-03-12 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-11 | Device Not Retu | No | Reuse | 2013
42011006 | 2012-10-05 | Malfunction |  | It was reported that during a davinci si prostatectomy procedure, the customer noted that the clamp was dislocated at the base of the precise bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. Additional observations not reported by site were main tube damage and bipolar pin bent. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. The bottom bipolar pin was also found to be bent towards the top pin. Engineering concluded that damages were likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PRECISE TM BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3026333 | 3000506 | 2897040 | 2955842201300799 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-12 | 2013-03-12 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | Yes | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42011006 | 2013-02-06 | Malfunction |  | It was reported that during a davinci si prostatectomy procedure, the customer noted that the clamp was dislocated at the base of the precise bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable located at distal clevis hub. No damage was found at the clevis. Frayed segment was approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PRECISE TM BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3026384 | 3000578 | 2897112 | 2955842201300800 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-12 | 2013-03-12 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | Yes | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-28 | Malfunction |  | It was reported that during da vinci si hysterectomy procedure the 8mm large needle driver instrument kept sticking in the open position while in use. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of needle driver sticking in the open position. Instrument placed on the is3000 system and driven. Instrument recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The functional performance testing did not find any issues. Engineering found various scratch marks on the main tube with light material removal on the distal end. The scratches were. 060 -. 150 in length and not aligned with the tube axis. Engineering concluded that the main tube damage was likely due to mishandling. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3026541 | 3000704 | 2897238 | 2955842201300802 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-12 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42018907 | 2013-02-12 | Malfunction |  | The double fenestrated grasper instrument was returned without any specific allegations of malfunction. No information regarding the details of when the breakage occurred were specified. No further information was received. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | DOUBLE FENESTRATED GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3026557 | 3000705 | 2897239 | 2955842201300803 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-12 | 2013-03-12 | No | Yes | Health Professional | 42018907 | Device Returned To Manufacturer | Yes | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-06 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure the megasuturecut needle driver instrument cable snapped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a snapped cable. The one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3026540 | 3000717 | 2885067 | 2955842201300801 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-11 | 2013-03-12 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-02-11 | Device Not Retu | No | Reuse | 2013
42004906 | 2013-01-31 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure that the cadiere forceps instrument was not recognized by the system. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found various scratch marks with light material removal at the distal end of the main tube. The scratches were. 080 -. 190 in length and not aligned with the tube axis. Engineering concluded that damage to the main tube was likely due to mishandling. No other damage found. Per the customer reported complaint engineering evaluation found that the instrument was checked for recognition on an in-house is3000 system and the system successfully recognized the instrument; it showed 0 uses left and displayed an instrument is expired message as expected. The pogo pins did not stick and were not contaminated. Engineering was not able to replicate the recognition issue. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3026610 | 3000744 | 2897278 | 2955842201300805 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-12 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-02-12 | Malfunction |  | It was reported during reprocessing that the permanent cautery spatula was noted to have insulation around the head of the instrument scratched away. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the insulation around the head of the instrument is scratched away. Pieces of the ceramic sleeve are missing around spatula tip. No signs of charring on the tip. Engineering concluded that damage was likely due to mishandling. Electrical continuity testing passed. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3026581 | 3000772 | 2897306 | 2955842201300804 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-12 | 2013-03-12 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-12 | Malfunction |  | The megasuturecut needle driver instrument was returned without any specific allegations of malfunction. There were no missing or fallen pieces reported. Information regarding the details of when the breakage occurred were not specified. No further information was provided. | The instrument was returned and evaluated. No complaint was reported. Engineering found the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the edge and on the surface. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3026582 | 3000773 | 2897307 | 2955842201300806 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-12 | 2013-03-12 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-03 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the prograsp forceps instrument was noted to have a broken wire at the tip. The instrument was exchanged during the procedure. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was frayed. The one grip close cable was frayed at the force amp pulley. The frayed strands stuck out at the wrist. The other cables at wrist were not damaged. Engineering also found the cable was derailed cable and main tube damage. The same frayed grip close cable was derailed at the distal idler pulley. The grips could still open and close, but movement may not be precise. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 085 -. 125 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3026583 | 3000774 | 2897308 | 2955842201300807 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-14 | 2013-03-12 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-02-14 | Device Not Retu | No | Reuse | 2013
|  | Injury | Required Intervention | On (b)(4) 2012, isi received a legal filing concerning a patient who underwent a da vinci si hysterectomy surgery. The legal filing alleged that the patient sustained damage to her uterus, 'the patient continues to suffer from abdominal pain, pelvic pain, dyspareunia, bloating, abdominal distention, fatigue and decreased energy and stamina. Due to the injury sustained to her vaginal cuff and bowel during the davinci robotic hysterectomy, patient had to have multiple additional medical tests and physician consultations and has suffered pain, loss of function, emotional distress, and permanent injury. Patient has suffered the loss of consortium. '. | There was no hospital name/contact information provided for the reported event. Therefore, a follow up on the alleged patient's condition and event description could not be made. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3023701 | 2997910 | 2894444 | 2955842201300790 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-19 | 2013-03-11 | Yes | No | Health Professional |  | No | No | 2012-12-19 | No | No | Reuse | 2013
42022703 | 2013-02-06 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the pk dissecting forceps instrument was not recognized by the system. The instrument was reportedly not used in the planned procedure. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results could not confirm the alleged issue that the system did not recognize the instrument. The instrument was checked for recognition on an (b)(4). The system successfully recognized instrument. The pogo pins did not stick and were not contaminated. Dallas chip programmer (dcp) verification of the instrument passed. The alleged recognition issue could not be replicated. Engineering evaluation also found a broken cable, bent grip, and tube damage. The pitch down cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The pitch up cable was found to be frayed at the clevis hub. One grip was found to be bent causing side to side misalignment of the grips. There was a. 045 offset at the tips. The bent grip had separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. In addition, the distal end of instrument's main tube had various scratch marks with light material removal. The scratches were. 085 -. 190 in length and not aligned with the tube axis. Engineering evaluation concluded that the scratches were likely due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable and main tube scratch issues if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3024076 | 2998171 | 2894705 | 2955842201300791 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-11 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-14 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgical staff allegedly noticed that the mega suturecut needle driver instrument had a broken cable at the tip. No patient harm or adverse outcome was reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken cable. One grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely. The cable segment was reportedly sticking out at the instrument's wrist. Other cables at the wrist were not damaged. Engineering evaluation also found the idler pulley to have damages on the edge and on the surface. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3024195 | 2998286 | 2894820 | 2955842201300792 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-03-11 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-01 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the surgical staff allegedly noticed that the fenestrated bipolar forceps instrument would no longer work. No patient harm or adverse outcome was reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue that the instrument would no longer work. One grip was found to be bent, causing side to side misalignment of the grips. There was a. 142 offset at the tips, indicating overloading at the tip. Electrical continuity testing of the instrument passed. Engineering evaluation also found a scratch on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering evaluation concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issue if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3024291 | 2998379 | 2894913 | 2955842201300793 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-02-14 | 2013-03-11 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-21 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgical staff allegedly noticed that the prograsp forceps instrument had a wire that was derailed. No patient harm or adverse outcome was reported. The planned surgical procedure was completed with a backup instrument and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a derailed cable. Engineering evaluation found a derailed grip cable at the distal idler pulley. The grips could not fully open and close. Movement and functionality of the instrument was not precise. The instrument was returned expired. Engineering evaluation also found the derailed grip cable to be frayed at various sections. The customer reported complaint does not itself constitute a mdr reportable event; however the reported frayed cable issue if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3024279 | 2998391 | 2894925 | 2955842201300794 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-14 | 2013-03-11 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-08 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, the surgical staff allegedly noticed that the mega suturecut needle driver instrument had a broken wire. No patient harm or adverse outcome was reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of a broken wire. One grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and was not damaged. A cable segment was observed to be sticking out at the instrument's wrist. Other cables at the wrist were not damaged. Engineering evaluation also found an indentation on the edge of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3024318 | 2998440 | 2894974 | 2955842201300795 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-15 | 2013-03-11 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-15 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-28 | Malfunction |  | It was reported that during da vinci si lobectomy procedure, the prograsp forceps instrument was noted to be in the wrong position during removal of the instrument. The planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with the distal end assembly detached from the main tube at the proximal clevis and all of the distal end cables were cut. The proximal clevis exhibited damage on one side, where it meets the main tube. Engineering concluded that the issue experienced by the site, was likely caused by the dislodged proximal clevis at the main tube interface due to excessive loading at the distal end and that the cables were likely cut to allow for removal of instrument from cannula. Engineering evaluation also found that one side of the main tube interface wall was broken and missing pieces that were likely related to dislodgment of the clevis. The distal end of the main tube also had various scratch marks with light material removal. The scratches were. 070 -. 125 in length and were not aligned with the tube axis. Engineering concluded that the scratched on the main tube may have been caused by mishandling /misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, isi followed up with the isi clinical sales representative (csr) who was present during the surgical procedure. The csr indicated that no pieces from the instrument fell into the patient. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3024440 | 2998535 | 2895069 | 2955842201300796 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-11 | 2013-03-11 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-02-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-01 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable on the fenestrated bipolar forceps instrument broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with the pitch cable broken. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3024458 | 2998564 | 2895098 | 2955842201300797 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-11 | 2013-03-11 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-02-11 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-07 | Malfunction |  | It was reported that during a da vinci si inguinal hernia repair procedure, a wire of the large needle driver instrument snapped and fell into the patient. The wire was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken located at the distal idlers. The idler pulley spins freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional observations not reported by the site were indentations and scratches marks on edges and surface of the distal pulley. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3019903 | 2993763 | 2890297 | 2955842201300769 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-07 | 2013-03-08 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-07 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-02-07 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted a broken wire on the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken located at the distal idlers. The idler pulley spins freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional observations not reported by the site were indentations and scratches marks on edges and surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3019933 | 2993786 | 2890320 | 2955842201300770 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-07 | 2013-03-08 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-07 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-02-07 | Malfunction |  | The bowel grasper instrument was returned without an initial complaint. No reason for the returned was provided. | No reason for return was provided. Engineering observed that the instrument was returned with a dislodged cable and main insulation tube damage. The instrument snake wrist was found to be bent and it could not be straightened. A couple of wrist cables were also found to be kinked due to loose tension; however, not visibly broken or frayed cable was observed. Engineering removed the instrument's blue housing and a backend drive cable was found to be dislodged from the gimbal plate hole. The cable had slipped from the wedge feature and the wrist motion was not intuitive as a result. The main tube insulation was also found to be damaged at the distal end. The insulation was gouged on one side and exhibited a couple of 0. 040 x 0. 015 pieces missing roughly 4. 7 above the snake wrist. No other damage was found. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3020661 | 2994397 | 2891074 | 2955842201300771 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-07 | 2013-03-08 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-02-07 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-07 | Malfunction |  | The fenestrated bipolar forceps instrument was returned without an initial complaint. No reason for the return was provided. | No reason for return was provided. Engineering observed that the instrument was returned with a broken cable. The pitch down cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3020756 | 2994462 | 2890996 | 2955842201300772 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-07 | 2013-03-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-07 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-02-07 | Malfunction |  | The bowel grasper instrument was returned without an initial complaint. No reason for the returned was provided. | No reason for return was provided. Engineering observed that the instrument was returned with main insulation tube damage (bending and insulation gouged). Main insulation tube was found to be damaged located at the distal end. The insulation was gouged on one side and it exhibited various sections lifted up. There were a couple of exposed sections of the main tube, but exposed sections mate with lifted up section of insulation, likely no pieces are missing. Tube was also found to be bent roughly 6 above the snake wrist. Roll motion at distal end does not match proximal end due to bending. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | BOWEL GRASPER 5MM INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3020816 | 2994540 | 2891074 | 2955842201300773 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-07 | 2013-03-08 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-02-07 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-02-08 | Malfunction |  | It was reported that during a da vinci si surgical procedure the 5mm bowel grasper instrument was found to have the coating coming off the shaft of the instrument. The surgeon then found a piece of the instrument coating on the patient's liver. The instrument fragment was removed and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation observed the coating is coming off the shaft of the instrument. The main tube has a couple of different damaged sections with the tube insulation removed. The damaged areas are 1. 1 x. 100 and. 425 x. 085, located roughly 3. 8 and 2. 7 above the snake wrist. No other damage was found. | BOWEL GRASPER 5MM INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3020891 | 2994636 | 2891170 | 2955842201300775 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-08 | 2013-03-08 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-02-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-16 | Malfunction |  | It was reported that prior during set up of a da vinci si surgical procedure, the tips on the fenestrated bipolar forceps instrument did not meet. No missing or fallen pieces were reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as evaluation of the instrument found that the grip tips meet properly in a fully closed position. The grip tips did not exhibit any damage. Engineering evaluation also found that the pitch cable was broken at the wrist. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the damage to the instrument's pitch cable could likely cause or contribute to an adverse event if the malfunction recurred. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3021266 | 2994878 | 2861451 | 2955842201300776 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
IS3000 A60P2 | 2013-02-08 | Other |  | It was reported that the hospital has had numerous infections post da vinci surgical procedures. No further information was provided. | As of the date of this report no further information has been received regarding this allegation. No instrument or accessory has been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. Intuitive surgical contacted the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3021636 | 2995350 | 2887882 | 2955842201300779 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-02-08 | 2013-03-08 | Yes | Yes | Health Professional | IS3000 A60P2 | No | No | 2013-02-08 | No | No | Reuse | 2013
42017207 | 2013-01-25 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff allegedly noticed that the maryland bipolar forceps instrument had exposed cables. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of an exposed cable. One grip close cable was found to be loose at the distal clevis. A broken grip cable was found inside the instrument's housing which created the loose tension on the distal clevis. Other cables at the instrument's wrist were not damaged. Electrical continuity testing of the instrument passed. Engineering evaluation also found a loose conductor wire at the yaw pulley exit. The wire was sticking out from its connection at the grip. The yaw pulley did not show signs of arcing. In addition, engineering evaluation also found corrosion on the clamping pulley. Engineering evaluation concluded that the corrosion was most likely due to improper cleaning. Also, the distal end of the instrument's main tube was observed to have various scratch marks showing light material removal. The scratch was short in length and not axially aligned with the tube. Engineering evaluation concluded that the scratch was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported main tube scratch issue if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3021735 | 2995442 | 2891976 | 2955842201300781 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-08 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-08 | Malfunction |  | The fenestrated bipolar forceps instrument was returned without an initial complaint. No reason for the returned was provided. | No reason for return was provided. Engineering observed that the instrument was returned with main tube damage. The distal end of main tube exhibited various scratch marks with light material removal. Scratches were 0. 110 - 0. 220 in length and they were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3021846 | 2995546 | 2892080 | 2955842201300782 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-08 | 2013-03-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-31 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that wires were present at the tip of the fenestrated bipolar forceps instrument. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3021863 | 2995565 | 2892099 | 2955842201300783 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-08 | 2013-03-08 | No | Yes | Health Professional | 42020505 | No | No | 2013-02-08 | No | No | Unkown | 2013
42017910 | 2013-02-12 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci s surgical procedure, the surgical staff allegedly noticed that the monopolar curved scissors instrument had frayed wires. The reported issue was found during set up and was not used in the planned procedure. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of frayed wires. Engineering evaluation found a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | HOT SHEARS MONOPOLAR CURVEDSCISSORS INSTRUMENT | Reuse | 3021926 | 2995600 | 2892134 | 2955842201300785 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-02-12 | 2013-03-08 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42020505 | 2013-01-25 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the fenestrated bipolar forceps instrument was not being recognized by the davinci system robot. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was installed on an in-house da vinci robot system to be checked for recognition test. The da vinci robot system successfully recognized the instrument, showing 5 uses left. The pogo pins did not stick and they were not contaminated. Instrument passed electrical continuity test. The instrument was driven and moved intuitively, grips were able to open and close. Additional observation not reported by site was a frayed pitch cable at distal idler with scratches on the distal pulleys. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3021937 | 2995604 | 2892138 | 2955842201300786 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-08 | 2013-03-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-08 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-12 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci s surgical procedure, the surgical staff allegedly noticed that the monopolar curved scissors instrument had frayed wires. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the alleged issue of frayed wires. Engineering evaluation found a broken grip cable at the instrument's wrist. The idler pulley was able to spin freely. A cable segment was sticking out at the instrument's wrist. Other cables at wrist area were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | HOT SHEARS MONOPOLAR CURVEDSCISSORS INSTRUMENT | Reuse | 3021914 | 2995620 | 2892134 | 2955842201300784 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-02-12 | 2013-03-08 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42020706 | 2013-02-07 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted the cables of the tenaculum forceps instrument were frayed and broken and would not allow surgeon to move the wrist. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3021959 | 2995645 | 2892179 | 2955842201300774 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-07 | 2013-03-08 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-02-07 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-31 | Malfunction |  | It was reported that during a davinci si cholecystectomy procedure, the fenestrated bipolar forceps instrument was not being recognized by the davinci system robot. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was installed on an in-house da vinci robot system to be checked for recognition test. The da vinci robot system successfully recognized the instrument, showing 9 uses left. The pogo pins did not stick and they were not contaminated. Instrument passed electrical continuity test. The instrument was driven and moved intuitively, grips were able to open and close. Additional observations not reported by site were broken pitch cable and misalign tip. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. One grip was also found to be bent, causing side to side misalignment of grips. There was a 0. 535 offset at the tips, indicating overloading at the tip. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3021978 | 2995673 | 2892207 | 2955842201300787 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-08 | 2013-03-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-08 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-07 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the customer noted a frayed wire on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be frayed at the distal clevis hub. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3022011 | 2995699 | 2892233 | 2955842201300789 | 1 | NAY | Manufacturer | DistributorHealth Professional | Other | Initial | 2013-02-08 | 2013-03-08 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-08 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2013-02-08 | Other | Other | It was reported that during a da vinci s cardiac procedure, the surgeon was unable to focus one of the images through the high resolution stereo viewer (hrsv) on the surgeon side cart. The site troubleshot the vision issue by swapping the endoscopes and the camera head, however, the issue persisted. Unable to resolve the issue, the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the reported issue experienced by the customer was associated with the site's 30 degree and 0 degree endoscopes. During functional testing of the 30 degree endoscope the fse was unable to focus both of the eye channels at the same time and when the fse tested the 0 degree endoscope, he found that it had a slight glare spot in the right eye channel. The system endoscope projects light from the illuminator onto the surgical site and the video image of the surgical site captured by the endoscope is sent back through left and right channels to the camera head. The fse advised the customer to replace the affected endoscopes. The 30 degree endoscope was returned to the original equipment manufacturer (oem) for evaluation. The oem confirmed the vision issue experienced by the customer. The endoscope exhibited mechanical damage to the proximal end, which resulted in broken optical components and 3d optical misalignment. On (b)(4), 2013, isi contacted the isi clinical sales representative (csr) who initially reported this event and who was present during the surgical procedure. The csr indicated that she spoke to the surgeon concerning the patient's status a week after the reported event. The csr indicated that the surgeon reported to her that the patient tolerated the open procedure well and that the patient recovered well. As of (b)(6), 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3022012 | 2995700 | 2892234 | 2955842201300788 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-08 | 2013-03-08 | Yes | Yes | Health Professional | IS2000 A51P8 | Yes | No | 2013-02-08 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-07 | Malfunction |  | It was reported that during set up of a da vinci si surgical procedure the wire at the distal end of the pk dissecting forceps instruments was noted to be frayed. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3047456 | 3000097 | 2896631 | 2955842201300780 | 1 | NAY | Manufacturer | Health ProfessionalOther | Other | InitialFollowup | 2013-02-07 | 2013-03-08 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-03-13 | Yes | No | Reuse | 2013
42030901 | 2013-02-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the cable on the megasuturecut needledriver instrument was found to be broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that on grip cable was broken at the distal clevis and there was damage on idler pulleys. There were many dents along the edges of the idler pulley, which likely caused the cable to break. Engineering concluded that the dents on the idler pulley were likely caused by mishandling. The broken cable segment stuck out at the wrist approximately. 5155". No other damage was found. The instrument and accessories user manual specifically states: general precautions and warnings. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT NAY | 3047431 | 3000115 | 2896649 | 2955842201300778 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2013-02-07 | 2013-03-08 | No | Yes | Health Professional | 42030901 | Yes | No | 2013-02-07 | Yes | No | Reuse | 2013
42030901 | 2013-02-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the cable on the megasuturecut needledriver instrument was found to be broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The surface of the distal pulley also exhibited scratches. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT NAY | 3047345 | 3000117 | 2896649 | 2955842201300777 | 1 | NAY | Manufacturer | User facility | NOT APPLICABLE | Initial | 2013-02-07 | 2013-03-08 | No | Yes | Health Professional | 42030901 | Yes | No | 2013-02-07 | Yes | No | Initial | 2013
42000606 | 2013-01-11 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure a cable on the large needle driver instrument broke. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the reported broken cable. One grip cable was broken at the distal clevis and there was damage on idler pulleys. Around the edges of the pulley, there were nicks and dents creating a sharp edge which is likely the cause of the cable breaking. Evidence not conclusive, but damage is likely due to mishandling. The broken segments stuck out approximately. 2740 at the wrist. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3018422 | 2992493 | 2889029 | 2955842201300757 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-05 | 2013-03-07 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-15 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the fenestrated bipolar forceps instrument was not being recognized by the davinci robot system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to replicate the reported failure mode. The instrument was installed on the in-house davinci robot system and it was successfully recognized, showing 9 uses left. The instrument was driven and moved intuitively, the grips opened and closed properly. The pogo pins did not stick and were not contaminated. Instrument passed electrical continuity test. Additional observations not reported by site were bipolar pin damage and main tube damage. The bottom pin was found to be bent towards the top pin. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damages were likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3018543 | 2992582 | 2889118 | 2955842201300762 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-07 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
40018010 | 2013-01-30 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the mcs tip cover accessory installed on the monopolar curved scissors (mcs) instrument tore due to the blades on the mcs instrument were open. No missing or fallen pieces were reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found that the tip cover was returned with a crack in the tip cover wall which likely occurred as a result of installation of the tip cover onto the monopolar curved instrument while the blades on the instrument were open. The crack in tip cover was approximately. 4210 long. No other damage to the tip cover was found. The instruments and accessories user manual instructions for tip cover accessory installation indicates the following: tip cover accessory installation - before use: close the scissor blades. Straighten the wrist of the instrument. Grasping the tip cover accessory with the installation tool, slide the tip cover accessory onto the distal end of the instrument until it comes to the stop. A twisting motion can be used to ease installation. The distal end of the instrument should be facing away from you during installation. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3019449 | 2993280 | 2889814 | 2955842201300763 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-05 | 2013-03-07 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | Yes | Unkown | 2013
IS3000 A60P7 | 2013-02-05 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the monopolar curved scissor's instrument located on patient side manipulator (psm) arm 1 failed. The customer contacted isi technical support however due to issues unrelated to da vinci si surgical system, the surgeon decided to convert the procedure to traditional open surgical techniques to complete the planned procedure. | The isi clinical sales representative indicated that the surgeon experienced a problem due to a broken cable located on the back of the da vinci si surgical system patient side manipulator (psm) arm 1 resulting in the surgeon not able to advance the monopolar curved scissor's instrument into the patient. The surgical staff moved the monopolar curved scissor's instrument from psm arm 1 to psm arm 3 to continue with the procedure. As the surgeon continued with the procedure a hole was discovered to have been made in the patient's bladder and the surgeon made the decision to convert the case to traditional open surgery to repair the bladder. The case was completed with no further harm to the patient, the patient status was stable following the procedure. The isi field service engineer visited the account the following day to repair the patient side manipulator (psm) arm. As of the date of this report, no further related incidents have occurred. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3019540 | 2993400 | 2889936 | 2955842201300764 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-02-05 | 2013-03-07 | Yes | Yes | Health Professional | IS3000 A60P7 | Yes | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
40018010 | 2013-01-30 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the mcs tip cover accessory installed on the monopolar curved scissors (mcs) instrument tore due to the blades on the mcs instrument were open. No missing or fallen pieces were reported. | The tip cover accessory was returned and evaluated. Per the customer reported complaint, engineering found the tip cover was returned with a crack in its wall. Engineering concluded that the crack was likely due to installation of the tip cover onto the monopolar curved scissors instrument while the blades on the instrument were open. The crack on the tip cover was approximately. 1380 long. There was no other damage to the tip cover found. The instruments and accessories user manual instructions for tip cover accessory installation indicates the following: tip cover accessory installation - before use: close the scissor blades. Straighten the wrist of the instrument. Grasping the tip cover accessory with the installation tool, slide the tip cover accessory onto the distal end of the instrument until it comes to the stop. A twisting motion can be used to ease installation. The distal end of the instrument should be facing away from you during installation. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3019694 | 2993546 | 2889814 | 2955842201300765 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-05 | 2013-03-07 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | Yes | Unkown | 2013
41029902 | 2013-02-04 | Malfunction |  | Event Description: --/-- On (b)(6) 2013, intuitive surgical received the suction irrigator instrument in a damaged condition. No event details concerning the instrument were provided by the site. | Manufacturer Narrative: --/-- On (b)(6) 2013, isi contacted the initial report concerning the returned suction/irrigator instrument. The initial reporter indicated that at this time he does not recall the specific complaint details concerning the instrument as originally provided to him by the surgical staff; however, he indicated that there was no report by the surgical staff that any piece(s) from the affected instrument fell into a patient. The initial reporter also indicated that there was no report by the surgical staff that any patient harm, adverse outcome or injury occurred involving the returned instrument. --/-- Additional Manufacturer Narrative: --/-- Engineering evaluation found that the proximal end cap disk was dislodged from female-female disk. The disks and distal tip exhibited various scratch marks. The two drive cables were cut at the snake wrist and the entire wrist assembly was detached from the main tube. Engineering concluded that damage to the snake wrist disk required cutting to allow for removal of the instrument from the cannula and that dislodgment was likely due to overloading at the distal end of the instrument. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The failure analysis finding does not constitute a mdr reportable event; however, recurrence of the device malfunction could likely cause or contribute to an adverse event if the malfunction were to recur. | SUCTION IRRIGATOR INSTRUMENT | Reuse | 3019771 | 2993608 | 2890142 | 2955842201300766 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-03-06 | 2013-03-07 | No | Yes | Health Professional | 41029902 | Device Returned To Manufacturer | No | 2013-03-06 | Device Not Retu | No |  | 2013
42017206 | 2013-01-02 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the maryland bipolar forceps instrument would not energize. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as functional testing of the instrument on an is3000 in-house system found that the instrument was successfully recognized and the instrument passed cautery testing with smoke coming from a wet paper towel. The pogo pins do not stick, and there were no signs of damage to the dallas chip, bipolar pins, or conductor wire. The instrument also passed electrical continuity testing. Engineering evaluation also found that the distal end of the main tube exhibits scratches. From the proximal end to approximatly 3 in, there are several deep scratches, leaving a rough surface. Engineering concluded that the damage to the main tube is likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The failure analysis finding does not constitute a mdr reportable event; however; the damage to the instrument's main tube, could likely cause or contribute to an adverse event if the malfunction were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3019738 | 2993616 | 2890150 | 2955842201300767 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-05 | 2013-03-07 | No | Yes | Health Professional | 42017206 | Device Returned To Manufacturer | Yes | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42806503 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure, the single site port ripped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The port has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SINGLE SITE PORT | ENDOSCOPIC ACCESSORY | 3019802 | 2993659 | 2890193 | 2955842201300768 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-07 | No | Yes | Health Professional | 42806503 |  | No | 2013-02-06 | Device Not Retu | No | Unkown | 2013
42020505 | 2013-02-06 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted frayed cables on the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be frayed at the distal clevis hub. The frayed strands stick out at the wrist. Both pitch cables were also found to be loose at the distal end. The clevis did not exhibit any other damage. Additional observation not reported by site were derailed backend cable and main tube damage. The instrument's housing was removed to find one pitch cable derailed from the back idler pulleys. The cable was found to be wedged between two pulleys and it had lost tension. The clamping pulley screws were still tight. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were 0. 075 - 0. 200 in length and they were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3016889 | 2991165 | 2887699 | 2955842201300739 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-10 | Malfunction |  | It was reported that during a da vinci si pancreatectomy (whipple procedure) procedure the fenestrated bipolar instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was broken. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3017054 | 2991296 | 2887849 | 2955842201300743 | 1 | NAY | Manufacturer | DistributorHealth Professional | Other | Initial | 2013-02-05 | 2013-03-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-02-06 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the 8mm small grasping retractor instrument was noted to have exposed wires. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the of instrument with wires exposed was confirmed. Instrument was found with frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 04 in length. There was no other cable damage found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR 8MMINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3017047 | 2991338 | 2887872 | 2955842201300741 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-06 | 2013-03-06 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | Yes | 2013-02-06 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-04 | Malfunction |  | It was reported during a da vinci si prostatectomy procedure that the fenestrated bipolar instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was broken. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3017050 | 2991342 | 2887876 | 2955842201300742 | 1 | NAY | Manufacturer | DistributorHealth Professional | Other | Initial | 2013-02-05 | 2013-03-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-06 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted frayed cables on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with frayed pitch cable located at the distal end. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Frayed strands were approximately 0. 0625 in length. No other cable damaged was found. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3017052 | 2991344 | 2887878 | 2955842201300740 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-06 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-05 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the 8mm pk dissecting forceps instrument was not working. The scrub technician pulled the instrument out and noticed the wire was coming out of the area. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. Engineering also found scratches on the main tube at the distal end that showed light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage to the main tube was likely due to mishandling. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3017055 | 2991346 | 2887880 | 2955842201300744 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-05 | 2013-03-06 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
IS3000 A60P4 | 2013-02-08 | Other | Required Intervention | It was reported that prior to starting a da vinci si prostatectomy procedure, the surgical staff experienced a system error code 25521. Upon extending the camera arm, it created a system error code. The surgical staff was able to reboot the system and clear the fault code, however, upon moving the camera arm, the system immediately faulted again with the same error code. Isi technical support informed the customer that this is likely to continue to occur anytime the camera arm is moved. The surgeon made the decision to abort the surgery prior to placing ports and prior to docking the robot. No patient harm was reported. | The investigation conducted by field service engineering concluded that system error code 25521 was associated with the endoscopic camera manipulator (ecm). The ecm is the camera arm located on the patient side cart and provides the sterile interface for the 3d endoscope. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error code 25521 appears when the system detects that it has lost communication of sensor information between two printed circuit assemblies in the arm control system. In the event of these communication issues, the system records the fault and transitions to a safe state. The ecm was returned to intuitive surgical for failure analysis investigation. A visual inspection was performed on the ecm, and no problems were found. The arm was mounted and configured on a system and powered up in maintenance mode. All sensor checks and switch testing passed. The sine cycle, which runs sinusoidal motion trajectories to exercise the joints and the flat flexible cables (ffc's) going to and from the escc/espm (embedded serializer for the camera cannula/embedded serializer patient manipulator) was executed for about (b)(4) without any problems. Engineering was unable to recreate the reported error; however, it was confirmed in the system log. Engineering recommended the escc, ffc's, main harness be replaced and the espm be upgraded as a precaution. Axis 1 housing produced a loud grinding noise during the sine cycle while being exercised. Engineering was unable to determine the exact source of the noise. As a fix, engineering recommended upgrading the axis 1 motor and the axis 1 and axis 2 brakes. The clinical sales representative indicated the case was aborted post-anesthesia, no ports had been placed; the robot had not been docked. The patient did not return with any complications. The procedure was not completed. The surgery was elective and therefore, postponed. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3017057 | 2991348 | 2887882 | 2955842201300745 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-02-08 | 2013-03-06 | Yes | Yes | Health Professional | IS3000 A60P4 | Device Returned To Manufacturer | No | 2013-02-08 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-04 | Malfunction |  | It was reported during a da vinci s surgical procedure that the megasuturecut needle driver instrument cables were broken at the tip of the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the cables were broken at tip. The one grip cable was broken at the distal clevis and there was evidence of wear on the idler pulleys. Engineering concluded the damage was likely linked to cable derailing off pulley, causing the cable to rubbing against idler pulley and breaking. The broken segment stuck out approximately. 6370 at the wrist. Engineering also found frayed cables, corroded bearings and damaged blades. On the opposite side, a broken grip cable was frayed at the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at the wrist were not damaged. The frayed cable was possibly linked to the broken cable. The bearings lying under each clamping pulley, exhibited signs of corrosion. Each axis still rotated freely with no resistance. In addition, the blades on the grips were damaged. There were multiple indentations along both bottom and top blades. Engineering concluded that damage to the blades was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3017124 | 2991449 | 2887982 | 2955842201300747 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-06 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-01-22 | Malfunction |  | It was reported during a da vinci si hysterectomy procedure that the monopolar curved scissors instrument was noted to be dull. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint. Engineering was unable to perform the cut test due to a broken grip cable. Engineering found a broken grip cable. The one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at wrist were not damaged. The electrical continuity testing passed. Additionally, engineering found an indentation on the edge of the distal pulley. The last finding was print removed from extension tube. The part of the print was removed from the extension tube measured roughly about 0. 356 x 0. 188. Engineering concluded the damage was most likely due to improper cleaning. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3017129 | 2991454 | 2887987 | 2955842201300749 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the wire broke off the tip of the large suturecut needle driver instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken cable. The one grip cable was broken at the distal clevis and there was wear on the idler pulleys. Engineering concluded the damage was likely linked to the cable derailing off the pulley, causing the cable to rub against the idler pulley and break. The broken segment stuck out approximately. 4795 at the wrist but no other cables were damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLEDRIVER8MM INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3017159 | 2991455 | 2887988 | 2955842201300748 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-08 | 2013-03-06 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-02-08 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-28 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, while the customer was preparing to install the prograsp forceps instrument, they noted that the instrument's wheel wouldn't turn. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3017192 | 2991459 | 2887992 | 2955842201300746 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-06 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-12 | Malfunction |  | It was observed during central processing that the mega suturecut needle driver instrument had frayed wires. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of a frayed wire. The instrument was found to have a broken grip cable. One grip close cable was broken at the distal idlers. The idler pulley was able to spin freely and was not damaged. A cable segment was observed to be sticking out at the instrument's wrist. Other cables at the wrist were not damaged. Engineering evaluation also found an indentation on the edge of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3017205 | 2991482 | 2888016 | 2955842201300750 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-06 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
40018010 | 2013-01-15 | Malfunction |  | It was reported that during a da vinci s low anterior resection procedure, the surgeon noticed that the tip of the 8mm monopolar curved scissors (mcs) tip cover accessory used in conjunction with the mcs instrument had breakage. According to the reporter, the surgical staff exchanged the mcs tip cover accessory and the planned procedure was completed. There were no missing or fallen pieces reported. There was no patient harm, adverse outcome or injury reported. | The instrument accessory was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of mcs tip cover accessory breakage. The tip cover accessory was returned with various mechanical tears and a hole burned though the silicone. The silicone exhibited localized melting around the hole which is indicative of arcing. The hole was found to be oblong, measuring roughly. 035 x. 020, and was located roughly. 500 below the distal end of the tip cover accessory. The edge of the hole had a tear, and there were also various tears in the general vicinity of the hole. Axially oriented tears were observed at the distal end opening of the tip cover accessory. The instruments and accessories user manual specifically states: general precautions and warnings: inspect the tip cover accessory periodically during use. If any damage or tears are observed, replace the tip cover accessory with a new one. Do not use another instrument to clean the tip cover accessory. O to prevent damage to the tip cover accessory insulation, always straighten the instrument wrist under endoscopic visualization before removing the instrument through the cannula o do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3017326 | 2991627 | 2888160 | 2955842201300751 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-04 | 2013-03-06 | No | Yes | Health Professional | 40018010 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | Yes | Unkown | 2013
42009308 | 2013-01-24 | Malfunction |  | It was reported that during a davinci si prostatectomy procedure, the customer noted that the wires were off track on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with frayed pitch cable located at the distal idler with loose tension. Additional observation not reported by site was indentation damage on the pulley. The instrument exhibited indentations and scratch marks on the distal pulley surface. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3017405 | 2991696 | 2888230 | 2955842201300752 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-06 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
IS3000 A60P3 | 2010-04-26 | Injury | HospitalizationLife ThreateningRequired Intervention | On (b)(6) 2013, intuitive surgical received a legal summons indicating that post op a da vinci si lysis of adhesions with left ovarian endometrioma aspiration procedure performed on (b)(6) 2010 at (b)(6), it was discovered that the patient had sustained a bowel injury. Reportedly, post-operatively, the patient experienced symptoms compatible with a bowel injury and as a result the patient was hospitalized until (b)(6) 2010. The legal summons indicated that on (b)(6) 2010, the patient was admitted into the emergency room at (b)(6) with peritoneal signs. Reportedly, the patient had to be resuscitated and a ct scan performed found that the patient had developed massive pneumoperitoneum with a massive amount of free fluid. Reportedly, the patient underwent an exploratory laparotomy with resection and anastomosis of the distal ileum and a washout of the peritoneal contamination. Reportedly, the patient has continued to experience complications as a result of the bowel injury and infection. | Isi has contacted the hospital several times to obtain additional information concerning the reported event; however, no additional information has been provided. A follow-up medwatch report will be submitted to the fda if additional information concerning the reported event is provided. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3017446 | 2991714 | 2888247 | 2955842201300753 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-02-04 | 2013-03-06 | Yes | No | Health Professional | IS3000 A60P3 | No | No | 2013-02-04 | No | No | Reuse | 2013
42017207 | 2013-01-24 | Malfunction |  | It was reported that during reprocessing that the tips did not align on the maryland bipolar forceps instrument. There were no missing pieces or fallen pieces reported. | The instrument was returned and evaluated. per the customer reported complaint engineering evaluation found that the instruments tips did not align. One grip was bent, causing side to side misalignment of grips. There was a. 04 offset at the tips. The bent grip had separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. Evidence not conclusive, but damage likely due to mishandling. An electrical continuity test passed. Additional observation found a pitch cable frayed at the distal celvis hub. No damage was found at the clevis. The frayed segment is approximately 0. 02 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3017461 | 2991723 | 2855631 | 2955842201300754 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-04 | 2013-03-06 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-02-01 | Malfunction |  | The bowl grasper instrument was returned but no reason for the return was provided. There were no missing pieces or fallen pieces reported. | The instrument was returned and evaluated. No reason for return was provided so the complaint cannot be confirmed or denied. Engineering's observations not reported by site were tube damage and a broken housing. The black tube insulation was damaged at the distal end. The insulation was gouged on one side and had a couple of pieces missing roughly. 675 and 1 above the snake wrist. The housing had three snaptabs and two locating pins completely broken off. One side of the housing could be lifted up as a result of the damage. Evidence not conclusive, but the tube and housing damage are likely due to mishandling/misuse. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3017490 | 2991761 | 2888295 | 2955842201300756 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-04 | 2013-03-06 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-21 | Malfunction |  | It was reported that during a davinci si prostatectomy procedure, the customer noted that the wires were off track on the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional observation not reported by the site were indentations and scratch marks on the pulley. The instrument exhibited indentations located at the edge of the distal pulley and scratches on the surface. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3017465 | 2991768 | 2888301 | 2955842201300755 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-06 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-12 | Malfunction |  | It was observed during central processing that the mega suturecut needle driver instrument had frayed wires. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of frayed wires. The instrument was found to have a frayed pitch cable at the wrist. Frayed segments were observed in various lengths. The frayed cable was also found to be derailed. The grips could not open and close properly. Engineering evaluation noted that movement and functionality might not have been precise. In addition, engineering evaluation also found damage to the idler pulley rim. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3017557 | 2991848 | 2888016 | 2955842201300758 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-12 | 2013-03-06 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-12 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-02-05 | Malfunction |  | It was reported that during reprocessing of the fenestrated bipolar forceps instrument, the cable on instrument was noted to be frayed and separated. | The instrument was returned and evaluated. Per the customer reported complaint, with the following clarification that the pitch cable on the instrument was broken. The pitch cable was broken at the wrist and the cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3017633 | 2991898 | 2888432 | 2955842201300761 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-05 | 2013-03-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42014304 | 2012-11-02 | Malfunction |  | It was reported that during a da vinci si tongue base resection procedure, the cable on the maryland dissector instrument broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode of a broken cable; however, evaluation found that the main tube insulation had scratches and gouge marks, with materials removed at the distal end. Engineering concluded that damage to the main tube was likely caused by excess force contact on surface of the main tube. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | 5MM MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3017605 | 2991910 | 2888444 | 2955842201300759 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-02-05 | 2013-03-06 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, one of the cables on the pk dissecting forceps instrument frayed. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint with the following clarification that the pitch cable was broken. The pitch cable was broken at the wrist and the cable segment that contains the crimp was remained installed in clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3017614 | 2991912 | 2888446 | 2955842201300760 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-05 | 2013-03-06 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
420093 | 2013-02-22 | Injury | Other | During surgical procedure two robotic instruments broke. |  | INTUITIVE SURGICAL | PROGRASP FORCEP | 3027042 | 2999598 | 2896132 | MW5029293 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-03-06 | 2013-03-06 | Yes | Yes | Health Professional | 420093 | Yes | Yes |  |  |  |  | 2013
420093 | 2013-02-22 | Injury | Other | During surgical procedure two robotic instruments broke. |  | INTUITIVE SURGICAL | PROGRASP FORCEP | 3027042 | 2999598 | 2896132 | MW5029293 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-03-06 | 2013-03-06 | Yes | Yes | Health Professional | 420093 | Yes | Yes |  |  |  |  | 2013
| 2013-02-27 | Malfunction |  | Large fibroid removed from uterus - fibroid being held with davinci tenaculum. Monopolar curved scissors used to cut fibroid in half. Tenaculum and scissor never touched and kept a good distance apart while cutting. Having trouble getting scissor to work and burn noted to peritoneum. No damage to bowel or bladder - no other procedure needed. During time monopolar scissor appeared to be not working it apparently arced and energy transferred to tenaculum. Tenaculum near the joint of the instrument is what caused the burn on peritoneum. No one saw it arc - it was assumed by staff and doctor that it arced. Possibly from dissecting adhesions. No visible damage or imperfections noted. Insulation intact. What was the original intended procedure?davinci myomectomy. |  | DA VINCI ENDOWRIST TENACULAR FORCEPS | SYSTEMSURGICALCOMPUTER CONTROL | 3040628 | 3014770 | 2911284 | 3014770 | 1 | NAY | User Facility |  |  | Initial | 2013-03-06 | 2013-03-06 | No | Yes | Physician |  | No | No Answer Provided |  |  |  |  | 2013
42018404 | 2013-01-13 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff alleged that the permanent cautery spatula instrument had an exposed wire. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue of an exposed wire. Engineering evaluation noted that the wire was exposed because the yaw pulley notch in between the distal clevis was broken which probably caused the conductor cap and wire to pop out of place. The conductor cap wire slipped out during the engineering evaluation. Engineering evaluation also found corroded clamping pulleys. All of the clamping pulleys had signs of moderate corrosion which was adding friction on the cables. Engineering evaluation concluded that the damage was likely due to improper cleaning. Electrical continuity testing of the instrument passed. The cleaning and sterilization user manual specifically states: o when scrubbing the tip of cautery instruments, take care not to damage the insulation. O do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. O prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken yaw pulley notch issue if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULAINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3015368 | 2989899 | 2886432 | 2955842201300728 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-08 | 2013-03-05 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2013-02-08 | Device Not Retu | No | Initial | 2013
42009308 | 2013-01-29 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a broken wire at the prograsp forceps instrument's tip. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. 48 - additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3015465 | 2989982 | 2886516 | 2955842201300729 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-04 | 2013-03-05 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-01-29 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a broken wire at the tip of the small grasping retractor instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3015622 | 2990074 | 2886608 | 2955842201300730 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-04 | 2013-03-05 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-11 | Malfunction |  | It was reported that prior to starting a davinci s surgical procedure, the surgical staff allegedly noticed that the mega suturecut needle driver instrument had frayed cables. The alleged issue was found during setup. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue of a frayed wire. The instrument was found to have a broken grip cable. One grip close cable was broken at the distal idlers. The idler pulley was able to spin freely and was not damaged. A cable segment was observed to be sticking out at the instrument's wrist. Other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3015668 | 2990128 | 2886661 | 2955842201300731 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-11 | 2013-03-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-11 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-14 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer was unable to install the maryland bipolar forceps instrument in the davinci robot system, the black instrument's shaft was curved. The instrument was replaced to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was inspected for recognition using an in-house davinci robot system. The davinci robot system successfully recognized the instrument, showing 2 uses left. The pogo pins did not stick and were not contaminated. Instrument passed electrical continuity test. Instrument was driven and moved intuitively, the grips were able to open and close. The shaft was straight. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3015795 | 2990163 | 2886697 | 2955842201300732 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-04 | 2013-03-05 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-02-01 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly noticed that the jaws of the small clip applier instrument were bent. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue of bent jaws. The jaws of the instrument were found to be broken at the mid-point. The grip broken surface was not smooth and a sign of slight torsion was observed. Engineering evaluation concluded that the damage was likely due to mishandling during clip loading. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, isi contacted the initial reporter for follow-up information. The reporter denied that any fragments fell. She also denied that any patients were harmed or injured. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken tip issue if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3015788 | 2990195 | 2886729 | 2955842201300733 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-11 | 2013-03-05 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-02-11 | Device Not Retu | No | Initial | 2013
42030901 | 2013-01-21 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the wrist mechanism of the mega suture cut needle driver instrument was not articulating correctly. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on in-house davinci robot system. Recognition and engagement test passed. Instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The instrument's scissors cut cleanly through 0-silk suture. Instrument was found to be fully functional. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 3015890 | 2990246 | 2886780 | 2955842201300734 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-04 | 2013-03-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
40017208 | 2013-01-30 | Malfunction |  | It was reported that after completion of a da vinci prostatectomy procedure, the surgical staff reported seeing a loose cable with the maryland bipolar forceps instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue of a loose cable. The bipolar insert was observed to be missing from the back end. The pins and conductor wires were found to be sticking out of the chassis. Engineering evaluation noted that the insert might not have been fully seated in the chassis or might have been pulled out when the bipolar cord was removed. Engineering evaluation also noted that some bipolar cords are a tight fit on the bipolar pins and may require a high removal force, potentially causing the insert to pull out. In this case, the cord was also returned with the instrument and a test was done. The fitting was reportedly ok. Engineering evaluation concluded that it was possible that the insert was loosely fitted into the plug riser. Electrical continuity testing of the instrument passed. Engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering evaluation concluded that the scratches were likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issue if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3016010 | 2990324 | 2886857 | 2955842201300736 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-02-11 | 2013-03-05 | No | Yes | Health Professional | 40017208 | Device Returned To Manufacturer | No | 2013-02-11 | Device Not Retu | No | Reuse | 2013
42018406 | 2012-12-31 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted cracked insulation on the permanent cautery spatula instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal yaw pulley was missing a 0. 0845 piece from the ceramic insulation which is located close to grip tip. No signs of char on tip were observed. Instrument passed electrical continuity test. Engineering concluded that damage was likely due to mishandling. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3015994 | 2990328 | 2832689 | 2955842201300735 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-05 | 2013-03-05 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2013-02-05 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2013-02-01 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the nurse observed that the lens was missing from the endoscope. | The system endoscope projects light from the illuminator onto the surgical site and the video image of the surgical site captured by the endoscope is sent back through left and right channels to the camera head. The endoscope was returned to the original equipment manufacturer (oem) for evaluation. The oem confirmed that a lens is missing and found that the endoscope was returned with mechanical damage to the distal window assembly hermetic seal. Engineering concluded that damage to the hermetic seal resulted in fluid invasion causing subsequent major damage to the optical components. The da vinci s surgical system user's manual specifically states: pre-operative inspection - use care when handling the endoscope to prevent damage to the internal lens. Working with the endoscope at the patient side - the operator must ensure that the endoscope is handled with great care, as the instrument is very delicate and can be easily broken if dropped or struck. Instrument and accessories cleaning and sterilization - exercise care when cleaning and handling the distal end of the endoscope. Do not exert excessive force on the distal windows and never clean with sharp objects or instruments. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | DA VINC S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3016397 | 2990662 | 2887196 | 2955842201300737 | 1 | NAY | Manufacturer |  | Other | Initial | 2013-02-04 | 2013-03-05 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42018404 | 2012-12-10 | Malfunction |  | It was reported that during a da vinci si lobectomy procedure, the cable on the permanent cautery spatula instrument broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode of a broken cable; however, evaluation found that the distal clevis and pulley notch between the clevis was broken. The yaw pulley and conductor wire were dislodged and the yaw cables were found to be derailed off of the distal pulley. There was also a small crack on the proximal clevis at the distal end. Engineering concluded that the damage was likely due to excessive side loading or other type of mishandling. The broken pieces off the distal clevis and conductor wire cap were missing. Engineering evaluation also found that the main tube near the distal end had various scratches. Two of the longest scratches were. 6350 and. 2785 long. The banana plug was found bent at the back end of the housing. A piece of white string was found tangled around the pulley. Engineering evaluation concluded that the damage was likely due to mishandling. The instrument also passed electrical continuity testing. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3016518 | 2990784 | 2887317 | 2955842201300738 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-02-05 | 2013-03-05 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-05 | Malfunction |  | It was observed during central processing that the mega suturecut needle driver instrument had a broken cable. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue that the instrument had a broken cable. The instrument was found to have a broken grip cable at the base of the needle driver grip jaw. A strand of wire was sticking out between the jaws. A cable was also found to be loose at the clamping pulley. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012127 | 2986483 | 2882997 | 2955842201300681 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-05 | 2013-03-04 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-09 | Malfunction |  | It was reported during a da vinci si pulmonary wedge resection procedure that the maryland bipolar forceps instrument would not articulate the surgeon's movements. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated engineering was unable to confirm the reported failure mode. The instrument was placed on an in-house system and driven. Recognition and engagement testing passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The instrument is fully functional. The electrical continuity testing passed. Engineering also found various scratch marks exhibiting light material removal on the distal end of the main tube. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage to the main tube was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012328 | 2986700 | 2883214 | 2955842201300685 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-02-01 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the monopolar curved scissors had a chip on the blade that was identified during setup. No missing or fallen pieces were reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found that the instrument had a chip on the blades. There were circular shaped dents and damage found on blade edges. The dents and damage acted as hard stop and prevented the scissor blades from closing completely. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012307 | 2986703 | 2883217 | 2955842201300695 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42018304 | 2012-12-07 | Malfunction |  | It was reported that prior to starting a davinci s surgical procedure, plugs were noted to be loose on the permanent cautery hook instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. During failure analysis engineering no damage to the banana plug was observed. Instrument passed electrical continuity test. Additional observations not reported by site were broken cable, missing conductor cap, cracked proximal clevis, and corroded clamping pulleys. Engineering observed that one yaw cable was found to be broken near the proximal clevis opening. The broken cable segment sticks out at wrist. Other wrist cables were not damaged. The conductor cap was missing from distal end. The cap likely detached after the cable broke, as breakage allows the yaw pulley to slide out from distal clevis. Proximal clevis had a crack in the axial direction. The crack was located halfway between clevis ears. The location and orientation of the crack suggests overloading at the tip with the wrist pitched. The housing was removed to find clamping pulleys exhibiting pitting corrosion. Engineering concluded that damage was likely due to improper cleaning. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012319 | 2986708 | 2883222 | 2955842201300683 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-25 | Malfunction |  | It was reported that prior to starting a da vinci s hysterectomy procedure the pk dissecting forceps instrument was noted to have an exposed wire at the wrist. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the instrument had an exposed wire. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012327 | 2986719 | 2860561 | 2955842201300686 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-10 | Malfunction |  | It was reported that during a davinci si surgical procedure, the customer noted that the tip on the prograsp forceps instrument would not close. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. Instrument was placed on the in-house davinci robot system and driven. Instrument passed recognition and engagement test. Instrument moved intuitively with full range of motion in all directions. The grips were able to open and close properly. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012329 | 2986720 | 2883234 | 2955842201300684 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
ISI2000 A51P8 | 2013-01-25 | Other | Required Intervention | It was reported that prior to starting a procedure using the da vinci s surgical system, while setting up the illuminator, the site experienced a system error stating no light from illuminator. The nurse exchanged the illuminator light bulb however, the system error recurred and the site was unable to locate a replacement illuminator. The patient was under anesthesia when the surgeon made the decision to abort the planned surgical procedure. No further information is available. No patient harm was reported. | The field service engineer (fse) has replaced the illuminator to resolve repeated errors during startup. The system tested to specification. The illuminator was returned and evaluated. Engineering was able to replicate the customer reported issue when installing the illuminator onto a test system; it powered up but failed to turn on light. The illuminator has been sent to the oem (original equipment manufacturer) for repair. Isi has contacted the site to verify additional information concerning the status of the patient; however, no additional information has been provided at the time of this report. A follow up mdr report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3012334 | 2986732 | 2883246 | 2955842201300692 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-25 | 2013-03-04 | Yes | Yes | Health Professional | ISI2000 A51P8 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-28 | Malfunction |  | It was reported prior to starting a da vinci si procedure that the prograsp forceps instrument was noted to have a broken cable, the cable was in a loop. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the instrument had a broken cable. The one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found there were scratches on the surface of the distal pulley. Additionally, there were various scratch marks on the distal end of the main tube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012355 | 2986741 | 2883255 | 2955842201300688 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-30 | Malfunction |  | It was reported during a da vinci si hysterectomy procedure that the fenestrated bipolar forceps instrument cable broke during the procedure. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was broken. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012351 | 2986747 | 2883261 | 2955842201300696 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-30 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted that the cable was 'cable hanging out at end on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the one grip cable was found to be broken and it was sticking out at the wrist. The idler pulley was able to spins freely and it exhibited pulley face and rim damage. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012354 | 2986749 | 2883263 | 2955842201300694 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-01-30 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was noted that orange piece between scissor tip and end of instrument was splitting on the monopolar curved scissors instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that extension tube was found to be crack at the distal end. The crack was. 0250 wide. No other damage around instrument tip. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012364 | 2986777 | 2883291 | 2955842201300700 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-28 | Malfunction |  | It was reported that prior to starting a da vinci si procedure the pk dissecting forceps instrument was noted to have a broken cable. The cable was said to be in a loop. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found the instrument had a broken cable. The pitch cable was broken at distal clevis. There was no sign of damage around the area of cable breakage. The pitch cable was broken at the distal clevis hub and the cable did not stick out. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering also found several deep scratch marks at the distal end of the main tube exhibiting light material removal and a rough surface finish. Engineering concluded that damage was likely due to mishandling. Electrical continuity testing passed. There was no other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012365 | 2986780 | 2883294 | 2955842201300689 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure the pk dissecting forceps instrument was noted to have an exposed wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken pitch cable. The pitch cable was broken at the wrist. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012379 | 2986783 | 2883297 | 2955842201300701 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-02-01 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the thoracic grasper instrument was noted to have insulation is peeling off of the main tube. No missing or fallen pieces were reported. No missing or fallen pieces were reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the insulation was peeling off of the instrument. The main tube of instrument exhibited insulation coating peeled off near the distal end. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER5MM INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012380 | 2986784 | 2883298 | 2955842201300699 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-08-07 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the surgeon believed the cable inside the large needle driver instrument was broken. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012390 | 2986795 | 2883309 | 2955842201300693 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-08-01 | 2013-03-04 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-08-01 | No | No | Reuse | 2013
42029602 | 2013-01-25 | Malfunction |  | It was reported that during a davinci si sacrocolpopexy w/hysterectomy procedure, the customer noted a broken cable on the large suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one cable was found to be broken around grove near grip. The broken cable segment sticks out at wrist but there were no signs of damage where cable broke off. The idler pulley was able to spin freely and it was not damaged. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012388 | 2986807 | 2883321 | 2955842201300691 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-01-22 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure a broken wire was found on the small clip applier instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken wire. The instrument was found with a grip cable broken at the clamping pulley. Additional observation found clamping corroded and rusted. Evidence not conclusive, but damage likely due to mishandling during cleaning. Instrument. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. O prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3012462 | 2986881 | 2883395 | 2955842201300703 | 1 | NAY | Manufacturer | ForeignUser facility | Nurse | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | Yes | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-28 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure the wire on the 8mm prograsp forceps instrument was noted to be broken. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported complaint of broken wire. Visual inspection showed no damage to the drive cables at the distal end. The grips were able to open/close fine when input discs were manually rotated. Engineering found various scratch marks on the main tube with light material removal at the distal end. The scratches were. 085 -. 225 in length and not aligned with the tube axis. Engineering concluded that the main tube damage was likely due to mishandling. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012476 | 2986883 | 2883397 | 2955842201300690 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-28 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure that the tips of the maryland bipolar forceps instrument did not meet. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the reported complaint. One grip was bent, causing side to side misalignment of the grips. There was a. 056 offset at the tips, indicating overloading at the instrument's tip. An electrical continuity test passed. Additional observation not reported by site were scratches on the maintube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded the damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported main tube scratch issue were to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012463 | 2986884 | 2883398 | 2955842201300706 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-03 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure a wire broke on the wrist of the large needle driver instrument. The instrument was replaced with a new one. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a wire broke on the wrist of this instrument. One grip cable was broken at the distal clevis and there was wearing on the idler pulleys. Engineering evaluation concluded that the damage on the idler pulley was likely due to wear. The broken segment stuck out approximately. 3675 at the wrist. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3012477 | 2986885 | 2883399 | 2955842201300707 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-01-25 | Malfunction |  | It was reported that during a davinci si sacrocolpopexy w/hysterectomy procedure, the customer noted a broken cable on the large suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one cable was found to be broken around the grove near the grip. The broken cable segment sticks approximately 0. 50 out at wrist but there were no signs of damage where cable broke off. The idler pulley was able to spin freely and it was not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012471 | 2986888 | 2883321 | 2955842201300687 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci si procedure there were frayed wires near the tip of the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a frayed wire on the instrument. The instrument was found with one frayed cable along proximal and distal area. The cable derailed off the pulley and was rubbing against idler pulley, causing wear and some visible damage on idler pulley. Evidence not conclusive, but cable probably frayed from the cable friction against idler pulley, during use. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3012485 | 2986902 | 2883416 | 2955842201300709 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-30 | Malfunction |  | It was reported that prior to a da vinci s procedure a wire on the prograsp forceps instrument was sticking out near the tip. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a wire was sticking out near the tip of the instrument. The instrument was found with a frayed pitch cable at distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 04 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3012481 | 2986907 | 2883931 | 2955842201300710 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42014304 | 2013-01-30 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the sheath was damaged and the insulation was tearing on the maryland dissector instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that main tube exhibited severe discoloration leaving the tube with a greyish coating all around. Engineering concluded that damage was likely due to improper cleaning. Engineering also observed that the insulation was damage with small tears along midpoint of the main tube. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The instruments and accessories user manual specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | MARYLAND DISSECTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012486 | 2986911 | 2883425 | 2955842201300697 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-30 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was noted that one of the wires that operates the jaws was broken on the fenestrated bipolar instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed one pitch cable was found to be broken. The pitch down cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist are not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012468 | 2986927 | 2883441 | 2955842201300698 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-08 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy a wire on the pk dissecting forceps broke. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the broken wire. Engineering found a broken conductor wire segment at the distal end. The broken wire segment stuck out approximately. 6215 at the wrist but there were no signs of damage where the cable broke off. No other cables were damaged. Additional observations not reported by site were a deep scratch and cracked housing. Around the distal end of the main tube, there was one scratch mark exhibiting light material removal and a rough surface finish. Evidence not conclusive, but damage may be due to mishandling. Also, there was a small crack on the housing where the main tube enters housing. Evidence is not conclusive, but damage is likely due to mishandling. An electrical continuity test failed. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3012513 | 2986946 | 2883460 | 2955842201300705 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42011704 | 2013-01-09 | Malfunction |  | Event Description: --/-- It was reported that during a davinci si laryngoscopy procedure, while using the needle driver instrument to hold the omniguide laser introducer, one of the needle driver jaws broke off and fell into the patient, the clinical staff was able to retrieve the piece, and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- Isi contacted initial reporter biomed, he indicated that the needle driver instrument failed after of approximately 1 hour of use. The procedure was completed without any harm to the patient. Patient was released from the hospital as planned. Biomed provided some pictures of the needle driver instrument. Based on the pictures provided by the biomed, engineering observed that one grip appeared to be broken near the junction between the grip arm and the hub. The junction between the arm and the hub seem to be a sharp corner, which could become an area of high stress concentration. The rest of the pictures were too blurry to make any other observations. As of the date of this report, the instrument has not been returned for evaluation. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be broken off at the base, the broken piece was not returned. The grips were found to be fractured where the arm meets with the hub. The other grip was not damaged. | NEEDLE DRIVER INSTRUMENT | Reuse | 3012552 | 2986984 | 2883498 | 2955842201300712 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-05-02 | 2013-03-04 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | No | 2013-05-02 | Device Not Retu | No |  | 2013
40018012 | 2013-02-01 | Malfunction |  | It was reported during a da vinci si surgical procedure that the 8mm monopolar curved scissors (mcs) tip cover accessory used in conjunction with the mcs instrument, melted. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3012591 | 2986994 | 2883508 | 2955842201300702 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 40018012 | No | No | 2013-02-01 | No | Yes | Unkown | 2013
42030901 | 2013-01-25 | Malfunction |  | It was reported that during a davinci s hysterectomy procedure, the customer noted that wires were sticking out of the mega suture cut needle driver instrument's jaw. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip cable was found to be broken at the distal idlers. The cable segment was sticking out at the wrist. Additional observation not reported by the site was physical damage to the idler pulley. The idler pulley exhibited some indentations. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3012556 | 2987009 | 2883523 | 2955842201300713 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-31 | 2013-03-04 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure a wire broke in the endowrist on the pk dissecting forceps instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the reported broken wire. Engineering evaluation found a broken pitch cable at the wrist. The cable segment that contains the crimp was still installed on the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENET | ENDOSCOPIC ELECTROSURGICAL INS | 3012565 | 2987021 | 2883535 | 2955842201300704 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-01 | 2013-03-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci s procedure a wire was sticking out of the large needle drive instrument. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a wire was sticking out on the instrument. One grip cable was broken at the distal clevis. The broken segment stuck out approximately. 3755 at the wrist. The idler pulley spun freely but also exhibited some wear. The cable likely derailed off the pulley and was rubbing against the idler pulley causing wear and some visible damage. The other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3012577 | 2987023 | 2883537 | 2955842201300711 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-30 | Malfunction |  | It was reported that prior to starting a procedure on a da vinci s system there was a wire sticking out near the tip of the prograsp forceps instrument. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3013030 | 2987397 | 2883931 | 2955842201300708 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-03-04 | No | Yes | Health Professional | 42009308 | No | No | 2013-01-30 | No | No | Unkown | 2013
42019408 | 2013-02-05 | Malfunction |  | It was observed during central processing that the mega needle driver instrument had a broken cable. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results did not confirm the alleged issue of a broken cable. Visual inspection of the instrument showed no damage to the drive cables at the distal end. The grips were able open/close fine when the input discs were manually rotated. Engineering evaluation also found various scratch marks on the distal end of the instrument's main tube with light material removal. The scratches are. 100 -. 145 in length and not aligned with the tube axis. Engineering evaluation concluded that the scratches were likely due to mishandling/misuse. Zero uses left. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, isi contacted the initial reporter of this complaint and confirmed that no patients were harmed because of the alleged issue. The initial reporter denied that there was any patient involvement with the instrument in relation to the alleged issue. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issue if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3013685 | 2988049 | 2884563 | 2955842201300715 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-05 | 2013-03-04 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-01 | Malfunction |  | It was observed during central processing that the pk dissecting forceps instrument had a broken cable. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue of a broken cable. The instrument was found to have a frayed pitch cable at the wrist. Frayed segments were observed in various lengths. No damage was found at the clevis. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3013837 | 2988226 | 2884740 | 2955842201300716 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-05 | 2013-03-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-02-04 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted frayed wires on the thoracic grasper instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the sneak wrist. The cable was found to be broken at the wrist disc at the distal end. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3014043 | 2988350 | 2884864 | 2955842201300717 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-04 | 2013-03-04 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-02-04 | Malfunction |  | It was observed during central processing that the pk dissecting forceps instrument had a frayed cable. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue of a frayed cable. The instrument was found to have a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3014091 | 2988398 | 2884912 | 2955842201300718 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-05 | 2013-03-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-31 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a broken wire at the tip of the fenestrated bipolar instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Instrument passed electrical continuity test. Additional observation not reported by the site were main tube damage and bent bipolar pins. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were short in length and were not aligned with the tube axis. Both bipolar pins were also found to be bent. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3014117 | 2988446 | 2884960 | 2955842201300719 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-04 | 2013-03-04 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42018404 | 2013-01-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff alleged that the permanent cautery spatula instrument had a broken cable. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue of a broken cable. The pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULAINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3014206 | 2988512 | 2885026 | 2955842201300721 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-05 | 2013-03-04 | No | Yes | Health Professional | 42018404 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-31 | Malfunction |  | It was reported that during a davinci s surgical procedure, the customer noted frayed cables on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely but exhibited damages on the edge and on the surface. The cable segment sticks out at the wrist. Other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 3014213 | 2988515 | 2885029 | 2955842201300720 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-04 | 2013-03-04 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-01-31 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff alleged that the permanent cautery spatula instrument stopped moving. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue that the instrument stopped moving. Customer initially reported instrument stopped moving. The instrument was found to have non-intuitive motion due to a broken pitch cable at the wrist. The clevis did not exhibit any damage or wear marks. The cable crimp remained in the clevis. Engineering evaluation also found a derailed yaw cable. As a result, the movement and functionality of the spatula might not have been precise and smooth. The idler pulley also exhibited damage to the pulley rim. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULAINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3014235 | 2988522 | 2885036 | 2955842201300722 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-03-04 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-28 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure that the prograsp forceps instrument upon insertion in the robot was noted to have a broken wire prior to being used. The instrument was removed and replaced. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken wire. The instrument was found with a frayed cable at the distal idler pulley. The frayed cable section was approximately 0. 21 in length. No damage was found on pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3014262 | 2988551 | 2875490 | 2955842201300724 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-04 | 2013-03-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-07 | Malfunction |  | It was reported that during a da vinci si nissen fundoplication procedure, the reporter indicated that the mega suturecut needle driver instrument was acting in counterintuitive fashion. After removing the instrument and reseating the face plate, the instrument was reinserted. At that point, the reporter claimed that a wire broke when the surgeon attempted to take control of the instrument. The instrument was replaced by the surgical staff and the planned surgical procedure was completed with no patient harm, adverse outcome or injury. There were no missing or fallen pieces reported. | The tip cover has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3014249 | 2988553 | 2885067 | 2955842201300725 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-02-07 | 2013-03-04 | No | Yes | Health Professional | 42030901 | No | No | 2013-02-07 | No | No | Unkown | 2013
42034301 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci si pulmonary wedge resection procedure, the surgical staff alleged that the tube of the thoracic grasper instrument was splintering. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue that the instrument's tube was splintering. The instrument's main tube had two damaged sections with tube insulation removed. No pieces were observed to be missing. A piece of tape was found to be attached to the main tube but it was not concealing any damage. The tube insulation appeared to have scratches and gouge marks. No other instrument damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issue if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3014279 | 2988559 | 2885073 | 2955842201300727 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-07 | 2013-03-04 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-02-07 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the surgeon noted while performing (b)(6) stitch that the large needle driver instrument had a broken cable at the wrist. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a broken cable. The one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damages on the surface and the edge. Engineering concluded the edge damage was likely due to excess force contact with the cannula. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3014246 | 2988579 | 2885094 | 2955842201300723 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-02-05 | 2013-03-04 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-02-05 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci si pulmonary wedge resection procedure, the surgical staff alleged that the tube of the thoracic grasper instrument was splintering. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue that the instrument's tube was splintering. The main tube insulation appeared to have scratches and gouge marks, with material removed at the distal end. Engineering evaluation concluded that the damage was likely caused by excess force contact on the main tube surface. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issue if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3014275 | 2988586 | 2885073 | 2955842201300726 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-07 | 2013-03-04 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-02-07 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2013-02-01 | Malfunction |  | It was reported that one glass was falling off of the endoscope. There was no allegation of any harm, injury, or adverse outcome to a pt. | The system endoscope projects light from the illuminator onto the surgical site and the video image of the surgical site captured by the endoscope is sent back through left and right channels to the camera head. The endoscope was returned to the original equipment mfr (oem) for eval. Per the oem, the endoscope was rec'd with mechanical damage to the distal window assembly hermetic seal resulting in fluid invasion and subsequent minor damage to optical components. As of (b)(6) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3080913 | 3037789 | 2934303 | 2955842201300682 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-01 | 2013-03-01 | No | Yes | Health Professional | IS2000 A51P8 | Yes | Yes | 2013-02-07 | Yes | No | Reuse | 2013
| 2013-01-25 | Malfunction |  | On (b)(4) 2013, a (b)(6) video was posted online by dr (b)(6) with the following information: video - documented prograsp instrument malfunction. Notice the instrument's wire is torn cut. Video depicts cut wire, around 5 o'clock location of the video monitor at the end of the video clip. It shows around the middle of the prograsp instrument's end, sticks out in a right angle. Prograsp is normally changed after 10 case uses. This particular prograsp was used only once before. No injury happened in this case, because the cut was noticed and instrument disposed/replaced. If not noticed, it could have the potential of causing serosal bowel injury, due to the sharp edge of the cut metal wire. Perhaps, the manufacturer better add a sleeve around the bare wires for protection, similar to the protecting sleeve around the monopolar hot shear. As always, vigilance is needed in any case, any instrument, any surgery, anytime. | Intuitive surgical contacted the initial reporter regarding this complaint to request more information regarding the event; however, as of the date of this report no response has been received. The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3007947 | 2982625 | 2879139 | 2955842201300660 | 1 | NAY | Manufacturer | Health Professional | Other | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional |  | No | No | 2013-01-29 | No | No | Reuse | 2013
42022703 | 2013-02-04 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff allegedly noticed that the pk dissecting forceps instrument had a frayed wire. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint that the instrument had a frayed cable. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3008318 | 2982947 | 2879461 | 2955842201300661 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-04 | 2013-02-28 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff alleged that the fenestrated bipolar forceps instrument had a wire sticking out of the jaws. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint that the instrument had a wire sticking out of the jaws. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was missing in clevis. Electrical continuity testing of the instrument passed. Engineering evaluation also found indentations and scratches on the distal pulley. In addition, engineering evaluation found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering evaluation concluded that the scratches might have been due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable and main tube scratch issues if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3008387 | 2982999 | 2879513 | 2955842201300662 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-02-04 | 2013-02-28 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
400180 | 2013-01-31 | Injury | Other | On (b)(6) 2013, a video was posted on (b)(6) by dr (b)(6) with the following description: da vinci robotic hysterectomy. Video-documented insulation failure of monopolar hot shear instrument. Instrument replaced. No injury to the patient. Caution should be taken, when using monopolar electrocautery near ureter, small bowel, large bowel or any other organ. | The instrument and tip cover accessory have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Intuitive surgical attempted to contact the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided or obtained as of the date of this report. | 8 MM TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3008459 | 2983082 | 2879596 | 2955842201300663 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-29 | 2013-02-28 | Yes | Yes | Health Professional | 400180 | No | No | 2013-01-29 | No | Yes | Unkown | 2013
400180 | 2013-01-31 | Injury | Other | On (b)(6) 2012, a video was posted on (b)(6) by dr (b)(6) with the following description: video documented insulation failure of robotic monopolar hot shear instrument, during robotic assisted hysterectomy procedure. Insulation failure seen as sparks, firing against the posterior wall serosa of the uterine corpus. This is related to malfunction of the monopolar instrument (hot shear) sleeve. Careful handling of the instrument (and it's shortcomings) is a must, to avoid injury to surrounding structures in the operative field (such as bowel). Frequent change, as needed, of the hot shear monopolar instrument is need to avoid such failure. | The instrument and tip cover accessory have not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. Intuitive surgical attempted to contact the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided or obtained as of the date of this report. | 8 MM TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3008623 | 2983145 | 2879596 | 2955842201300664 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-29 | 2013-02-28 | Yes | Yes | Health Professional | 400180 | No | No | 2013-01-29 | No | Yes | Unkown | 2013
42027804 | 2013-01-25 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a broken cable on the grasping retractor instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3008834 | 2983355 | 2879869 | 2955842201300667 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2013-01-29 | Injury | Required Intervention | It was reported that after of a successful davinci si hysterectomy procedure, while the surgical staff was removing the cannula from the patient, it was noted that the patient sustained burn marks on the surface of the skin associated with patient side manipulator (psm) arm 2. | Intuitive surgical contacted davinci robotic coordinator, she indicated that the patient condition was stable; the patient didn't suffer any further complications. According to the information provided by her, the ground pads were correctly positioned on the patient. The da vinci robotic coordinator also indicated that a similar event occurred on (b)(6) 2013, but isi was not notified at the time of the event. Mdr 2955842-2013-00665 was created to address that event. The investigation conducted by the field service engineer (fse) could not find any davinci systems malfunctions. Fse performed an electrical safety test and system verification test. Fse verified that electrical safety passed on the vision site cart, surgeon console cart, and patient site cart. He also verified that all 3 of the cannula mounts were electrically isolated from the system through the resistance checks. The davinci robotic coordinator also indicated that the maryland bipolar forceps instrument was installed on patient side manipulator (psm) arm 2 and the monopolar curved scissors instrument was installed on patient side manipulator (psm) arm 1. All energy levels on the electro surgical unit (esu) were confirmed to be set at 45 and the procedure was side docked. The following morning, the staff had the valleylab generator (force triad) replaced and tested. Biomed reported that valleylab generator passed electrical safety tests and power test. The maryland bipolar forceps and the monopolar curved scissors instruments were returned and evaluated by isi; both instruments were found to be within specifications. Investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of (b)(6) 2013, there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3009043 | 2983512 | 2880026 | 2955842201300666 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-29 | 2013-02-28 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2013-01-25 | Injury | Required Intervention | It was reported that after of a successful davinci si hysterectomy procedure, while the surgical staff was removing the cannula from the patient, it was noted that the patient sustained burn marks on the surface of the skin. This mdr was created to address the event occurred on (b)(6) 2013; reported by davinci coordinator on related mdr 2955842-2013-00666. | Intuitive surgical contacted initial reporter of this event, she indicated that the patient condition was stable; the patient didn't suffer any further complications. According to the information provided by the surgical staff, the ground pads were correctly positioned on the patient. The investigation conducted by the field service engineer (fse) could not find any davinci systems malfunctions. Fse performed an electrical safety test and system verification test. Fse verified that electrical safety passed on the vision site cart, surgeon console cart, and patient site cart. He also verified that all 3 of the cannula mounts were electrically isolated from the system through the resistance checks. The davinci robotic coordinator indicated that all energy levels on the electro surgical unit (esu) were confirmed to be set at 45 and the procedure was side docked. The valleylab generator (force triad) used on this procedure was tested by the customer. Biomed reported that valleylab generator passed electrical safety tests and power test. The fenestrated bipolar forceps instrument used on this event was returned and evaluated by isi, the instrument was found to be within specifications. Investigation has shown that it is possible for the electrosurgical unit (esu) currents to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. However, the root cause of the customer reported failure mode cannot be definitively determined. A follow-up mdr will be submitted if additional information is received. As of (b)(4) 2013, there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3009045 | 2983513 | 2880027 | 2955842201300665 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-29 | 2013-02-28 | Yes | Yes | Health Professional | IS3000 A60P6 | Yes | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-17 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the mega suture cut needle driver instrument was not recognized by the davinci system robot. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. Instrument was checked for recognition using an in-house davinci robot system and it was successfully recognized. Additional observation not reported by the site was a broken grip cable located at the distal idlers pulley. The idler pulley was able to spin freely and it was slight damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3009244 | 2983646 | 2847593 | 2955842201300668 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-17 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the tips of the maryland bipolar forceps instrument were misaligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent, causing side to side misalignment of grips. There was a. 820 offset at the tips, indicating overloading at the tip. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3009226 | 2983655 | 2880169 | 2955842201300669 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci s procedure a wire on the tenaculum forceps instrument broke. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer reported complaint of a broken wire was confirmed. The pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3009259 | 2983676 | 2880341 | 2955842201300670 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-17 | Malfunction |  | It was reported that during a davinci s surgical procedure, the customer reported that the fenestrated bipolar instrument was unable to rotate. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a broken pitch cable at the wrist. The clevis did not exhibit any damage or wear marks. The cable crimp remained in the clevis. Engineering concluded that the non-intuitive motion was probably caused by the broken pitch cable. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3009314 | 2983728 | 2880242 | 2955842201300671 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-01-24 | Malfunction |  | Event Description: --/-- It was reported that during central processing, wires were found exposed at the tip of the mega needle driver. There were no missing or fallen pieces reported. No patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. The customer initially reported wires exposed at tips; however, for clarification, the damaged instrument component was a broken grip cable. Based on this clarification, the customers reported complaint is confirmed. The grip cable was broken and a cable segment stuck out at the wrist. Additionally, the idler pulley spun freely but exhibited pully face and rim damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. The customer initially reported wires exposed at tips; however, for clarification, the damaged instrument component was a broken grip cable. Based on this clarification, the customers reported complaint is confirmed. The grip cable was broken and a cable segment stuck out at the wrist. Additionally, the idler pulley spun freely but exhibited pully face and rim damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER | Reuse | 3009369 | 2983782 | 2880296 | 2955842201300672 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-24 | 2013-02-28 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No |  | 2013
42017207 | 2012-11-16 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff alleged that the maryland bipolar forceps instrument was only working from time to time. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results could not confirm the alleged complaint that the instrument would only work from time to time. The instrument was checked on the dallas chip programmer (dcp). The dcp successfully recognized instrument. Also, the instrument was placed on a system and driven. Recognition and engagement of the instrument passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The pogo pins did not stick and were not contaminated. Engineering evaluation also found deep scratches on the instrument's main tube, with potential fragment missing. Both scratches were approximately. 25. Engineering evaluation concluded that the scratches were likely due to mishandling. Electrical continuity testing of the instrument passed. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issues if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3009371 | 2983784 | 2880298 | 2955842201300673 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-02-04 | 2013-02-28 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci s procedure a wire on the tenaculum forceps instrument broke. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer reported complaint of a broken wire was confirmed. The pitch down cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The other cables at wrist were not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3009413 | 2983827 | 2880341 | 2955842201300674 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | Yes | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci s procedure a wire on the fenestrated bipolar forceps instrument broke. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The complaint of broken wire was confirmed. The pitch up cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed on the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 080 -. 185 in length and not aligned with the tube axis. Evidence not conclusive, but damage may have been caused by mishandling/misuse. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3009494 | 2983912 | 2868952 | 2955842201300675 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci s procedure a wire broke on the mega suturecut needle driver instrument. There were no missing pieces or fallen pieces reported. There were no missing pieces or fallen pieces reported. | The instrument was returned and evaluated. The reported complaint of a broken wire was confirmed. One grip close cable was broken. The idler pulley spun freely and did not exhibit damage. The broken cable segment stuck out at the wrist. Other cables at wrist are not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3009540 | 2983938 | 2880452 | 2955842201300676 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-09-07 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the surgical staff alleged that the prograsp forceps instrument would not open anymore. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results could not confirm the alleged complaint that the instrument would not open. The instrument was placed on a system and driven. Recognition and engagement of the instrument passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The instrument was found to be fully functional. Engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering evaluation concluded that the scratches were likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issue if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3009602 | 2984000 | 2880514 | 2955842201300677 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-02-04 | 2013-02-28 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci s procedure a wire on the tenaculum forceps instrument broke. There were no missing pieces or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed a broken cable. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Additional findings not reported by the customer were an indentation on the edge of the distal pulley. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering evaluation concluded that indentation and scratch mark damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT ACCESSOR | 3009624 | 2984047 | 2880341 | 2955842201300678 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-12-05 | Malfunction |  | It was reported that after completion of a da vinci si prostatectomy procedure, the surgical staff reported seeing a frayed cable with the monopolar curved scissors instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue that the instrument had a frayed wire. The instrument was found to have a broken grip cable at the distal idlers. The idler pulley was observed to spin freely and was not damaged. The cable segment was sticking out at the wrist. Other cables at the wrist were not damaged. Electrical continuity testing of the instrument passed. Engineering evaluation also found an indentation on the edge of the distal pulley. In addition, engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering evaluation concluded that the damages were likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issue if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3009753 | 2984168 | 2880682 | 2955842201300679 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-02-04 | 2013-02-28 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-01-29 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure that the monopolar cautery instrument was noted to have a tear in the shaft of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the nonconformance was damage to the main tube insulation at the distal end. The insulation was found with a gouge mark with material missing. The mark is not axially aligned with the tube. The main tube also exhibited a couple of different damaged sections with tube insulation with deep scratches but no material removed. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3009868 | 2984304 | 2880818 | 2955842201300680 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-29 | 2013-02-28 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
400156 | 2013-01-28 | Malfunction |  | It was reported that during a davinci si colectomy procedure, the monopolar cautery hook tip was loose and it fell into the patient. The surgical staff was able to retrieve the monopolar cautery hook tip from the patient and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | Isi contacted initial reported clinical sales rep, he indicated that based on the information provided by the surgical staff, the monopolar cautery hook tip was not screw correctly into the monopolar cautery instrument. Clinical sales rep visited the account to review the monopolar cautery hook tip installation process; the surgical staff was able to install the monopolar cautery hook tip into the monopolar cautery instrument with no problem. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Electrocautery tip accessory installation - to install the electrocautery tip accessory (the electrode), insert the metal pin of the electrocautery tip accessory into the slot at the distal tip of the instrument. Grasping the silicone sleeve that protects the electrocautery tip accessory, rotate the accessory until it contacts the sleeve stop. The metal electrode will not rotate with the sleeve. Do not over tighten. Remove the silicone sleeve. The silicone sleeve may be discarded or it may be retained to aid in removing the electrocautery tip accessory after the clinical application. | MONOPOLAR CAUTERY HOOK 5MM ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3004944 | 2979594 | 2876108 | 2955842201300640 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-01-28 | 2013-02-27 | No | Yes | Health Professional | 400156 | No | No | 2013-01-28 | No | No | Unkown | 2013
42030901 | 2013-01-28 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a broken wire on the mega suture cut needle driver instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken near the proximal pulleys and the proximal clevis cable opening. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment sticks out at wrist. Additional observation not reported by site was proximal clevis wearing. The proximal clevis cable opening exhibited wear on one edge. Other cables located at the wrist were not damaged. Evidence not conclusive, but wear on hole may have contributed to the cable breakage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3004955 | 2979610 | 2876124 | 2955842201300641 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-28 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted frayed wires on the maryland bipolar forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a frayed pitch cable was found at the distal idler which was also found to be broken. There was no damage found on pulley or clevis. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and they were not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3004956 | 2979611 | 2876126 | 2955842201300642 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-27 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-28 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted frayed wires on the maryland bipolar forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was found to be broken with fraying too at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and a rough surface finish. The scratches were short in length and they were not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3004945 | 2979612 | 2876126 | 2955842201300643 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-27 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-28 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, the wire on the mega suturecut needle driver instrument was twisted so the instrument did not have full range of motion. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was frayed at the distal idler pulley and the frayed strands were stuck out at the wrist. The other cables at wrist were not damaged. Engineering evaluation also found the grip cable was derailed. Although the cable was derailed at the distal idler pulley the jaws still opened and closed but the movement was not precise. Engineering concluded that cable derailment was likely due to the cable losing contact with pulley during wrist articulation thus causing the non-intuitive motion experienced by the customer. The fleet angle between proximal and distal pulleys may have contributed to the derailment. In addition, engineering evaluation also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3005869 | 2980721 | 2870247 | 2955842201300644 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the surgical staff allegedly noticed that the mega suturecut needle driver instrument had a broken cables. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint that the instrument had a broken cable. The grip cable was found to be broken at the wrist. The cable segment was observed to be sticking out at the wrist. Engineering evaluation also found the pulley with damage to the face and rim. The idler pulley was observed to spin freely. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3005957 | 2980776 | 2877290 | 2955842201300645 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-02-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
40020504 | 2013-01-14 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the fenestrated bipolar forceps instrument did not close. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the clamp did not close. The one grip was bent, causing side to side misalignment of grips. There was a. 035 offset at the tips. The bent grip had separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. The same grip was also twisted so that one side of grips had a. 050 gap when grips were in the closed position. The grasping functionality was diminished due to the grip damage. Additional observation was various scratch marks with light material removal at the distal end of the main tube. The scratches were. 060 -. 200 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3006013 | 2980812 | 2877326 | 2955842201300646 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-01-29 | 2013-02-27 | No | Yes | Health Professional | 40020504 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-01-29 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff allegedly noticed that the large suturecut needle driver instrument had a broken wire at the distal pivot end. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint that the instrument had a broken wire at the distal pivot. One grip close cable was found to be broken near the proximal pulleys and proximal clevis cable hole. The cable segment was found to be sticking out at the wrist. The idler pulley was observed to spin freely and did not exhibit damage. Engineering evaluation also found wearing at the proximal clevis. The proximal clevis cable hole was found to exhibit wear on one edge. Other cables at the instrument's wrist were not damaged. Engineering evaluation concluded that the wear on the hole might have contributed to the cable breakage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3006075 | 2980869 | 2877383 | 2955842201300648 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-02-27 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42004906 | 2012-12-03 | Malfunction |  | It was reported prior to starting a da vinci s surgical procedure that the cadiere forceps instrument tip did not work. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the nonconformance was a broken grip cable. The intuitive motion was not being experienced because the grip cable was broken at the distal idlers. The idler pulley spun freely and exhibited rim pulley damage. The cable segment stuck out at the wrist. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3006069 | 2980871 | 2877385 | 2955842201300647 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-01-29 | 2013-02-27 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-18 | Malfunction |  | It was reported during reprocessing that the pk dissecting forceps instrument was noted to have a broken cable inside, the forceps were only moving in one direction. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The electrical continuity testing passed. 48 - engineering also found the one grip was bent, causing side to side misalignment of the grips. There was a. 049 offset at the tips, indicating overloading at the tip. There were also various scratches on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage to the main tube was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3006142 | 2980903 | 2877417 | 2955842201300649 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-02-27 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-18 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the surgical staff allegedly noticed that the prograsp forceps instrument had a broken wire at the distal pivot end. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint that the instrument had a broken wire. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. 48 - engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering evaluation concluded that the scratches were likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable and main tube scratch issues if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3006170 | 2980931 | 2877445 | 2955842201300650 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-02-27 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-08-02 | Malfunction |  | It was reported during a da vinci si surgical procedure that the maryland bipolar forceps instrument tips would not open. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found there was a frayed pitch cable. The instrument was found with frayed pitch cable at the distal idler with scratches on both distal pulleys. Both distal pulleys had scratches on the surface. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering found misaligned tips. The one grip was bent, causing side to side misalignment of grips. There was a. 076 offset at the tips, indicating overloading at the tip. The electrical continuity testing passed. Engineering concluded that damage to the main tube was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3006198 | 2980965 | 2877479 | 2955842201300651 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-08 | 2013-02-27 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-08 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-31 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the pk dissecting forceps instrument was noted to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that there was a broken pitch cable. The pitch cable was broken at the wrist. The cable segment that contained the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS ISNTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3006251 | 2981057 | 2877571 | 2955842201300652 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-31 | 2013-02-27 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
41032204 | 2013-01-31 | Malfunction |  | It was reported during a da vinci si surgical procedure that a piece of the endowrist one vessel sealer instrument broke off and fell into the patient's pelvic cavity. The piece was retrieved without harm to the patient. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The representative was contacted and indicated that the instrument piece was retrieved from the patient laparoscopically. The surgeon was performing the surgical task of lar when the instrument malfunction occurred. It was indicated that the vessel sealer instrument did not make contact with any other instrument during the surgical procedure. There was no harm to the patient and the planned surgical procedure was completed. The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | ENDOWRIST ONE VESSEL SEALER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3006313 | 2981129 | 2877643 | 2955842201300654 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-31 | 2013-02-27 | No | Yes | Health Professional | 41032204 | No | No | 2013-01-31 | No | Yes | Unkown | 2013
41032204 | 2013-01-28 | Malfunction |  | It was reported that during a davinci si surgical procedure, the customer noted broken wires on the vessel sealing instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | VESSEL SEALING INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3006357 | 2981169 | 2877683 | 2955842201300653 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-27 | No | Yes | Health Professional | 41032204 | No | No | 2013-01-28 | No | Yes | Unkown | 2013
42003305 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy procedure the black diamond micro forceps instrument tips did not align. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one tip was bent close to the tip area. The grips should be straight against the other tip. Engineering concluded that damage to the tips were likely due to mishandling. Engineering also found indentations on the edges of the pulley and some scratches on the surface of the pulley. There were also various scratches on the distal end of the maintube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage to the main tube was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | BLACK DIAMOND MICRO FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3006365 | 2981173 | 2877687 | 2955842201300655 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-06 | 2013-02-27 | No | Yes | Health Professional | 42003305 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-02-01 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff allegedly noticed that the mega suturecut needle driver instrument had broken wires. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint that the instrument had a broken wire. One grip close cable was found to be broken at the distal idlers pulley. The idler pulley was observed to spin freely. The cable segment was sticking out at the instrument's wrist. Other cables at the wrist were not damaged. Engineering evaluation also found damage on the edge and on the surface of the idler pulley. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3006483 | 2981252 | 2877766 | 2955842201300656 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-02-27 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-01-21 | Malfunction |  | It was reported that during a davinci si colon resection procedure, while using the bowel grasper instrument, the cable broke during the case, nothing fell into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed broken cables segment sticking out at snake wrist of the instrument, no other signs of damage were noted. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3006661 | 2981384 | 2877898 | 2955842201300657 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-28 | 2013-02-27 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | Yes | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-11 | Malfunction |  | It was reported that during a davinci si surgical procedure, the pk dissecting forceps instrument stopped moving, nothing fell in the patient and no patient injury. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3006788 | 2981517 | 2878031 | 2955842201300658 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-28 | 2013-02-27 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-28 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff allegedly noticed that the fenestrated bipolar forceps instrument had a torn cable and was twisted. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint that the instrument had a torn cable. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was missing in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. Engineering evaluation also found a bent tip. One grip was bent, causing side to side misalignment of the grips. There was a. 111 offset at the tips, indicating overloading at the tip. Engineering evaluation concluded that the bent tip was most likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3007055 | 2981767 | 2878281 | 2955842201300659 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-02-01 | 2013-02-27 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
| 2013-01-28 | Injury | Required Intervention | The vessel sealer was noted during the case to have wires unraveling. |  | ENDOWRIST ONE VESSELSEALER | VESSEL SEALER | 3020139 | 2993192 | 2889728 | MW5029210 | 1 | NAY | Voluntary |  | RISK MANAGER | Initial | 2013-02-21 | 2013-02-27 | No | Yes | Health Professional |  | Yes | Yes |  |  |  |  | 2013
420309 | 2013-02-22 | Malfunction |  | Davinci mega suture cut needle driver quit working during case. Device was removed from field, inspected and noted to have broken parts. What was the original intended procedure?laparoscopic nissen fundoplication. Device #1is this a laboratory device or laboratory test?no. |  | ENDOWRIST MEGA SUTURECUT | SYSTEM SURGICAL | 3059877 | 3032439 | 2928953 | 3032439 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-02-27 | 2013-02-27 | No | Yes | Physician | 420309 | Yes | No Answer Provided |  |  |  |  | 2013
40027202 | 2013-02-08 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, a piece got stuck at the distal end of the harmonic ace curved shears insert, and the surgical staff was unable to remove it; this was identified during set up, the harmonic ace curved shears insert was not used in the case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that entire teflon pad was missing from clamp arm. The clamp arm was still able to open/close, but insert was not fully functional without the teflon pad. Engineering concluded that damage was likely due to mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HARMONIC ACE CURVED SHEARS INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 3004119 | 2978894 | 2875408 | 2955842201300628 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-28 | 2013-02-26 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | Yes | Reuse | 2013
42009308 | 2013-01-28 | Malfunction |  | This instrument was returned without an initial complaint. Failure analysis found visible main tube damage on the prograsp forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were. 080 -. 125 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3004243 | 2978976 | 2875490 | 2955842201300629 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-26 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the surgical staff allegedly noticed that the fenestrated bipolar forceps instrument had broken cables. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint of a broken cable. A conductor wire was found to be broken at the yaw pulley exit. The wire was detached from its connection at the grip. Electrical continuity testing of the instrument failed. Engineering evaluation also observed that the yaw pulley had a missing piece opposite the conductor break which allowed the grip to move within the pulley and have a larger range of motion in yaw. Engineering evaluation was unable to determine how the piece of yaw pulley broke. The instrument was returned expired. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken pulley issue if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3004378 | 2979047 | 2875866 | 2955842201300630 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-02-26 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy (bso) procedure, the surgical staff allegedly noticed that the fenestrated bipolar forceps instrument had a broken wire. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint of a broken cable. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Engineering evaluation also found a bent grip tip. One grip was bent, causing side to side misalignment of grips. There was a. 3020 offset at the tips. Engineering evaluation concluded that the damage was likely due to mishandling. Electrical continuity testing of the instrument passed. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3004556 | 2979211 | 2874122 | 2955842201300631 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-31 | 2013-02-26 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-31 | Malfunction |  | It was reported that prior to starting a da vinci si total benign hysterectomy procedure, the surgical staff allegedly noticed that the mega suturecut needle driver instrument had a broken wire. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint of a broken wire. One grip close cable was found to be broken at the distal idlers. The idler pulley was observed to spin freely and was not damaged. The cable segment was found to be sticking out at the wrist. Other cables at wrist are not damaged. Engineering evaluation also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 3004613 | 2979241 | 2875755 | 2955842201300632 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-31 | 2013-02-26 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-25 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s surgical procedure the fenestrated bipolar forceps instrument was noted to be cut. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | Reuse | 3004598 | 2979260 | 2875772 | 2955842201300633 | 1 | NAY | Manufacturer | DistributorOther | Other | InitialFollowup | 2013-03-20 | 2013-02-26 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No |  | 2013
42020505 | 2012-09-23 | Malfunction |  | Event Description: --/-- It was reported during a da vinci si cholecystectomy procedure the fenestrated bipolar forceps instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was broken. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | Reuse | 3004651 | 2979263 | 2875775 | 2955842201300635 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-03-01 | 2013-02-26 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No |  | 2013
42020505 | 2013-01-23 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff alleged that the wheels on the bottom of the blue housing of the fenestrated bipolar forceps instrument would turn but the device would not respond. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint that the wheels on the bottom of the blue housing would turn but does not do anything. The pitch input disc was observed to rotate freely without corresponding output motion at the instrument's wrist. Visual inspection of the instrument showed the pitch down cable to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3004653 | 2979300 | 2875814 | 2955842201300634 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-31 | 2013-02-26 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-28 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure the prograsp forceps instrument was noted to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3004655 | 2979304 | 2875818 | 2955842201300636 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-26 | No | Yes | Health Professional | 42009308 | No | No | 2013-01-28 | No | No | Unkown | 2013
42019408 | 2013-01-28 | Malfunction |  | It was reported prior to starting a da vinci s surgical procedure the mega needle driver instrument was noted to have fibers exposed at the end of the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the nonconformance was a broken grip cable. The grip cable was broken at the wrist. The idler pulley spun freely and exhibited pulley face damage. The cable segment stuck out at the wrist. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3004658 | 2979306 | 2875820 | 2955842201300637 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-26 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-28 | Malfunction |  | It was reported that during central reprocessing the megasuturecut needle driver instrument was noted to have a broken cable. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely but had damage on the edge and on the surface. The cable segment stuck out at the wrist. The other cables at wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3004604 | 2979310 | 2875824 | 2955842201300638 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-26 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-23 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff allegedly noticed that the fenestrated bipolar forceps instrument had a broken cables. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint that the instrument had a broken cable. The pitch cable was found to be broken at the wrist. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3004682 | 2979352 | 2875866 | 2955842201300639 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-01 | 2013-02-26 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-02-01 | Device Not Retu | No | Reuse | 2013
410322 | 2013-02-26 | Malfunction | Other | During robotic procedure, the vessel sealer was connected to robotic arm. While in pt, the machine/computer reported malfunction, alarmed with exposed blade reported. However, no blade noted by any staff in the room. Vessel sealer removed from pt and removed from field. New vessel sealer opened and put into service w/o issue. No harm to pt. |  | ENDOWRIST ONE VESSEL SEALER | VESSEL SEALER | 3020338 | 2993165 | 2889721 | MW5029202 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-02-26 | 2013-02-26 | No | Yes | Health Professional | 410322 | Yes | Yes |  |  |  |  | 2013
42017704 | 2012-11-07 | Malfunction |  | It was reported that during a davinci si gastric bypass procedure, the bowel grasper instrument stopped working. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with broken cable inside the snake wrist. Additional observation not reported by the site was main insulation tube damage. The main insulation tube exhibited multiple damage with some materials removed. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | BOWEL GRASPER 30MM JAW INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3000916 | 2975638 | 2872152 | 2955842201300607 | 1 | NAY | Manufacturer | DistributorUser facility | Other | Initial | 2013-01-23 | 2013-02-25 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-25 | Malfunction |  | It was reported that during reprocessing of the mega suture cut needle driver instrument the wires at the distal end of the instrument were observed to be frayed. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that grip cable was broken at the distal idlers. The cable segment stuck out at the wrist. The idler pulley was found to be physically damaged. The idler pulley exhibited indentations. Engineering concluded that the damage was likely due to misuse. The instruments and accessories user manual specifically states: general precautions and warnings - handle instrument with care. Avoid mechanical shock or stress that can cause damage to the instruments - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3000955 | 2975655 | 2872169 | 2955842201300610 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-25 | 2013-02-25 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2010-09-29 | Injury | Required Intervention | On (b)(6) 2013, a patient contacted isi employee to report that on (b)(6) 2010, he had a mitral valve repair surgery performed with a da vinci s surgical robot system at (b)(6). According to the information statement provided by the patient, it was indicated that he did not feel well after the surgery, and he had to return for a second open surgery to repair three problems, which he was told were a result of the first da vinci procedure. Reportedly, the patient was informed by the second surgeon who performed the corrective surgery that the first da vinci robotic surgery procedure took 90 minutes longer than it was originally planned, and that the problems he suffered were due to errors made by the first surgeon. The patient indicated that he still has atrial fibrillation issues, and he is now wearing a pacemaker. | On (b)(4) 2013, isi received the medical records from the patient. According to the medical records, on (b)(6) 2010, the patient underwent a da vinci mitral valve repair consisting of 2 neo chords 4-0 gore-tex suture, reduction annuloplasty using a flexible annular ring and a cryomaze procedure. According to the medical records, the patient tolerated the procedure well, was then transferred to the intensive care unit (icu) in stable condition in normal sinus rhythm. On (b)(6) 2011, the patient had a severe symptomatic recurrent mitral regurgitation procedure performed. According to the medical records, a radical reconstruction redo mitral valve repair, resection of p2 rupture chord, removal of a previous annuloplasty ring, and placement of another annuloplasty ring was performed. Per medical records, the patient was then transferred to the icu in stable condition. A review of the system logs shows no system malfunction occurring during the procedure performed on (b)(6) 2010. A follow-up mdr will be submitted if a device if additional information is received. We subsequently learned that the surgeon who performed the initial procedure is deceased as of (b)(6) 2012. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3000943 | 2975656 | 2872170 | 2955842201300608 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-24 | 2013-02-25 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2013-01-24 | No | No | Reuse | 2013
42022703 | 2013-01-16 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the site was unable to release the grips on the pk dissecting forceps instrument. No missing or fallen pieces were reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering that the instrument was returned with the grips fully closed and the grips did not open when torque was manually applied to the input discs. The instrument grips popped open during engagement testing of the instrument after it was installed on an in-house is3000. A visual inspection of the instrument grips found no excessive biodebris. Engineering evaluation also found that the pitch down cable was frayed at the distal clevis hub. The frayed strands stuck out at the wrist. The other cables at the wrist were not damaged. The distal end of the main tube had a scratch mark with light material removal. The scratch was. 270 in length and was not aligned with the tube axis. Engineering concluded that the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instrument with care. Avoid mechanical shock or stress that can cause damage to the instruments - do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument pitch down cable and main tube found during failure analysis could likely cause or contribute to an adverse event if the malfunctions were to recur. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3000951 | 2975699 | 2872213 | 2955842201300609 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-24 | 2013-02-25 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-24 | Malfunction |  | It was reported that during reprocessing the maryland bipolar forceps instrument, it was noted that the tips do not align. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip was bent, causing side to side misalignment of the grips. There was a 0. 061 offset at the tips. The grip tips did not meet together when fully closed. Engineering concluded that damage was likely due to mishandling. Electrical continuity testing passed. Engineering also found a frayed pitch cable at the wrist. The clevis does not exhibit any damage or wear marks. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3000952 | 2975701 | 2855631 | 2955842201300611 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-04 | 2013-02-25 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-04 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-01-24 | Malfunction |  | Event Description: --/-- It was reported that during dissection of the bilateral ima mobilization, the surgeon used the micro bipolar forceps instrument to clean the permanent cautery spatula instrument and broke the collar around the permanent cautery spatula element, the item was retrieved without any harm to the patient. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; however, based on the event description provided during the initial complaint. It was concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that an uneven piece of the ceramic sleeve was missing exposing the shaft of the spatula. Engineering concluded that damage was likely due to mishandling. Additional observation not reported by site was a derailed yaw cable. The yaw cable was derailed at the wrist. The spatula movement and functionality could not be precise. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction of broken ceramic sleeve if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | Reuse | 3000991 | 2975769 | 2872283 | 2955842201300612 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-03-12 | 2013-02-25 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-03-12 | Device Not Retu | No |  | 2013
42000606 | 2013-01-18 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a broken wires on the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was found to be frayed. The frayed cable sections were sticking out at the wrist in various lengths. Additional finding not reported by the site was idler pulley exhibited rim damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3001164 | 2975957 | 2872471 | 2955842201300614 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-24 | 2013-02-25 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
IS3000 A60P8 | 2013-01-29 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the surgeon experienced a yellow halo in the left eye image through the high resolution stero viewer. With the assistance of an isi technical support engineer, the customer swapped to their backup camera head, camera cable and light guide; however, the issue occurred with the right eye image. The tse instructed the sight to swap the core video cables; however, the vision issue persisted. With the assistance of the tse, the site performed additional troubleshoot techniques; however, the issue was not resolved. The site informed the tse that the surgeon was going to perform the surgical procedure in 2d and with the assistance of the tse, the site converted the system to 2d mode. Approximately 1. 5 hrs later the isi clinical sales representative contacted isi technical support and indicated the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. | The investigation conducted by the field service engineer (fse) concluded that the vision issue experienced by the site was associated with the 1st camera cable and the backup camera cable. Inspection of the camera cables by the fse found that both were worn and were bulged on the side that connects to the camera head. Also, during the fse's investigation he was advised by the isi clinical sales representative that the site had a 2nd back up camera cable; however, the site was unable to locate it when the vision issue occurred. The fse installed and tested the cable on the site's system found that it functioned within specifications, as the left and right images were observed to be normal. The camera cable connects the camera to the system's vision cart, which then transmits the image to the surgeon's console. The system was repaired by replacing the affected camera cables. The fse also inspected the site's camera heads and found that one of the cameras exhibited cracks on the housing and that the light guide assembly on the camera head was loose. The camera head produces three dimensional images to the da vinci vision system through right and left optical paths. The images are then acquired through separate optical channels of the endoscope and are directed to the camera head and processed independently. On (b)(4) 2013, isi contacted the initial reporter concerning the reported event. The initial reporter indicated that the surgeon made the decision to complete the planned surgical procedure using open surgical techniques as a result of the camera cable malfunctions. The initial reporter indicated that to the best of her knowledge there was no injury to the patient as a result of the camera cable malfunctions. The initial reporter also indicated that to the best of her knowledge the patient tolerated the open surgical procedure well and the patient has not returned to the hospital due to any post surgical complications as a result of the reported event. The initial reporter indicated that the 2nd backup camera cable was not located until after the surgeon had converted the planned surgical procedure to open. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3001193 | 2976020 | 2872534 | 2955842201300613 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-29 | 2013-02-25 | Yes | Yes | Health Professional | IS3000 A60P8 | Yes | Yes | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-17 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was noted that blue handle fell apart on the maryland bipolar forceps. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the housing had one rear snap tab and two locating pins completely broken off. The front snap of the same side had the retaining tab feature broken. One side of the housing could be lifted up as a result of damage. Additional observations not reported by site were tube damage, missing luer plate, and bent bipolar pins. The distal end of main tube exhibited various scratch marks with light material removal. The scratches were. 100 -. 210 in length and not aligned with the tube axis. The luer plate was missing from the chassis. The flush tube was still installed. The plate likely had come off after the housing broke. Bottom bipolar pin was bent upwards. The bipolar cord could not be properly installed. Both pins exhibited scratch marks. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3001252 | 2976072 | 2872586 | 2955842201300617 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-24 | 2013-02-25 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-24 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a broken cable at the tip of the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley spins freely and they did not exhibit any damage. Cable segment sticks out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3001262 | 2976080 | 2864706 | 2955842201300615 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-24 | 2013-02-25 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-01-18 | Malfunction |  | It was reported that during a davinci si prostatectomy procedure, the customer noted a broken cable on the permanent cautery spatula instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3001271 | 2976135 | 2872649 | 2955842201300616 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-24 | 2013-02-25 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
40016907 | 2013-01-25 | Malfunction |  | Event Description: --/-- It was reported that during a davinci si hysterectomy procedure, the customer used 3 harmonic curved shears insert, according to the information provided by the customer, the 1st and the 2nd harmonic curved shears insert broke after 15 minutes of use. During the use of the 2nd harmonic curved shears insert, the surgeon noticed weakness on the harmonic curved shears insert and decided to remove it from the patient. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm became detached. The harmonic curved shears insert was returned without the clamp arm. The distal end of the shaft exhibited features that mate with the clamp arm; the pin bent backwards on one side. The direction of the bending indicated hyperextension of the clamp arm. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- Based on additional information provided by the nurse, he indicated that no fragments fell into the patient during the use of the 2nd or 3rd harmonic curved shears inserts. The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. First harmonic curved shears mdrs 2955842-2013-00618 and 3rd harmonic curved shears insert mdr 2955842-2013-00620. | HARMONIC CURVED SHEARS INSERT | Unkown | 3001311 | 2976164 | 2872697 | 2955842201300619 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2013-03-01 | 2013-02-25 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | Yes | 2013-03-01 | Device Not Retu | Yes |  | 2013
40016907 | 2013-01-25 | Injury | Required Intervention | Event Description: --/-- It was reported that during a davinci si hysterectomy procedure, the customer used 3 harmonic curved shears insert, according to the information provided by the customer, the 1st and the 2nd harmonic curved shears insert broke after 15 minutes of use. On the first instrument, a tip fell into the patient, surgeon was able to remove it. Chief surgeon confirm that, after performing radio exam and a ct scan to the patient, there was no fragment left. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm was found to be missing from the wrist and the metal casing was bent. The harmonic curved shears insert blade was bent backwards. Engineering concluded that damage was likely due to excessive side loading or other type of mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. 2nd harmonic curved shears mdrs 2955842-2013-00619 and 3rd harmonic curved shears insert mdr 2955842-2013-00620. | HARMONIC CURVED SHEARS INSERT | Unkown | 3001228 | 2976183 | 2872697 | 2955842201300618 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2013-03-01 | 2013-02-25 | Yes | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | Yes | 2013-03-01 | Device Not Retu | Yes |  | 2013
40016907 | 2013-01-25 | Malfunction |  | Event Description: --/-- It was reported that during a davinci si hysterectomy procedure, the customer used 3 harmonic curved shears insert, according to the information provided by the customer, the 3rd harmonic curved shears insert broke, but no fragments fell into the patient. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the clamp arm became detached. The harmonic curved shears insert was returned without the clamp arm. The distal end of the shaft exhibited features that mate with the clamp arm; the pin bent backwards on one side. The direction of the bending indicated hyperextension of the clamp arm. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- Based on additional information provided by the nurse, he indicated that no fragments fell into the patient during the use of the 2nd or 3rd harmonic curved shears inserts. The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. First harmonic curved shears mdrs 2955842-2013-00618 and 2nd harmonic curved shears insert mdr 2955842-2013-00619. | HARMONIC CURVED SHEARS INSERT | Unkown | 3001306 | 2976213 | 2872697 | 2955842201300620 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2013-03-01 | 2013-02-25 | No | Yes | Health Professional | 40016907 | Device Returned To Manufacturer | Yes | 2013-03-01 | Device Not Retu | Yes |  | 2013
42009308 | 2012-11-06 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure, the surgical staff alleged that one of the control wires was broken on the prograsp forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of a broken wire. The instrument was found to have a frayed pitch cable at the distal clevis hub with loose tension. No damage was found at the clevis. The other cables of the instrument's wrist did not exhibit any damage. In addition, the instrument's housing was removed and the pitch cable was found to be derailed on the clamping pulley. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | Reuse | 3002017 | 2976915 | 2873429 | 2955842201300621 | 1 | NAY | Manufacturer | ForeignUser facility | Other | FollowupInitial | 2013-03-20 | 2013-02-25 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No |  | 2013
42020505 | 2013-01-22 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff alleged that the tips of a fenestrated bipolar forceps instrument were twisted and stuck. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results did not confirm the alleged complaint that the instrument's tips were twisted and stuck. The grip tips were found to be aligned. The grip tips met when the grips were opened and closed. Engineering evaluation also observed a broken pitch cable. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was found to be missing. The clevis did not exhibit any damage or wear marks. On (b)(4) 2013, isi contacted the initial reporter of this complaint to obtain additional information. The initial reporter denied that any fragments fell and no patients were harmed because of the alleged issue. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3002069 | 2976978 | 2854184 | 2955842201300622 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-29 | 2013-02-25 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-22 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff alleged that the tips of a prograsp forceps instrument were twisted. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results did not confirm the alleged complaint that the instrument's tips were twisted. Based on visual inspection, the instrument's tips were not twisted. The instrument was checked for movement on an in-house system. The instrument did not move intuitively, which required additional observation. Engineering evaluation also found frayed pitch cables. The instrument was found with frayed pitch cable. The cables on both yaw clamping pulleys were found to be derailed. Also, the left clamping pulley had excessive debris. Engineering evaluation concluded that the debris could have been due to the cleaning process and the observations were likely the cause of the instrument's poor performance. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. On (b)(6) 2013, isi contacted the initial reporter of this complaint to obtain additional information. The initial reporter denied that any fragments fell. He also denied that any patients were harmed because of the alleged issue. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported frayed cable issue if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3002247 | 2977098 | 2873612 | 2955842201300623 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-29 | 2013-02-25 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that the fenestrated bipolar forceps instrument had broken wires. The surgical staff also claimed that the instrument was acting weird when the cautery was fired. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint of a broken wire. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. Engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering evaluation concluded that the scratches were likely due to mishandling of the instrument. Another finding was scratches on the surface of the distal pulley. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable and main tube scratch issues if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3002527 | 2977313 | 2873827 | 2955842201300624 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-30 | 2013-02-25 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-30 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the surgical staff alleged that the pk dissecting forceps instrument had broken wires. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint of a broken wire. The pitch cable was found to be broken at the wrist. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3002620 | 2977398 | 2873912 | 2955842201300625 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-31 | 2013-02-25 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-31 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff alleged that the pk dissecting forceps instrument had a broken cable at the tip. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint of a broken cable. The pitch cable was found to be broken at the wrist. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3002756 | 2977532 | 2874046 | 2955842201300626 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-31 | 2013-02-25 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy (bso) procedure, the surgical staff alleged that the fenestrated bipolar forceps instrument had a broken wire. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint of a broken cable. The pitch cable was found to be broken at the wrist. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3002840 | 2977608 | 2874122 | 2955842201300627 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-31 | 2013-02-25 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-01-20 | Malfunction |  | It was reported that prior to starting a da vinci si myomectomy procedure, the cable on one side of the permanent cautery hook instrument tip was coming out. No patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | PERMANENT CAUTERY HOOK | ENDOSCOPIC ELECTROSURGICAL INS | 2999107 | 2973417 | 2869933 | 2955842201300600 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-22 | No | Yes | Health Professional | 42018306 | No | No | 2013-01-28 | Device Not Retu | No | Unkown | 2013
42030901 | 2013-01-25 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, a broken cable was noticed on the mega suturecut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken at the distal idlers. Idler pulley spins freely and does not exhibit damage. Cable segment sticks out at wrist. Other cables at wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2999424 | 2973731 | 2870247 | 2955842201300598 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-25 | 2013-02-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-23 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted broken wires at the tip of the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be frayed at distal clevis hub. No damage was found at the clevis. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2999859 | 2974171 | 2870705 | 2955842201300601 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-22 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
40018012 | 2013-01-03 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the end of the tip cover burned. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received and/or post engineering analysis if the instrument accessory is returned. | 8 MM MCS TIP COVER ACCESSORY | Unkown | 2999933 | 2974290 | 2867077 | 2955842201300602 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-01-24 | 2013-02-22 | No | Yes | Health Professional | 40018012 | No | No | 2013-01-24 | Device Not Retu | Yes |  | 2013
42009308 | 2013-01-26 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the wire of the prograsp forceps instrument was broken. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation confirmed the alleged issue that the instrument had a broken wire. Engineering evaluation found the pitch cable broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. Engineering evaluation indicated that the clevis did not exhibit any damage or wear marks. The other cables of the instrument's wrist did not exhibit any damage. Engineering evaluation also found a deep scratch on the instrument's main tube. The distal end of the main tube had a scratch that exhibited light material removal and a rough surface finish. The scratch was found to be approximately. 6270 long at distal end. Engineering concluded that the main tube damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP TM FORCEPS | Reuse | 3000038 | 2974429 | 2870945 | 2955842201300603 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-02-22 | 2013-02-22 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-03-01 | Device Not Retu | No |  | 2013
420172 | 2013-01-11 | Injury | Required Intervention | It was reported that during a da vinci si presacral neurectomy procedure, when the surgeon attempted to move the patient's sigmoid colon out of the patient's pelvis using the maryland bipolar forceps instrument, cautery energy was observed coming from the instrument, causing a. 5 cm burn to the patient's sigmoid colon. Reportedly, the instrument energized without activation from the foot pedal by the surgeon at the surgeon side console. Investigation performed by the site found that the maryland bipolar forceps energy cable was inadvertently installed into the monopolar receptacle on the site's electrical surgical unit (esu). The bipolar energy cable was installed in proper receptacle on the esu and the maryland bipolar instrument functioned properly. The damage to the patient's sigmoid colon was repaired by the site's general surgeon, who over sewed the affected area with a vicryl suture. Reportedly, there was no evidence that the patient's sigmoid colon was perforated or any other injury to the patient occurred. The patient was kept overnight for observation and was discharged the following morning. | The instrument has not been returned for evaluation; however, based on the information provided by the initial reporter, the instrument did not malfunction in a way that caused or contributed to the patient's injury, but was caused by the site's inadvertent installation of instrument's energy cable into the improper receptacle on the electrical surgical unit. On (b)(4) 2013, isi contacted the surgeon, who was the initial reporter of this event. The surgeon indicated that the planned surgical procedure was completed, the patient has had two follow up visits since the surgery, is recovering well and has not experienced any post-surgical complications as a result of the reported event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3000405 | 2974811 | 2871325 | 2955842201300604 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-23 | 2013-02-22 | Yes | No | Health Professional | 420172 | No | No | 2013-01-23 | No | No | Unkown | 2013
42030901 | 2013-01-23 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted that the mega suture cut needle driver instrument's cable was sticking out at the end. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be frayed at the distal idler pulley. The frayed strands stick out at the wrist. Other cables at wrist were not damaged. Additional observation not reported by site was a derailed cable. Same grip close cable that was frayed was also found to be derailed from the distal idler pulley. The grips were able to open and close, but their movement could not be precise. Evidence not conclusive, but cable derailment was likely due to cable losing contact with pulley during wrist articulation. Derailment likely contributed to the fraying cable. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 3000474 | 2974854 | 2871370 | 2955842201300605 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-21 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the maryland bipolar forceps instrument was not recognized. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument successfully passed recognition and engagement testing on an in-house is3000 system and showed that the instrument had 3 uses remaining. The pogo pins did not stick and there was no evidence of damage to the dallas chip. The instrument moved intuitively and the grip opened and closed when test driven. Engineering evaluation also found that the bottom pin was not aligned with the top pin and it was slightly bent towards the right. Engineering concluded that damage to the pin was likely due to mishandling as result of inadvertent contact of the instrument with an object while moving the instrument from and to the system. Main tube shows various scratches near distal end and is missing material. The main tube exhibited two scratches along the proximal end which measured. 25 (closer to proximal end) inches and. 5 inches and is missing material. The instrument passed electrical continuity testing. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; damage to the instrument's main tube found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3000505 | 2974869 | 2871385 | 2955842201300606 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-25 | 2013-02-22 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-17 | Malfunction |  | It was reported that during a davinci s tongue base resection procedure, the customer noted a broken wires on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2997753 | 2972286 | 2868802 | 2955842201300585 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-21 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci si surgical procedure the cable on the large suturecut needledriver 8mm instrument broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLEDRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2997773 | 2972302 | 2868818 | 2955842201300577 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-22 | 2013-02-21 | No | Yes | Health Professional | 42029602 | No | No | 2013-01-22 | No | No | Unkown | 2013
42004906 | 2013-01-14 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si surgical procedure the cadiere forceps instrument had broken wires. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that the damaged instrument component was a frayed pitch cable. The instrument was found with frayed pitch cable at the distal clevis hub. No damage was found at the clevis. No other cable damage was found. Engineering also found indentations and scratches on the edge of the pulley plus scratches on the surface of the distal pulley. There were also various scratch marks on the main tube at the distal end showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | Reuse | 2997780 | 2972308 | 2868824 | 2955842201300578 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-03-07 | 2013-02-21 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-03-07 | Device Not Retu | No |  | 2013
42009308 | 2013-01-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure that the wires came off the track of the prograsp forceps instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the wires were not off the track. Engineering found the maintube had various scratches at the distal end that showed light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2997850 | 2972359 | 2868875 | 2955842201300579 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-02-14 | 2013-02-21 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-02-14 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-14 | Malfunction |  | It was reported that during a da vinci si surgical procedure the megasuturecut needle driver instrument cable broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument had a frayed grip cable. The frayed cable sections were in various lengths sticking out at the wrist. The grip cable was derailed as well. The grips could still open and close, but movement and functionality may not be precise. The idler pulley also exhibited pulley face and rim damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2997851 | 2972381 | 2868894 | 2955842201300576 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-31 | 2013-02-21 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-17 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, it was noted that one of the wheels of the fenestrated bipolar instrument was not moving the wrist properly. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the manual input of the disc wheel did not generate intuitive motion because the pitch cable was found to be broken at the distal clevis hub. The cable crimp was missing in the clevis. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2997890 | 2972395 | 2868911 | 2955842201300586 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-21 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-01-31 | Malfunction |  | It was reported prior to starting a da vinci s surgical procedure that the tenaculum forceps instrument was noted to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch up cable was broken at the distal clevis hub. The cable segment that contained the crimp is missing. The clevis did not exhibit any damage or wear marks. The other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2997939 | 2972427 | 2868940 | 2955842201300582 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-31 | 2013-02-21 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-22 | Malfunction |  | It was reported during a da vinci s surgical procedure that the fenestrated bipolar forceps instrument was noted to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. Engineering also observed various scratch marks with light material removal on the distal end of the main tube. The scratches were. 090 -. 500 in length and not aligned with the tube axis. Engineering concluded that damage to the main tube was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2997912 | 2972436 | 2868952 | 2955842201300581 | 1 | NAY | Manufacturer | ForeignHealth Professional | Other | Initial | 2013-01-29 | 2013-02-21 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-01-31 | Malfunction |  | It was reported during a da vinci s thymectomy procedure the permanent cautery hook instrument tip was noted to broken. During the procedure the instrument tip was turned at an odd angle. The surgeon attempted to rotate the instrument to a neutral position but it still appeared to be folded on itself. Upon closer examination with the scope, the tip of the plastic end of the instrument appeared to be broken; the instrument was removed from the patient. It was determined after searching for loose parts in the patient, it was determined no pieces were in the patient. The case was completed with a different instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument was returned with the yaw pulley and hook missing from the distal end. Both yaw cables and the conductor wire were cut at the wrist. The distal clevis had one ear broken off near its base. The missing ear piece measures approximately. 235 x. 285. Engineering concluded the likely failure mode was a broken distal clevis ear as a result of overloading the tip. The cables and wire were likely cut after the clevis breakage to enable instrument removal from the cannula. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2997929 | 2972447 | 2868963 | 2955842201300583 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Health Professional | Initial | 2013-01-31 | 2013-02-21 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | Yes | 2013-01-31 | Device Not Retu | No | Reuse | 2013
IS3000 A60P3 | 2013-01-25 | Other |  | Event Description: --/-- It was reported that during a da vinci si cardiac-thoracic procedure, the surgeon experienced a dark image through the high resolution stereo viewer (hrsv). The surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | Manufacturer Narrative: --/-- The investigation conducted by the isi field service engineer (fse) was unable to replicate the vision issue experienced by the site. Functional testing of the site's da vinci surgical system found that the system functioned with in specifications. On (b)(4) 2013, isi received additional information from the isi clinical sales representative (csr) who initially reported this event. The csr indicated that the site did not contact isi for technical support assistance to resolve the vision issue. The csr indicated that the or nurse indicated to her that the illuminator failed intraoperatively and that when the endoscope was outside of the patient, the flow of light through the optic was observed. Per the csr, the or nurse indicated that after introducing the endoscope through the trocar the inside of the patient's body appeared dark. The csr indicated that the or nurse indicated to her that a restart of the system was performed; however, the issue recurred and the surgeon made the decision to convert the planned surgical procedure to open. The csr indicated that the next day after the customer's report, she went out to the site and tested the camera head and optic. The csr indicated that both the camera and optic functioned well. The csr indicated that the or nurse reported that the patient is doing fine. The csr also indicated that she has advised the or nurse they should contact isi technical support engineering immediately for troubleshooting assistance to prevent future conversions of da vinci surgical procedures. Isi had made multiple attempts to obtain additional information from the site concerning the reported event, however, no additional information has been provided by the site. If additional information is provided, a follow-up mdr will be submitted. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. --/-- Additional Manufacturer Narrative: --/-- On (b)(4) 2013, intuitive surgical inc. (isi) received additional information regarding this complaint. The following information was provided by the surgeon: the reported da vinci si surgical procedure was performed on (b)(6) 2013. The reported vision issue with the da vinci si system occurred approximately an hour after the surgical procedure had begun. Due to the alleged vision issue, the surgeon made the decision to finish the surgical procedure laparoscopically using the same incision ports and the patient did not sustain any intra-operative complications. According to the surgeon, the da vinci si cardiac-thoracic procedure was not converted to open surgical techniques as reported in the initial report of this complaint. The surgeon indicated that the patient was doing fine and had not experienced any post-operative complications. The surgeon did not know what caused or contributed to the alleged vision issue of the da vinci si system. The surgeon also denied that there were any issues during the 3-d calibration of the endoscope. An isi field service engineer (fse) evaluated the da vinci si system using the same endoscope that was involved with the alleged issue and performed a system verification that was found to be within specification. The fse indicated that the endoscope was not defective and will not be returned to isi for investigation. The surgeon indicated that the patient was (b)(6). Based on the follow-up information, this complaint is being reclassified as a non-reportable due to the following conclusion: there is no evidence that the patient experienced a serious injury because of the reported event. In addition, a recurrence of the alleged vision issue with the hrsv is not expected to be likely to cause or contribute to a death, serious injury or serious deterioration in the state of health of a patient. | DA VINCI SI SURGICAL SYSTEM | Reuse | 2997973 | 2972455 | 2868988 | 2955842201300587 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | InitialFollowup | 2013-01-25 | 2013-02-21 | Yes | Yes | Health Professional | IS3000 A60P3 | Yes | No | 2013-01-25 | Device Not Retu | No |  | 2013
42017207 | 2013-01-29 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure that the maryland bipolar forceps instrument was noted to have frayed wires at the distal end. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the grip cable was frayed at the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at wrist were not damaged. The instrument passed electrical continuity testing. Engineering found that the frayed grip cable was derailed at the distal idler pulley. The grips could still open and close, but movement was not precise. Engineering concluded the cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. The fleet angle between proximal and distal pulleys may have contributed to the derailment. In addition, engineering found the one grip was bent, causing side to side misalignment of grips. There was a. 024 offset at the tips indicating overloading at the tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2998030 | 2972501 | 2869017 | 2955842201300580 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-31 | 2013-02-21 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-31 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-15 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, during setup the jaws of the fenestrated bipolar forceps instrument were noticed to not be aligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent, causing side to side misalignment of the grips. There was a 0. 0512 offset at the tips. The grips did not meet together once the grips were closed. Evidence not conclusive, but damage likely due to overloading at the tip or excessive force applied normal to the grip. Instrument passed electrical continuity test. Additional observation not reported by the site was frayed pitch cable. The instrument exhibited frayed segments in various lengths. No damage was found at the clevis. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2998095 | 2972590 | 2869086 | 2955842201300588 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-21 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-01-18 | Malfunction |  | Event Description: --/-- It was reported that during da vinci si prostatectomy procedure the monopolar curved scissors instrument was noted to be dull and unable to cut tissue. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering found various scratch marks with light material removal at the distal end of the main tube. The scratches were. 150 -. 175 in length and were not aligned with the tube axis. The electrical continuity testing passed. No other damage was found. Engineering concluded that damage was likely due to mishandling. Per the customer reported complaint engineering evaluation found that when the instrument was placed on is3000 system for latex cut test, the scissors do not cut cleanly through. 006 latex. The latex got snagged at the scissor tips. The blade edges were not damaged, but edges exhibited wear at the tips, negatively affecting cut performance. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | HOT SHEARS MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 2998152 | 2972644 | 2869160 | 2955842201300584 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2013-02-01 | 2013-02-21 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-27 | Device Not Retu | No |  | 2013
42020505 | 2013-01-10 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the fenestrated bipolar instrument was not responding to the surgeon's control. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument's intuitive motion was experienced due to a broken pitch cable found at the distal clevis hub. The clevis did not exhibit any damage or wear marks. The cable segment that contains the crimp was still installed in clevis. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2998171 | 2972659 | 2841146 | 2955842201300589 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-21 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-09 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a bent at the tip. According to the information provided by the customer, the prograsp forceps instrument was not used on the patient and it was never inserted in the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that instrument was found broken at the proximal clevis to the main tube interface. The clevis was found to be dislodged from the main tube as a result. No missing piece was observed. Engineering concluded that the breakage was likely due to overloading at the tip. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2998212 | 2972698 | 2869194 | 2955842201300590 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-22 | 2013-02-21 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-01-10 | Malfunction |  | It was reported that during a davinci s surgical procedure, the customer noted a broken tip on the permanent cautery hook instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with the ceramic sleeve broken, a small piece was removed from the hook base. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2998411 | 2972803 | 2869319 | 2955842201300591 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-21 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2013-01-21 | Other | Other | It was reported that during a da vinci si upper right lobectomy procedure, a patient side manipulator (psm) arm on the patient side cart stopped working. No patient harm was reported. | The investigation conducted by the isi field service engineer (fse) was unable to replicate the issue experienced by the customer. During the fse's site visit, he spoke to the initial reporter of this event. The initial reporter to the fse that the surgeon experienced issues controlling the patient side manipulator (psm) 1 and that when the issue occurred, the psm was being used at the end of its range of motion. The da vinci coordinator indicated that the surgical staff did not reseat or replace the sterile adapter to resolve the issue, as the surgeon made the decision not to troubleshoot the issue and converted the planned surgical procedure to open. As a precaution, the fse replaced psm 1 and the remote arm controller (rac). A sine cycle performed and a test drive of the system found that the site's da vinci si surgical system functioned within specification. The psm is an instrument arm located on the patient side cart that provides the sterile interface for the endowrist instruments. The rac consists of five printed circuit assembly (pca) boards which operate together to provide control of the system arms. The psm and rac were returned to isi for failure analysis investigation. Engineering was unable to replicate the customer reported failure mode with the returned psm and rac. The psm installed on an in-house is3000 system found that the psm functioned within specification. The rac passed functional testing on an in-house pca test board system and was also exercised for one day and no system errors were generated. On (b)(4) 2013, isi contacted the initial reporter of this event. The initial reporter indicated that the surgeon was unable to control psm 1 after docking the patient side cart to the patient and that the procedure was completed using minithoracotomy techniques. The initial reporter indicated that there was no harm to the patient, the patient was discharged from the hospital on (b)(6) 2013 and that to the best of her knowledge, the patient has not returned to the hospital due to any post surgical complications as a result of the reported event. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2998783 | 2973152 | 2869668 | 2955842201300592 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-22 | 2013-02-21 | Yes | Yes | Health Professional | IS3000 A60P7 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-01-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure when the physician's assistant was placing the instruments in the robotic arms, the fenestrated bipolar forceps instrument was noted to be bent. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found various scratch marks that showed light material removal on the main tube at the distal end. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage to the main tube was likely due to mishandling. Per the customer reported complaint engineering evaluation found that the one grip was bent, causing side to side misalignment of grips. There was a. 161 offset at the tips, indicating overloading at the tip. The electrical continuity testing passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2998863 | 2973205 | 2869721 | 2955842201300593 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-02-06 | 2013-02-21 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-02-06 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the wire of the pk dissecting forceps instrument was broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still remained in the clevis. Clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2998978 | 2973286 | 2869802 | 2955842201300594 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-21 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-01-18 | Malfunction |  | It was reported that during a da vinci si surgical porocedure, the cables of the tenaculum forceps instrument were broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch cable was broken at the wrist. Cable segment that contains the crimp was still installed in the clevis. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS | ENDOSCOPIC INSTRUMENT | 2998981 | 2973300 | 2869816 | 2955842201300595 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-24 | 2013-02-21 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-22 | Malfunction |  | It was reported that during central processing, a broken cable on the fenestrated bipolar forceps instrument was noted. No patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2998986 | 2973318 | 2854184 | 2955842201300596 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-24 | 2013-02-21 | No | Yes | Health Professional | 42020505 | No | No | 2013-01-24 | No | No | Unkown | 2013
42020505 | 2013-01-22 | Malfunction |  | It was reported that during central processing, a broken cable on the fenestrated bipolar forceps instrument was noted. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch up cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. The clevis does not exhibit any damage or wear marks. Other cables at wrist are not damaged. Additional observation not reported by site is tube damage. Distal end of main tube has various scratch marks with light material removal. Scratches are. 085 -. 180 in length and not aligned with the tube axis. Evidence not conclusive, but damage may have been caused by mishandling/misuse. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2999002 | 2973336 | 2854184 | 2955842201300597 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-24 | 2013-02-21 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-25 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the maryland bipolar forceps instrument cable was frayed at the distal end. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | Reuse | 2999012 | 2973337 | 2869853 | 2955842201300599 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-01-25 | 2013-02-21 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No |  | 2013
42017207 | 2013-01-08 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted loose wires on the maryland bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. Engineering also observed residue located near the clamping pulley inside the housing. Upon opening the housing, residue and corrosion were found on the clamping pulley. Engineering concluded that damages found were likely due to mishandling and improper cleaning. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2995438 | 2969807 | 2866872 | 2955842201300550 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-21 | 2013-02-20 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-08 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted a loose wire on the maryland bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that drive cables were found loose at the wrist. The cables were not broken at the distal end, indicating a likely failure at the backend. Upon opening of the housing, the pitch cable and the grip cable were found broken at the back idler pulleys. The idler pulleys were able to spin freely. An orange discoloration at the back idlers and idler block was also found. Additional observations not reported by site were tube damage and corroded clamping pulleys. The main tube exhibited wearing of epoxy layer with light material removal and a rough surface finish throughout its length. The clamping pulleys exhibited excessive white residue around the cable grooves and a small darkened spots indicating possible signs of pitting corrosion. Engineering concluded that damages found were likely due to improper cleaning. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2996112 | 2970358 | 2866872 | 2955842201300551 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-21 | 2013-02-20 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 |  | Injury | Other | During a user interview, it was reported to isi that a nurse was injured as she tried to pull the patient cart of the da vinci surgical system backwards (date of event was unknown; however, it was stated that this has occurred a while ago). Reportedly, the nurse may have grabbed the handles too hard and the patient cart went back too far. Allegedly, the nurse's hand got stuck between the wall and the da vinci system, and she was unable to let go. According to the reporter of this complaint, the nurse was out for 6-8 weeks, but nothing was broken. | Intuitive surgical contacted the hospital to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report as the safety department was not able to identify this case due to the nurse's/patient's name and date of event being unknown. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2996177 | 2970408 | 2866924 | 2955842201300553 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-01-21 | 2013-02-20 | Yes | Yes | Health Professional | IS3000 A60P6 | No | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-22 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted frayed wires at the distal of the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found with frayed wires at the distal clevis hub. Frayed segment is approximately 0. 015 in length. No other cable damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2996197 | 2970418 | 2866932 | 2955842201300552 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-20 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-21 | Injury | Other | It was reported that during a da vinci si hysterectomy procedure, the fenestrated bipolar forceps instrument activated without the surgeon activating cautery energy from the foot pedal on the surgeon side cart. No patient harm was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. The instrument passed recognition and engagement testing when tested on an isi in-house is3000 system. The instrument moved intuitively with a full range of motion in all directions, the grips opened and closed properly and performed cautery function with no trouble found. On (b)(6) 2013, isi contacted the isi clinical sales representative that was present during the surgical procedure. The csr indicated that approximately 30 - 40 minutes into the surgical procedure, while the surgeon was grasping tissue that was being dissected for removal, the surgeon observed energy coming from the fenestrated bipolar forceps instrument without any activation of cautery energy from the foot pedal on the surgeon side cart (ssc). The csr indicated that the site performed troubleshooting of the issue and when the surgeon attempted to grasp tissue a 2nd time, the issue recurred. The csr indicated energy activated a 3rd time from the instrument when the surgeon demonstrated to him what had occurred during troubleshooting. The csr indicated that the fenestrated bipolar forceps instrument was immediately replaced and the issue was resolved. The csr indicated that a valley lab force triad electrical surgical unit (esu) and a valley lab disposable esu cable were used in conjunction with the fenestrated bipolar forceps instrument. The csr indicated that it is unknown if the site inspected the esu prior to use; however, the site did inspect the esu during troubleshooting and found that the esu cable was installed properly. The csr indicated that the settings on the ssc were correct and that the surgical procedure was not recorded. The csr indicated that the surgical procedure was completed and there was no impact to the patient, except to the tissue being removed as part of the surgical procedure. The csr also indicated that there was no report by the surgeon that the patient experienced any post-surgical complications as a result of the reported event. On (b)(6) 2013, the site's da vinci si surgical system was evaluated by an isi field service engineer. The fse was unable to replicate the failure mode experienced by the site as inspection of the site's surgical system found that it functioned within specifications. As of (b)(6) 2013, there have been no reported recurrences of the issue at this hospital. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2996187 | 2970427 | 2866941 | 2955842201300555 | 1 | NAY | Manufacturer | Company Representative | Nurse | Initial | 2013-01-21 | 2013-02-20 | Yes | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-21 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, a broken cable on the fenestrated bipolar forceps instrument was noted. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch down cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in clevis. The clevis does not exhibit any damage or wear marks. Other cables at the wrist are not damaged. Additional observation not reported by site is tube damage. Distal end of main tube has various scratch marks with light material removal. Scratches are. 115 -. 150 in length and not aligned with the tube axis. Evidence not conclusive, but damage may have been caused by mishandling/misuse. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the failure analysis findings; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2996204 | 2970433 | 2866949 | 2955842201300556 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-21 | 2013-02-20 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42805307 | 2013-01-08 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure, the medium-large hem-o-lok clip applier instrument would not clip down, causing the clip to fall out of the instrument and into the patient. The clip was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was able to hold a clip but the grip was not able to close properly and fully to engage the clip. The input disc was found to experience high friction during rotation and the grip jaws would not fully open and close when put tested on an isi in-house is3000 system. The jaws were able to be manually closed to ensure proper clip closure. On (b)(4) 2013, isi contacted the initial reporter of this complaint and obtained additional information. The initial reporter indicated that the instrument involved with this complaint was inspected prior to use. The initial reporter indicated that there was no video recording of the planned surgical procedure and did not know how long the instrument was in use prior to the event. The initial reporter indicated that the event occurred when the surgeon was attempting to clip the patient's cystic duct. She did not know what other instruments or tools were in use during the procedure. The clip that fell into the patient was retrieved using a laparoscopic grasper instrument. The initial reporter denied that any post-operative tests were performed or that any post-operative complications occurred. | MEDIUMLARGE CLIP APPLIERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2996205 | 2970453 | 2866947 | 2955842201300557 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-20 | No | Yes | Health Professional | 42805307 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-21 | Malfunction |  | It was reported that during central processing, fibers were sticking out at end of a pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that conductor wires are intact and undamaged, but drive cable is frayed. The pitch down cable is frayed at the distal clevis hub. Frayed strands stick out at the wrist. Other cables at wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2996213 | 2970455 | 2866969 | 2955842201300558 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-21 | 2013-02-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-31 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the fenestrated bipolar forceps instrument had a damaged wire at the wrist. The instrument was replaced and the procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that a wire was damaged at the wrist. The instrument was found with frayed cables along the proximal and distal area. Four frayed cable strands were sticking out. No damage was found on pulley and clevis area. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2996256 | 2970495 | 2867011 | 2955842201300559 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-22 | 2013-02-20 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-22 | Malfunction |  | It was reported that prior to starting a da vinci si hysterectomy procedure, frayed cables were noted on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2996266 | 2970510 | 2867026 | 2955842201300560 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-20 | No | Yes | Health Professional | 42009308 | No | No | 2013-01-22 | Device Not Retu | No | Unkown | 2013
42020706 | 2013-01-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the tip of the instrument bent. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the pitch cable frayed and loose at the distal clevis hub, and the cables were found derailed at the clamping pulley. Engineering evaluation also found that the cables and idler block pulley at the back end of the housing exhibited chemical attack and had rust. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2996330 | 2970543 | 2867059 | 2955842201300562 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-20 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-02 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable of the pk dissecting forceps instrument snapped. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch down cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Other cables at the wrist are not damaged. Additional observation not reported by site is tube damage. Distal end of main tube has various scratch marks with light material removal. Scratches are. 075 -. 115 in length and not aligned with the tube axis. Evidence not conclusive, but damage may have been caused by mishandling/misuse. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2996333 | 2970545 | 2867061 | 2955842201300563 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-01-02 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable of the large suturecut needle driver instrument snapped. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken near the proximal pulleys and prox clevis cable hole. The idler pulley spins freely and does not exhibit damage. Cable segment sticks out at wrist. Additional observation not reported by site is proximal clevis wearing. Proximal clevis cable hole exhibits wear on one edge. Other cables at the wrist are not damaged. Evidence not conclusive, but wear on hole may have contributed to cable breakage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2996304 | 2970553 | 2867047 | 2955842201300561 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-20 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
40018012 | 2013-01-03 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure, the surgical staff alleged that the monopolar curved scissors (mcs) tip cover accessory, installed on the mcs instrument, burned. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. On (b)(6) 2013, isi contacted the nursing clinical leader for robotic surgery at the site and obtained additional information. She indicated that the accessory, instrument, and cannula were inspected prior to use. To her knowledge, she also indicated that the mcs tip cover accessory was installed using the isi installation tool. She denied that electrolube was applied to the mcs instrument prior to installing the mcs tip cover accessory. She denied that electrolube was applied to the mcs instrument prior to installing the mcs tip cover accessory. She denied that the event was recorded, that the patient was injured as a result of the alleged arcing issue, that any post-operative complications occurred, or that any instruments were removed at any time prior to the arcing event. She also mentioned that the site uses valleylab force fx electrosurgical unit (esu). According to the nursing clinical leader, the surgeon involved with this event typically use 45 (coag), 45 (cut), and 25 (bipolar) settings. She did not know how long the instrument and accessory were in use before the arcing event occurred. On (b(6) 2013, isi also contacted the csr for the site, who was present during the event, and obtained additional information. He verified that the mcs tip cover accessory was melted on the side. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2996335 | 2970560 | 2867077 | 2955842201300554 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-24 | 2013-02-20 | No | Yes | Health Professional | 40018012 | No | No | 2013-01-24 | No | Yes | Unkown | 2013
| 2013-01-03 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure, the surgical staff alleged that the monopolar curved scissors (mcs) tip cover accessory, installed on the mcs instrument, burned. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. On (b)(4) 2013, isi contacted the nursing clinical leader for robotic surgery at the site and obtained additional information. She indicated that the accessory, instrument, and cannula were inspected prior to use. To her knowledge, she also indicated that the mcs tip cover accessory was installed using the isi installation tool. She denied that electrolube was applied to the mcs instrument prior to installing the mcs tip cover accessory. She denied that electrolube was applied to the mcs instrument prior to installing the mcs tip cover accessory. She denied that the event was recorded, that the patient was injured as a result of the alleged arcing issue, that any post-operative complications occurred, or that any instruments were removed at any time prior to the arcing event. She also mentioned that the site uses valleylab force fx electrosurgical unit (esu). According to the nursing clinical leader, the surgeon involved with this event typically use 45 (coag), 45 (cut), and 25 (bipolar) settings. She did not know how long the instrument and accessory were in use before the arcing event occurred. On (b)(4) 2013, isi also contacted the csr for the site, who was present during the event, and obtained additional information. He verified that the mcs tip cover accessory was melted on the side. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2996336 | 2970563 | 2867077 | 2955842201300564 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-24 | 2013-02-20 | No | Yes | Health Professional |  | No | No | 2013-01-24 | No | Yes | Unkown | 2013
42022703 | 2012-12-20 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the pk dissecting forceps instrument would not open and/or close. After the instrument was pulled out, it was discovered that a cable was broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken cable. The pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp has still remained in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2996351 | 2970574 | 2867108 | 2955842201300566 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-01-17 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgeon was unable to articulate the mega needle driver instrument. Upon removal of the mega needle driver instrument from the patient, it was noted that the wire on the instrument was looped. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the pitch cable broken at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2996431 | 2970584 | 2867118 | 2955842201300569 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-20 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-21 | Malfunction |  | Event Description: --/-- It was reported that during an atrial septal defect repair procedure, the large needle driver instrument was found to have a broken screw after it had been taken out. Reportedly, this issue delayed the procedure about 10 minutes. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering confirmed that the screw is broken. The clevis pin and the pulley were broken off but were returned with the instrument. Evidence not conclusive but damage most likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event. It was stated that nothing was observed falling into the patient. There was no patient harm or adverse outcome reported. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering confirmed that the screw is broken. The clevis pin and the pulley were broken off but were returned with the instrument. Evidence not conclusive but damage most likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. | LARGE NEEDLE DRIVER | Reuse | 2996328 | 2970592 | 2867088 | 2955842201300565 | 1 | NAY | Manufacturer | ForeignUser facility | Other | FollowupInitial | 2013-01-22 | 2013-02-20 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No |  | 2013
42004906 | 2013-01-21 | Malfunction |  | It was reported that during a da vinci si surgical procedure, prior to installing the cadiere forceps instrument into the patient, the wire on the wrist of the instrument was observed to be broken. No missing or fallen pieces were reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2996366 | 2970610 | 2867106 | 2955842201300567 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-21 | 2013-02-20 | No | Yes | Health Professional | 42004906 | No | No | 2013-01-23 | No | No | Unkown | 2013
42022703 | 2013-01-22 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted frayed cables on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable located at the distal idler. No damage was found on pulley and clevis. Instrument passed electrical continuity test. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal and rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2996454 | 2970679 | 2867195 | 2955842201300568 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-17 | Malfunction |  | It was reported that after successful completion of a da vinci si hysterectomy procedure, the maryland bipolar forceps instrument became stuck inside of the port during removal of the instrument from the patient. After the instrument was removed, inspection of the instrument by the or nurse found that one wire at the distal tip of the instrument was broken. No missing or fallen pieces were reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp, had one crimp missing. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering evaluation also found that the distal end of the main tube has various scratch marks that had light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage was likely due to misuse/mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2996531 | 2970755 | 2867271 | 2955842201300570 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-24 | 2013-02-20 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-20 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the cable on the megasuturecut needle driver instrument broke and the grips would not open and close. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one cable was broken around the grove near grip. The broken cable segment stuck out at the wrist, however, there were no signs of damage where the cable broke off. There were scuff marks on the idler pulley where the cable was resting on, but the pulley moved freely. The blades and grips did not exhibit damage. Engineering evaluation also found that the other grip cable was derailed at the opposite distal idler pulley of the broken cable. The cable seemed to be tensioned as the grips still moved side to side. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2996560 | 2970785 | 2867299 | 2955842201300572 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-01-24 | 2013-02-20 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-18 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the cables at the tip of the prograsp forceps instrument were observed to be broken and loose. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch down cable was broken at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2996571 | 2970795 | 2867311 | 2955842201300571 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-24 | 2013-02-20 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tips on the pk dissecting forceps instrument were observed to be misaligned. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was bent, causing side to side misalignment of grips. There is a. 055 offset at the tips, indicating that overloading at the tip occurred. The instrument passed electrical continuity testing. Engineering evaluation also found that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit any damage or wear marks. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2996561 | 2970801 | 2854289 | 2955842201300573 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-25 | 2013-02-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tips on the large needle driver instrument were misaligned. No missing or fallen pieces were reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode; however, engineering found that one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment sticks out at the wrist. The other cable at the wrist were not damaged. Engineering evaluation also found that the distal pulley exhibited an indentation on the edge of the distal pulley and had some small scratches on the surface. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2996596 | 2970810 | 2845608 | 2955842201300575 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-25 | 2013-02-20 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tips on the pk dissecting forceps instrument were observed to be misaligned. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The grips had a very small offset at the tips measuring roughly. 005. This offset is below the maximum allowable specification of. 010, therefore the instrument was acceptable and ok to use. Engineering evaluation also found that the pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at wrist were not damaged. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 070 -. 215 in length and were not aligned with the tube axis. Engineering concluded that the damage to the tube was likely due to mishandling/misuse. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2996591 | 2970814 | 2854289 | 2955842201300574 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-25 | 2013-02-20 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-21 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted frayed wires on the mega suture cut needle driver instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional finding not reported by site was an indentation on edge of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2993433 | 2968190 | 2864706 | 2955842201300539 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-21 | 2013-02-19 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-12 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported seeing a cut cable with a maryland bipolar forceps instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of a cut cable. The instrument was found to have a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. Electrical continuity testing of the instrument passed. The other cables of the instrument's wrist did not exhibit any damage. Engineering evaluation also found scratches on the surface of the distal pulley. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | Reuse | 2993525 | 2968290 | 2864806 | 2955842201300540 | 1 | NAY | Manufacturer | ForeignUser facility | Other | FollowupInitial | 2013-03-21 | 2013-02-19 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-03-21 | Device Not Retu | No |  | 2013
42022703 | 2013-01-21 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, it was noted that the pk dissecting forceps had a frayed cable at the distal end. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the frayed cable, however, for clarification the damaged instrument component was a damaged insulation conductor wire. Based on this clarification, the customer reported complaint was confirmed. Engineering also found a deep scratch mark and black insulation was lifted up exposing the cable underneath. Engineering concluded that damage was likely due to mishandling. Additional observation not reported by site was a bent tip. Engineering found one grip was bent, causing side to side misalignment. There was a. 046 offset at the tips. The bent grip exhibited a separation at the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. The main tube had various scratch marks at the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, if the reported malfunction were to reoccur, it could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2994086 | 2968678 | 2865193 | 2955842201300542 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-21 | 2013-02-19 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-28 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported seeing frayed wires on the fenestrated bipolar forceps instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of a frayed wire. Engineering evaluation found a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. Engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering evaluation also found a bent bipolar pin. The bottom pin was observed to be bent to the side. Engineering evaluation concluded that the scratches and bent pin might have been due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable and main tube scratch issues if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2994130 | 2968682 | 2865197 | 2955842201300543 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-28 | 2013-02-19 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-28 | No | No | Reuse | 2013
42017207 | 2013-01-14 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that the maryland bipolar forceps instrument was not moving correctly. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint that the instrument would not move correctly. The instrument was found to have a broken pitch cable at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | Reuse | 2994119 | 2968688 | 2865203 | 2955842201300544 | 1 | NAY | Manufacturer | User facility | Other | FollowupInitial | 2013-02-25 | 2013-02-19 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No |  | 2013
42020505 | 2013-01-21 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the wire at end was broken and it was sticking out of the fenestrated bipolar forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Additional observation not reported by site was main tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were. 065 -. 145 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2994173 | 2968726 | 2865241 | 2955842201300541 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-21 | 2013-02-19 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-18 | Malfunction |  | It was reported that prior to starting a davinci s surgical procedure, the davinci system was unable to recognize the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. Instrument was checked for recognition on an in-house davinci system robot is3000, the davinci system robot successfully recognized the instrument, showing 4 uses left. The pogo pins did not stick and were not contaminated. Engineering was unable to replicate recognition issue. Additional observation not reported by site was a broken cable. One grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment sticks out at wrist. Other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2994183 | 2968753 | 2865268 | 2955842201300545 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-21 | 2013-02-19 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
40018012 | 2012-12-21 | Malfunction |  | It was reported that prior to starting a da vinci si hysterectomy with bilateral salpingo-oophorectomy (bso) procedure, the surgical staff reported that the mcs tip cover accessory installed on the monopolar curved scissors instrument melted and a hole formed. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. |  | MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2994196 | 2968767 | 2822380 | 2955842201300546 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-24 | 2013-02-19 | No | Yes | Health Professional | 40018012 | No | No | 2013-01-24 | No | Yes | Unkown | 2013
42020505 | 2013-01-10 | Malfunction |  | It was reported that the fenestrated bipolar forceps instrument could not work. No other information was available. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a corroded board, there were no other signs of chemicals observed. Engineering concluded that damage was likely due to improper cleaning. Additional observation not reported by site was a frayed pitch cable located at the distal clevis hub. The frayed section was approximately 0. 03 in length. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: warning: read all instructions carefully. Failure to properly follow instructions may cause improper functioning of the device. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2994162 | 2968770 | 2865285 | 2955842201300548 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-21 | 2013-02-19 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42018304 | 2013-01-29 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the surgical staff reported seeing wires sticking out of the permanent cautery hook instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint of a wire sticking out. The conductor wire was found to be undamaged at the distal end, but one pitch and one yaw cable were observed to be loose at the wrist. Motion was not intuitive in pitch/yaw as a result. The instrument was disassembled and one pitch and one yaw cable was found to be broken at the distal end of the hypotube. The cables exhibited a fuzzy appearance at the breakage. Engineering evaluation also found tube damage and corroded levers. The instrument's main tube exhibited wearing with light material removal and a rough surface in various locations along its length. Both levers exhibited pitting corrosion and a darkened discoloration. Engineering evaluation concluded that the damage was likely related to exposure of the instrument to improper cleaning chemicals. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. O prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issues if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOKINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2994166 | 2968778 | 2865313 | 2955842201300549 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-29 | 2013-02-19 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-22 | Malfunction |  | It was reported that during reprocessing of the pk dissecting forceps instrument the wire at the tip of the instrument was observed to be exposed. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode with the following clarification that device malfunction was a frayed pitch cable. The frayed pitch cable was found at the wrist. There was no damage found at the clevis. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2994167 | 2968780 | 2865315 | 2955842201300547 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-22 | 2013-02-19 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
M10121210 | 2013-02-19 | Injury | Required Intervention | The pulley system on the intuitive surgical, davinci megasuture cut endowrist broke during a surgical intervention. An identical unit was opened and utilized to finish the procedure. Reason for use: laparoscopic robotic total hysterectomy, bso, cysto, loa. |  | MEGASUTURE CUT ENDOWRIST | MEGASUTURE CUT ENDOWRIST | 3004978 | 2978469 | 2874983 | MW5029078 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-02-19 | 2013-02-19 | No | Yes | Health Professional | M10121210 | Yes | Yes |  |  |  |  | 2013
420227 | 2013-01-02 | Malfunction |  | Cable break noted during pre-op inspection. Item not used. No harm to pt. This item had 10 uses. |  | INTUITIVE SURGICAL | ROBOTIC PK DISSECTING FORCEPS | 3005499 | 2978548 | 2875062 | 2978548 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-01-10 | 2013-02-18 | No | Yes | Health Professional | 420227 | Yes | Yes |  |  |  |  | 2013
420227 | 2013-01-07 | Malfunction |  | Pk dissecting forceps not working properly per surgeon. Instrument replaced. Upon inspection of instrument, there was a frayed segment on the instrument. No harm to pt. No x-ray required, as item was intact. |  | INTUITIVE SURGICAL | ROBOTIC PK DISSECTING FORCEPS | 3005504 | 2978552 | 3029931 | 2978552 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-01-10 | 2013-02-18 | No | Yes | Health Professional | 420227 | Yes | Yes |  |  |  |  | 2013
IS3000 A60P7 |  | Death | Death | A wall street journal article indicated the following details related to patient death: 'a follow-up story referred to a case in which a (b)(6) man died following robotic surgery last summer at a (b)(6) hospital. An attorney for the man's family said the urologist who operated on him had never before performed the procedure he was attempting with the robot, according to the report. '. | Intuitive surgical contacted the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2991431 | 2965965 | 2862480 | 2955842201300537 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-01-17 | 2013-02-16 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2013-01-17 | No | No | Reuse | 2013
IS3000 A60P6 | 2002-10-11 | Death | Death | In a published citron research article titled intuitive surgical: angel with broken wings, or the devil in disguise? on january 17, 2013, it was stated in an embedded link on page 22 that in december 2003 the widow of a (b)(6) teacher filed a lawsuit after a doctor at [hospital] accidently cut his aorta and the vena cava while using the da vinci surgical robot to remove a cancerous kidney. The lawsuit alleged that the hospital allowed doctors inexperienced with the robot to perform the surgery. The claim further charged that the hospital was more interested in using its new device than in ensuring her husband's safety. According to the legal findings in this case, the complaint was filed on (b)(6) 2003 and the only named defendant was the hospital. The allegations stem from the claim that the decedent underwent a total nephrectomy procedure. The complaint also alleged that isi had estimated 18 robotic surgeries were necessary to be proficient with the system and that isi certified the doctor after only three training sessions. The complaint stated that the decedent suffered from a severed aorta and vena cava. In addition, it was alleged that a lap pad was left inside of the decedent that they had to re-open and recover. There was also reference to a missing needle, which was never explained by the hospital. The decedent died two days after the surgery. These allegations were not previously reported to intuitive surgical and have not been substantiated. Any further information received by the company will be provided via a follow up mdr. | Intuitive surgical was not able to obtain additional information concerning the reported event as of the date of this report. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2991410 | 2965986 | 2862502 | 2955842201300536 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-01-17 | 2013-02-16 | Yes | Yes | Health Professional | IS3000 A60P6 | No | No | 2013-01-17 | No | No | Unkown | 2013
|  | Death | Death | In a published citron research article titled intuitive surgical: angel with broken wings, or devil in disguise? on january 17, 2013, it was stated on page 22 in an embedded article link (chicago daily law bulletin - family gets $7. 5 million in death after spleen removal) that a (b)(6) man died one month after a da vinci surgery to remove his spleen in 2007. Allegedly, the decedent underwent the surgery because he suffered from a disorder that prevented his blood from properly clotting. The decedent allegedly suffered two holes in his duodenum, an organ that is 6 to 10 inches from the spleen. Furthermore, according to the article, he became sick and developed sepsis that caused several organs to fail post-surgery. It was also alleged that he suffered brain damage and died in (b)(6) 2007. A medical-malpractice lawsuit was filed against four doctors of the (b)(6) -based hospital, who reportedly denied the allegation that the duodenum holes occurred during the surgery. Intuitive surgical was not named as a defendant in this lawsuit. These allegations were not previously reported to intuitive surgical and have not been substantiated. Any further information received by the company will be provided via a follow up mdr. | Intuitive surgical was not able to obtain additional information concerning the reported event as of the date of this report. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2991438 | 2966014 | 2862529 | 2955842201300538 | 1 | NAY | Manufacturer | Literature | Other | Initial | 2013-01-17 | 2013-02-16 | Yes | No | Service Personnel |  | No | No |  | No | No | Unkown | 2013
410322 | 2013-01-16 | Injury | Required InterventionLife Threatening | It was reported that during a da vinci si splenectomy procedure, the surgeon nicked the patient's splenic vein, causing bleeding from the vessel. The surgeon successfully occluded the patient's splenic vein utilizing the da vinci vessel sealer instrument; however, the surgeon made the decision to repair the vessel using open surgical techniques, as it was his decision that damage to the patient's vessel could not be repaired using the da vinci si surgical system. The surgeon attempted to release the vessel sealer instrument from the patient's vessel using the emergency release tool; however, the surgeon was unable to release the instrument. Since the surgeon was unable to release the instrument using the emergency release tool, he pulled the instrument from the patient's vessel, thus causing additional damage to the patient's vessel. Reportedly, the additional damage to the patient's vessel was a small tear. | On (b)(4) 2013, additional information regarding the event was provided by the isi representative who initially reported this event. The isi representative indicated that he was not present during the surgical procedure and that it is unknown why the surgeon did not release the vessel sealer instrument utilizing the master tool manipulators (mtm) on the surgeon side cart. The isi representative indicated the site reported to him that no malfunction of the da vinci si surgical system, instruments and/or accessories occurred during the surgical procedure that caused the injuries to the patient's vessel. The isi representative indicated that the site reported to him that the system was not in an error state when the surgeon attempted to release the instrument from the patient's vessel at the patient side cart and that the surgeon removed the vessel sealer instrument from the patient's vessel after converting the procedure to open. The isi representative indicated that he advised the site that the surgeon could have released the jaws of the instrument from the patient's vessel using the mtms on surgeon side console since the system was not in an error state or if the system had been in an error, the jaw on the instrument could have been opened by manipulating the thumbwheel located on the housing of the instrument. The csr indicated that the site made the decision not to return the vessel sealer instrument to isi for failure analysis, since no malfunction of the instrument occurred. The master tool manipulator (mtm) refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The thumbwheel is located on the side of the vessel sealer housing and allows the user to manually open and close the instrument jaws for intraoperative cleaning and emergency grip release. The csr indicated that he has re-trained the site on how to properly release and remove the vessel sealer instrument should this type of event recur. The endowrist one vessel sealer system user manual specifically states: general precautions and warnings - in case of system failure while the instrument is grasping tissue, open the instrument jaws by turning the thumbwheel in the direction of the arrow to release the tissue. Once the tissue is released, close the jaws (in this case, it is ok to close the jaws on an exposed blade). Squeeze the release levers and carefully withdraw the instrument. - after removal of the instrument during a system or instrument failure, use caution when opening the jaws using the thumbwheel, to avoid potential injury due to an exposed cutting blade. Based on the information provided, it was determined that the da vinci si surgical system did not malfunction in a way that would cause nor contribute to the patient injuries. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | ENDOWRIST ONE VESSEL SEALERINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2989029 | 2963870 | 2860387 | 2955842201300516 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-01-16 | 2013-02-15 | Yes | No | Health Professional | 410322 | No | No | 2013-01-16 | No | No | Unkown | 2013
42030901 | 2013-01-25 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a wire was observed to be sticking out from the mega suturecut needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed with a large needle driver instrument and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of a wire sticking out. One grip close cable was found to be broken at the distal idlers. The idler pulley was observed to spin freely and was not damaged. A cable segment was found to be sticking out at wrist. Other cables at the instrument's wrist were not damaged. Engineering evaluation also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2989263 | 2964027 | 2852323 | 2955842201300517 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-25 | 2013-02-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Initial | 2013
42022703 | 2013-01-23 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy (bso) procedure, the surgical staff alleged that a broken wire was observed in the wrist of the pk dissecting forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of broken cable. The pitch down cable was found to be frayed at the distal clevis hub. Frayed strands were observed to be sticking out at the instrument's wrist. Other cables at the wrist were not damaged. Engineering evaluation also found main tube damage. The distal end of the instrument's main tube had various scratch marks with light material removal. The scratches were. 095 -. 135 in length and not aligned with the tube axis. Engineering evaluation concluded that the scratches were likely caused by mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable and main tube scratch issues if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2989306 | 2964046 | 2860561 | 2955842201300518 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-25 | 2013-02-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-23 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy (bso) procedure, the surgical staff alleged that a broken wire was observed in the wrist of the pk dissecting forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of a broken cable in the in the instrument's wrist. The pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was found to be still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2989517 | 2964188 | 2860561 | 2955842201300519 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-25 | 2013-02-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
42004904 | 2013-01-11 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff alleged that only one side of the cadiere forceps instrument would open. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation was unable to confirm the alleged complaint of only one side opening up on the instrument. The instrument's drive cables were not damaged at the distal end. The grips were able to open/close fine when the input discs were manually rotated. Engineering evaluation also observed tube damage on the instrument. The distal end of the instrument's main tube had various scratch marks with light material removal. The scratches were. 100 -. 250 in length and not aligned with the tube axis. Engineering concluded that the scratches might have been caused by mishandling/misuse. No other instrument damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported tube damage issue if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2989604 | 2964239 | 2860754 | 2955842201300521 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-25 | 2013-02-15 | No | Yes | Health Professional | 42004904 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
40015801 | 2013-01-18 | Malfunction |  | It was reported that after completion of a da vinci s surgical procedure, the snap-fit paddle blade installed on the snap-fit scalpel instrument fell into the patient and was retrieved. No patient harm, adverse outcome or injury was reported. | The instrument accessory has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the accessory is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2013, isi contacted the initial reporter concerning the reported event. The initial reporter indicated that upon removal of the scalpel instrument, with the snap-fit paddle blade accessory installed, the snap-fit paddle blade accessory made contact with the insertion tool. The initial reporter indicated that the snap-fit paddle blade accessory was not in a neutral position during removal of the instrument, which caused the snap-fit paddle blade accessory to bend and fall into the patient. The initial reporter indicated that the surgical technician informed her that there were no issues noted during installation of the snap-fit paddle blade accessory onto the snap-fit scalpel instrument and that the surgical technician indicated that there were no issues observed during use of the instrument with the accessory installed. The initial reporter indicated that the instrument was inspected prior to use and no visible damage was found. The reporter indicated, however, that the insertion tool exhibited a ding due to the blade rubbing against it during removal. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | SNAPFIT PADDLE BLADEACCESSORY | ENDOSCOPIC INSTRUMENT ACCESSOR | 2989956 | 2964516 | 2861051 | 2955842201300522 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-18 | 2013-02-15 | No | Yes | Health Professional | 40015801 | No | No | 2013-01-18 | No | No | Unkown | 2013
42022703 | 2013-01-23 | Malfunction |  | It was reported that during a da vinci si hysterectomy with bilateral salpingo-oophorectomy (bso) procedure, the surgical staff alleged that a broken wire was observed in the wrist of the pk dissecting forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of a broken cable. The pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2989966 | 2964556 | 2860561 | 2955842201300520 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-25 | 2013-02-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-01-10 | Malfunction |  | It was reported that during a da vinci si pulmonary lobectomy procedure, broken wires were noticed on the thoracic grasper instrument. As a result, there was limited movement and articulation possible. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that drive cables are broken at the snake wrist. The broken segments stick out at the wrist. Wrist will not straighten properly and motion is non-intuitive due to cable breakage. Other cables at wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER | ENDOSCOPIC INSTRUMENT | 2990226 | 2964806 | 2861321 | 2955842201300523 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-16 | 2013-02-15 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
40018012 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff observed arcing from the mcs tip cover accessory installed on the monopolar curved scissors (mcs) instrument. The mcs instrument with the tip cover installed was removed, and inspection of the tip cover by the surgical staff found that there was a tear in the middle area of the tip cover. The planned surgical procedure was completed with a replacement tip cover accessory and no patient harm, adverse outcome or injury was reported. | The tip cover accessory was not returned to isi for evaluation, as the initial reporter indicated that the tip cover was discarded, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the tip cover is returned (post engineering evaluation) or if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2990318 | 2964877 | 2822380 | 2955842201300524 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-12-21 | 2013-02-15 | No | Yes | Health Professional | 40018012 | No | No | 2012-12-21 | No | No | Unkown | 2013
42020505 | 2013-01-16 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, wires on the fenestrated bipolar forceps instrument were noted to be broken and the tips were bent. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch down cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The pitch up cable is frayed at the distal clevis hub. One grip is bent, causing side to side misalignment of grips. There is a. 030 offset at the tips. Bent grip does not exhibit separation of the yaw pulley and pulley cover at the glue joint, however, likely cause of bending is still overloading at the tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2990365 | 2964936 | 2861451 | 2955842201300525 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-16 | 2013-02-15 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-17 | Malfunction |  | It was reported that during central processing, frayed wires were noted on the mega suturecut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken grip cable. One grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. Cable segment sticks out at the wrist. Other cables at the wrist are not damaged. Additional finding not reported by the site are indentations on pulley. There are big and small indentations on the same pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2990347 | 2964940 | 2861455 | 2955842201300526 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-17 | 2013-02-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-11 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the system would not recognize the maryland bipolar instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that instrument was successfully recognized on an in-house is3000 system, showing 1 use left. Pogo pins do not stick and are not contaminated. The dallas chip programmer (dcp) verification of instrument passed. Engineering was unable to replicate recognition issue. Additional observation not reported by site is tube damage. Distal end of the main tube has various scratch marks with light material removal. Scratches are. 085 -. 150 in length and not aligned with the tube axis. Evidence not conclusive, but damage may have been caused by mishandling/misuse. On (b)(6) 2013 a follow up call was made with the customer regarding the failure analysis findings. It was stated that no fragments were observed falling into the patient. No further information was available. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2990376 | 2964957 | 2861474 | 2955842201300527 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-17 | 2013-02-15 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-08 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the system was unable to recognize the maryland bipolar instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that an in-house is3000 system successfully recognized the instrument, showing six uses left. Pogo pins do not stick and are not contaminated. Dallas chip programmer (dcp) verification of instrument passed. Engineering was unable to replicate the recognition issue. Additional observation not reported by the site is a frayed cable. The conductor wires are intact and undamaged, but drive cable is frayed. The pitch up cable is frayed at the distal clevis hub. Frayed strands stick out at the wrist. Other cables at wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2990487 | 2965047 | 2861564 | 2955842201300528 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-17 | 2013-02-15 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the fibers of the tenaculum forceps instrument were broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the grip cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. Other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS | ENDOSCOPIC INSTRUMENT | 2990554 | 2965132 | 2861627 | 2955842201300529 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-17 | 2013-02-15 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-20 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the wires of the fenestrated bipolar forceps instrument were noted to be broken. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that both pitch cables are broken at the distal clevis hub. It is unusual for more than one cable to break. The cable segments that contain the crimp are still installed in the clevis. The clevis does not exhibit any damage or wear marks. Other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2990717 | 2965297 | 2861812 | 2955842201300530 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-17 | 2013-02-15 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42806503 | 2013-01-16 | Malfunction |  | It was reported that after successful completion of a da vinci si single-site cholecystectomy procedure, the surgeon observed that two pieces from the single-site port had broken off and fell onto the fatty omentum of the patient's belly. The fragments were removed through the patient's umbilical port using the da vinci si surgical system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the single-site port was returned with two pieces of the lower rim detached. The lower rim is the part that goes inside of the patient. The detached pieces are roughly 1. 3 to 1. 4 long and each of the detached pieces exhibit teeth marks along one edge. There are also teeth marks along the tear line of the port and below the tear. Engineering concluded that the damage to the port was likely due to excessive clamping force. No other damage was found. The da vinci si single-site user manual specifically states: general precautions and warnings: single-site instruments should be handled and operated by trained personnel. Handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | SINGLESITE PORT | ENDOSCOPIC ACCESSORY | 2990904 | 2965482 | 2861999 | 2955842201300531 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-01-16 | 2013-02-15 | No | Yes | Health Professional | 42806503 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Unkown | 2013
42027804 | 2012-11-01 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the surgical staff reported that a broken cable was found on the grasping retractor instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2991218 | 2965804 | 2862321 | 2955842201300533 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-21 | 2013-02-15 | No | Yes | Health Professional | 42027804 | No | No | 2013-01-21 | No | No | Unkown | 2013
42009308 | 2013-01-14 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted that the wire was ripped on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed grip cable and a frayed pitch cable. The frayed grip cable was found to stick out at the wrist in various lengths. The frayed pitch cable was found at distal clevis hub. No damage was found at the clevis. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2991238 | 2965807 | 2862324 | 2955842201300534 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-17 | 2013-02-15 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-11 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted that the cautery post had come off on the maryland bipolar forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that bipolar pin was found to be dislodged. The top bipolar pin was pushed into the back end. The chassis feature that acts as a hardstop for the pins was broken. Engineering concluded that damage was likely due to mishandling which resulted on breaking the chassis feature. Additional observation not reported by site was char marks. Both yaw pulleys exhibited charring and localized melting at the grip base, between the grips. The charring was a sign of an arcing event. Evidence not conclusive, but charring may be due to a breach in the insulation, carbonized tissue creating a conductive path, or high generator settings. Internal laboratory testing has determined that the arc track failure mode is most likely to occur when a bipolar instrument is energized with no tissue between the grips. Internal arcing in bipolar instruments results in a potential loss of product function, but no transfer of energy to the patient. Intuitive has concluded that the likelihood of injury due to this failure is remote. User are instructed to retain back-up instruments in case that an instrument fails to perform, and are also instructed in proper technique to minimize the likelihood of internal arcing. Additional observation not reported by site was main tube damage. The distal end of main tube exhibited various scratch marks with light material removal. Scratches were. 085 -. 145 in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2991340 | 2965857 | 2862392 | 2955842201300535 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-18 | 2013-02-15 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-18 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2013-01-17 | Other |  | It was reported that during a da vinci si partial nephrectomy procedure, the site experienced system error code #16 and system error code #802. With the assistance of an isi technical support engineer (tse) the site powered down the surgeon side cart (ssc) and performed an emergency power off (epo) of the pt side cart (psc). After powering the system back on, the site confirmed that the system homed and no system errors occurred. At the time of the call the tse was unable to view the site's system log. Approximately 45 minutes later, the site contacted isi for technical support assistance indicating that they experienced system error code #31030. Review of the site's system log by the tse confirmed the error code #31030 experienced by the site. The tse also found that system error code #23034 occurred. The tse concluded that system error code 31030 was caused by a failed system restart attempt and system error code #23024 was associated with the cannula port clutch on a pt side manipulator (psm) arm. With the assistance of the tse, the site performed an epo of the psc multiple times; however, the issue persisted. The pt had been under anesthesia approximately 1. 5 hrs when the surgeon made the decision to abort the planned surgical procedure and rescheduled it for a later date. No pt harm was reported. | The system logs were also reviewed during the investigation conducted by the field service engineer (fse). The fse found that the site also experienced system error # 23. The fse concluded that the system error codes experienced by the site were associated with the set up joint (sujx), the remote arm controller (rac) board, #2 pt side manipulator (psm) arm, the embedded serializer setup joint (essj), and auxiliary power board (apb). The sujx is a non-motorized support arm which holds a psm during surgery. The rac consists of five printed circuit assembly (pca) boards which operate together to provide control of the system arms. The psm is an instrument arm located on the pt side cart that provides the sterile interface for the endowrist instruments. The essj is the printed circuit assembly (pca) inside a system arm that monitors the encoder for each of four endowrist instrument joints and their associated backup potentiometers. The auxiliary power board (apb) controls the battery system in the pt cart. The system was repaired by replacing the affected sujx, rac, psm, essj, sujx and apb. The system alarm (system generated fault code) functioned as designed and there was no injury to the pt. Upon determining this condition, the system records the faults and transitions to a safe state. The affected replaced parts were returned and evaluated. Engineering was unable to replicate the system error codes experienced by the site with the returned rac, psm #2, essj and apb as functional testing found that they functioned within specification. Engineering eval of the sujx was able to replicate the system error code #23 experienced by the site. Eval found that a harness within the sujx had a short, which triggered the system error code #23 and causing the issue experienced by the site. The affected harness was replaced to repair the sujx. On (b)(4), 2013, additional info regarding the event was provided by the initial reporter. The initial reporter indicated that to the best of her knowledge the pt did not experience any complications as a result of the reported event and that the pt was scheduled to undergo the re-scheduled surgical procedure using the da vinci si surgical system on (b)(6), 2013. As of (b)(4), 2013, there have been no reported recurrences of the issue at this hosp. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 3021760 | 2969665 | 2866181 | 2955842201300532 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-17 | 2013-02-15 | Yes | Yes | Health Professional | IS3000 A60P7 | Yes | Yes | 2013-01-17 | Yes | No | Reuse | 2013
ID SH1149 | 2013-02-13 |  | Required InterventionHospitalization | Patient was having a robotic proctosigmoidectomy with take down of colovaginal fistula. During the procedure the surgical pa noticed that the #1 robotic arm was not acting right. She stated that while visually examining the area around the entrance site to see if there was a problem, the robotic arm all of a sudden jerked forward moving further into the patient and then jerked back. It was describes as a stabbing motion. This caused an injury to the aortic artery. The abdomen was immediately opened and vascular surgeon called to the room to assist. After the repair of the vessel the surgery continued. Patient was transferred post op to icu. As documented. After the occurrence i talked with the staff in the room and then called the company spoke with someone with technical service, our service engineer, and our sales representative. Robot sequestered until follow up with everyone is completed. What was the original intended procedure?robotic proctosigmoidectomy with take down of colovaginal fistula. |  | DAVINCI | SYSTEM SURGICAL COMPUTER CONTR | 2999312 | 2973586 | 2904239 | 2973586 | 1 | NAY | User Facility |  |  | Initial | 2013-02-15 | 2013-02-15 | Yes | Yes | Physician | ID SH1149 | Yes | No Answer Provided |  |  |  |  | 2013
420227 | 2012-10-12 | Malfunction |  | During robotic myomectomy procedure physician stated that the pk dissecting forcep tips were not approximating correctly. On further inspection of the pk dissecting forcep, one of the wires on the tip of the pulley system that controls the dissectors tips was broken. The instrument had 5 uses left when it broke. Manufacturer response for pk dissecting forcep, intuitive surgical (per site reporter). Complaint was registered rma and device was returned for their review. What was the original intended procedure?robotic myomectomy. |  | ENDOWRIST PK DISSECTING FORCEP | SYSTEMSURGICALCOMPUTER CONTROL | 3005489 | 2979934 | 2876448 | 2979934 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-02-15 | 2013-02-15 | No | Yes | Physician | 420227 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
420277 | 2012-10-16 | Malfunction |  | During a robotic hysterectomy it was noted that a wire the tip of the instrument at the pulley system was broken. Surgeon did not note any unusual movements or issues handling the instrument. There were 4 uses left on this instrument. Manufacturer response for pk dissecting forcep, endowrist pk dissecting forcep (per site reporter). Complaint was registered rma and device was returned for their review. What was the original intended procedure?robotic assisted hysterectomy. |  | ENDOWRIST PK DISSECTING FORCEP | SYSTEMSURGICALCOMPUTER CONTROL | 3007445 | 2982170 | 2737314 | 2982170 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-02-15 | 2013-02-15 | No | Yes | Physician | 420277 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
420049 | 2013-02-12 | Malfunction |  | 8mm cardiere forceps broke during use. Wire at pulley system frayed and exposed. Manufacturer response for cardiere forceps, intuitive endowrist (per site reporter). Complaint was filed with company rma. Device to be returned to company once complaint is verified. What was the original intended procedure?robotic assisted three field esphagectomy. |  | DA VINCI ENDOWRIST CARDIERE FORCEPS | SYSTEMSURGICALCOMPUTER CONTROL | 3015005 | 2989481 | 2886014 | 2989481 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-02-15 | 2013-02-15 | No | Yes | Physician | 420049 | Yes | No Answer Provided |  |  |  |  | 2013
420309 | 2013-01-21 | Malfunction |  | During operation mega suturecut needle driver wire broke at tip of instrument. Needle driver removed from sterile field. What was the original intended procedure?robotic assisted hysterectomy, bso, sacrocolpopexy and tvt. |  | DA VINCI ENDOWRIST MEGA SUTURECUT NEEDLE DRIVER | SYSTEMSURGICALCOMPUTER CONTROL | 3015023 | 2989501 | 3015880 | 2989501 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-02-15 | 2013-02-15 | No | Yes | Invalid Data | 420309 |  | No Answer Provided |  |  |  |  | 2013
420309 | 2012-11-29 | Malfunction |  | During a robotic assisted hysterectomy it was noted by the surgeon that the wire was frayed at the tip of the instrument at the pulley system. The mega suturecut needle driver was removed from the patient and use. Manufacturer response for mega suturecut needle driver, intuitive endowrist (per site reporter). Company received complaint rma and device was returned to company for their review. What was the original intended procedure?robotic assisted total laparoscopic hysterectomy and sacrocolpopexy. |  | DA VINCI ENDOWRIST MEGA SUTURECUT NEEDLE DRIVER | SYSTEMSURGICALCOMPUTER CONTROL | 3015033 | 2989506 | 2886019 | 2989506 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-02-15 | 2013-02-15 | No | Yes | Physician | 420309 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
420227 | 2012-11-20 | Malfunction |  | During robotic assisted hysterectomy it was noted that the pulley cable at the tip of the pk endowrist dissector was frayed. The dissector was immediately removed from the patient and replaced. Manufacturer response for pk endowrist dissecting forceps, pk endowrist dissecting forceps (per site reporter). Complaint filed rma and device returned to the company for their review. What was the original intended procedure?robotic assisted hysterectomy and lso. |  | DA VINCI ENDOWRIST PK DISSECTING FORCEPS | SYSTEMSURGICALCOMPUTER CONTROL | 3015038 | 2989510 | 2886043 | 2989510 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-02-15 | 2013-02-15 | No | Yes | Invalid Data | 420227 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
420309 | 2012-11-19 | Malfunction |  | During a robotic assisted sacrocolpopexy the metal cord/wire at the tip of the instrument at the pulley system broke. The mega suture cut needle driver was immediately removed from service and replaced. Manufacturer response for mega suture cut needle driver, intuitive surgical endowrist (per site reporter). Complaint was filed rma and device was returned for their review. What was the original intended procedure?robotic assisted sacrocolpopexy. |  | DA VINCI ENDOWRIST MEGA SUTURECUT NEEDLE DRIVER | SYSTEMSURGICALCOMPUTER CONTROL | 3015054 | 2989528 | 2886063 | 2989528 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-02-15 | 2013-02-15 | No | Yes | Physician | 420309 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
42805901 | 2012-11-19 | Malfunction |  | It was reported that during a davinci si cholecystectomy procedure, the customer noted that the external shaft broke on the crocodile grasper single site instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the main tube was found to be bent around the midpoint, pointing towards the right. Evidence not conclusive, but damage could be due to using excessive force inserting the tube in/out of cannula. Additional observation not reported by site was deep scratches with missing fragments in same location of where main tube was found to be bent. Damage at midpoint exhibited deep gouge marks with material peeled off the tube, about 0. 3335 long. Engineering concluded that the damage was linked to the bent tube, scraping against cannula with excessive force. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The single site instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Insertion and removal of instruments use care when inserting single-site instruments into the cannula to avoid catching the tip on the seal or cannula bowl and bending the shaft excessively. | CROCODILE GRASPER SINGLESITEINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2987048 | 2962135 | 2858654 | 2955842201300507 | 1 | NAY | Manufacturer | DistributorHealth Professional | Other | Initial | 2013-01-15 | 2013-02-14 | No | Yes | Health Professional | 42805901 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-07 | Malfunction |  | It was reported that after completion of a da vinci si surgical procedure, the customer noted broken wires on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2987604 | 2962501 | 2853689 | 2955842201300508 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-16 | 2013-02-14 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
42024603 | 2013-01-15 | Malfunction |  | It was reported during a da vinci si mitral valve replacement procedure that the 8mm atrial retractor short right instrument was noted to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | ATRIAL RETRACTOR SHORT RIGHTINSTRUMENT 8MM | ENDOSCOPIC INSTRUMENT | 2987636 | 2962554 | 2859073 | 2955842201300509 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-17 | 2013-02-14 | No | Yes | Health Professional | 42024603 | No | No | 2013-01-17 | No | No | Unkown | 2013
42022703 | 2013-01-10 | Malfunction |  | Event Description: --/-- It was reported that after a da vinci si hysterectomy procedure the pk dissecting forceps instrument was noted to have a frayed cable at the jaw. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found frayed cable by the jaws. The instrument was found with frayed a pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 04 in length. No other cable damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | Reuse | 2987640 | 2962558 | 2859077 | 2955842201300510 | 1 | NAY | Manufacturer | User facility | Other | FollowupInitial | 2013-02-26 | 2013-02-14 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-02-26 | Device Not Retu | No |  | 2013
42009308 | 2013-01-07 | Malfunction |  | It was reported that after competition of a da vinci si surgical procedure, the customer noted broken wires on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. Additional observation not reported was a derailed and frayed grip cable. The frayed segments were in various lengths and stick out at the wrist. The grips were able to open and close, but their movement and functionality could not be precise due to the cable derailment. The idler pulley exhibited rim damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2987653 | 2962559 | 2841028 | 2955842201300511 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-16 | 2013-02-14 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2012-12-07 | Death | Death | It was reported that a patient expired approximately four weeks after an esophagectomy procedure was performed. The robotically-assisted surgery was converted to open surgical techniques, as it was stated to have been difficult for a number of reasons, according to the surgeon. The company representative was informed that the patient was discharged three weeks after the procedure for rehabilitation. The patient then died a week later of what the hospital staff believes to be due to a pulmonary embolism/other cardiorespiratory related event. No further information was available at this time. | Intuitive surgical contacted (via phone and email) the initial reporter in regards to this complaint in attempts to obtain additional information concerning the reported event. It was stated that the procedure lasted a total of 16 hours, and that there were no known defects of any intuitive surgical instrument(s)/accessory(ies) or da vinci si system. No further information was known. In addition, the robotics coordinator and or manager were contacted (via phone and email) for further information; however, no additional information was obtained as of the date of this report. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2987641 | 2962560 | 2859079 | 2955842201300512 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-01-15 | 2013-02-14 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-01-15 | No | No | Reuse | 2013
IS3000 A60P6 | 2013-01-04 | Injury | OtherRequired Intervention | It was reported that after a successful esophagectomy procedure, the patient was doing well for eight days post surgery, according to the surgeon. However, the patient leaked on the evening of the eighth day and had to go to the emergency room. Reportedly, the anastamosis had to be taken apart and cervical esophagectomy had to be performed instead, ending the esophagus as an ostomy at the neck. No further information was provided at this time. | Intuitive surgical contacted (via phone and email) the initial reporter and appropriate hospital staff in regards to this complaint in attempts to obtain additional information concerning the reported event; however, no additional information has been available as of the date of this report. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2987643 | 2962563 | 2859082 | 2955842201300513 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-01-15 | 2013-02-14 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2013-01-15 | No | No | Reuse | 2013
IS3000 A60P4 | 2013-01-22 | Injury | Required Intervention | An article entitled patient says broken da vinci robot caused a perforated colon that led to disturbing gi complications published on (b)(4) 2011 by (b)(4) was provided to intuitive surgical on (b)(4) 2013. The article alleges that during the patient's prostatectomy procedure in january 2007 the system malfunctioned and the surgical procedure was converted to traditional open surgical techniques to complete the planned procedure. The patient was reported to have gone into recovery in good condition, however, on his way home from the hospital he sensed that the prostatectomy had gone wrong. The article alleged that the patient began urinating out of his rectum and noticed that dark matter was coming out of his penis. | On (b)(6) 2013 the physician's office was contacted to request additional information regarding the reported complaint, however, as of the date of this report the site was not willing to provide further information. A review of system logs for the da vinci surgical system used in the procedure did not produce or confirm any information for the given system at the time period indicated. If additional information is received a follow up mdr will be submitted. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2987645 | 2962573 | 2859090 | 2955842201300491 | 1 | NAY | Manufacturer | Other | Health Professional | Initial | 2013-01-22 | 2013-02-14 | Yes | Yes | Health Professional | IS3000 A60P4 | No | Yes | 2013-01-22 | No | No | Reuse | 2013
40030901 | 2013-01-15 | Malfunction |  | Event Description: --/-- It was reported that during a davinci hysterectomy procedure, item broke at the tip of the mega suture cut needle driver instrument (wires frayed). No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment stuck out at wrist. Other cables at wrist were not damaged. No other damage was found. | MEGASUTURECUT NEDDLE DRIVER INSTRUMENT | Unkown | 2987668 | 2962581 | 2859098 | 2955842201300514 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-02-21 | 2013-02-14 | No | Yes | Health Professional | 40030901 | Device Returned To Manufacturer | No | 2013-02-21 | Device Not Retu | No |  | 2013
42027804 | 2013-01-14 | Malfunction |  | Event Description: --/-- It was reported during a da vinci si nissen fundoplication procedure that the grasping retractor instrument would not release the patient tissue, even with the emergency release. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- It was indicated by the isi representative that the surgeon was repairing a hernia when a non-recoverable error occurred on the system. The emergency wrench failed to open the instrument so the arm was stabilized while the system was restarted and then the surgeon was able to open the grasping retractor with the system and release the tissue. The procedure was delayed by about 10 minutes. There was no harm or injury to the patient, the patient recovered as expected. The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering placed the instrument on a system and the instrument was driven. Recognition and engagement passed testing. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. | GRASPING RETRACTOR 8MM INSTRUMENT | Initial | 2987716 | 2962630 | 2859169 | 2955842201300515 | 1 | NAY | Manufacturer | Company Representative | Other | InitialFollowup | 2013-02-22 | 2013-02-14 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-02-22 | Device Not Retu | No |  | 2013
| 2013-01-30 | Malfunction |  | During the robotic hysterectomy after the port was inserted and the gas was instilled into the abdomen, the camera failed to allow 3-d imaging. The staff was unable to fix the issue. The robotic approach was aborted. The surgeon elected to perform an open abdominal hysterectomy. |  | CAMERA CORD | SYSTEM SURGICAL COMPUTER CONTR | 2999324 | 2973634 | 2870150 | 2973634 | 1 | NAY | User Facility |  |  | Initial | 2013-02-14 | 2013-02-14 | No | Yes | Physician |  | Yes | No Answer Provided |  |  |  |  | 2013
42030901 | 2013-01-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the megasuturecut needle driver instrument, it was noted the instrument wires popped at the end. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip close cable was broken at the distal idler. The idler pulley spins freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering also found one pulley had scratches on the surface and the other one had indentations on the edge of the pulley. Additionally, engineering found various scratch marks on the distal end of the maintube showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2984442 | 2959064 | 2855583 | 2955842201300456 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-29 | 2013-02-13 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-01-25 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure that the grasping retractor instrument appeared to have a broken wire. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the grip cable is frayed at the distal idler pulley. The frayed strands stuck out at the wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR 8MMINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2984444 | 2959083 | 2855599 | 2955842201300455 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-28 | 2013-02-13 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-22 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the fenestrated bipolar forceps instrument, the paddles of the instrument were noted to be out of alignment. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument paddles were out of alignment. The one grip was bent, causing side to side misalignment of the grips. There was a. 010" offset at the tips. They were bent but within the acceptable specification. The electrical continuity testing passed. Engineering also found a broken pitch cable at the distal clevis hub. The cable segment that contained the crimp remained installed in clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2984454 | 2959093 | 2855612 | 2955842201300454 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-23 | 2013-02-13 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-24 | Malfunction |  | It was reported that during reprocessing the maryland bipolar forceps instrument, it was noted that the cored was broken at the tip of the instrument. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the conductor wire was found broken at the yaw pulley, with the wire tip frayed. The yaw pulley was also found broken, with a small piece removed. The electrical continuity testing failed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2984471 | 2959115 | 2855631 | 2955842201300453 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-24 | 2013-02-13 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-04 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, it was noted that the pk dissecting forceps instrument's tip were unable to open and close. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. Pk dissecting forceps instrument's tips were able to open and close without any issues. Additional observation not reported by site was frayed pitch cable located at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. No damage was found at the clevis. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2984650 | 2959256 | 2854289 | 2955842201300486 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-13 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-04 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, it was noted that the pk dissecting forceps instrument's tip were unable to open and close. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. Pk dissecting forceps instrument's tips were able to open and close without any issues. Instrument passed recognition and engagement tests. Additional observation not reported by site was frayed pitch cable located at the distal clevis hub. No damage was found at the clevis. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2985099 | 2959660 | 2854289 | 2955842201300487 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-13 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-01-07 | Malfunction |  | It was reported that prior during reprocessing of the bowel grasper instrument, the wire on the bowel grasper instrument was observed to be sticking out. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one drive cable was broken at the snake wrist. The broken cable segment stuck out of the snake wrist and the instrument motion was non-intuitive as a result of the broken cable. The other cables at the wrist were not damaged. 48 - engineering evaluation also found that the black tube insulation was damaged at the distal end. The insulation was gouged in various places and had a 0. 111 x 0. 075 piece of insulation lifted up and a missing piece that measured approximately 0. 109 x 0. 071 in size. The metal shaft was also exposed as a result. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(4) 2013, isi contacted the initial reporter concerning the reported event. The initial reporter indicated that there was no report by the surgical staff that the any piece from the reported affected instrument fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2985197 | 2959752 | 2841144 | 2955842201300488 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-13 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wires on the prograsp forceps instrument appeared to be cut and the tip of the instrument was not grasping properly. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2985243 | 2959780 | 2856279 | 2955842201300489 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wires on the bipolar forceps instrument appeared to be cut. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2985274 | 2959799 | 2856315 | 2955842201300490 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-22 | 2013-02-13 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-02 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure, the cable of the prograsp forceps instrument snapped. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Additional findings not reported by the site are scratches on the main tube. The distal end of the main tube has various scratch marks showing light material removal and a rough surface finish. Scratches are short in length and not axially aligned with the tube. Evidence was not conclusive, but damage may be due to mishandling. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2985315 | 2959821 | 2856340 | 2955842201300493 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-14 | 2013-02-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-28 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the cable of the mega suturecut needle driver instrument snapped. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. Other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2985338 | 2959829 | 2856388 | 2955842201300496 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-14 | 2013-02-13 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-14 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the cable of the mega suturecut needle driver instrument snapped. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken near the proximal pulleys and proximal clevis cable hole. The idler pulley spins freely and does not exhibit damage. The cable segment sticks out at the wrist. Additional observation not reported by site is proximal clevis wearing. The proximal clevis cable hole exhibits wear on one edge. Other cables at the wrist are not damaged. Evidence is not conclusive, but wear on hole may have contributed to cable breakage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2985308 | 2959857 | 2856373 | 2955842201300492 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-14 | 2013-02-13 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-18 | Malfunction |  | It was reported that during reprocessing of the megasuturecut needle driver instrument, it was noted that the wires on the instrument were frayed. No missing or fallen pieces were reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with one grip close cable broken at the distal idlers pulley. The idler pulley spun freely; however, the pulley exhibited damage on the edge and on the surface. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2985309 | 2959858 | 2856374 | 2955842201300494 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-13 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-17 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the prograsp forceps cable snapped. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Additional findings not reported by the site are scratches on the main tube. The distal end of the main tube has various scratch marks showing light material removal and a rough surface finish. Scratches are short in length and not axially aligned with the tube. Evidence was not conclusive, but damage may be due to mishandling. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2985332 | 2959866 | 2856382 | 2955842201300495 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-14 | 2013-02-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-08 | Malfunction |  | It was reported that during a da vinci si cystectomy procedure, the prograsp forceps instrument was not opening and closing properly. Upon inspection, it was noticed that the wires were frayed. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is frayed near the force amp pulley. Frayed strands stick out at the wrist. Grip open cable on opposite side is derailed from the force amp pulley. Wrist will not straighten and grips cannot open/close properly due to derailment. Additional observation not reported by site is tube damage. Distal end of the main tube has various scratch marks with light material removal. Scratches are. 080 -. 130 in length and not aligned with the tube axis. Evidence not conclusive, but damage may have been caused by mishandling/ misuse. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2985345 | 2959876 | 2841124 | 2955842201300498 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-14 | 2013-02-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-15 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, a frayed cable on the prograsp forceps instrument was noted. No patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2985378 | 2959890 | 2856449 | 2955842201300501 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-13 | No | Yes | Health Professional | 42009308 | No | No | 2013-01-15 | No | No | Invalid Data | 2013
42022703 | 2013-01-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the pk dissecting forceps instrument was not rotating. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2985403 | 2959891 | 2854289 | 2955842201300505 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-13 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42032702 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci si cholecystectomy procedure, the clips of the medium-large clip applier instrument would not stay fixed and kept falling out of the clip applier. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering could not confirm the customer reported complaint of instrument clips not staying fixed and falling out. The instrument was placed on system and driven. Recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. Grips opened and closed properly, and hold clips with no trouble found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEDIUMLARGE CLIP APPLIER | ENDOSCOPIC INSTRUMENT | 2985347 | 2959893 | 2856409 | 2955842201300499 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-14 | 2013-02-13 | No | Yes | Health Professional | 42032702 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-19 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the mega suturecut needle driver instrument's piano-type wire broke inside the shaft, making it inoperable. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2985368 | 2959910 | 2826818 | 2955842201300497 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-14 | 2013-02-13 | No | Yes | Health Professional | 42030901 | No | No | 2013-01-14 | No | No | Unkown | 2013
42017207 | 2012-12-13 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the wrist cable of the maryland bipolar forceps instrument broke. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity passed. Additional findings not reported by the site is scratch marks on the main tube. The distal end of the main tube has various scratch marks showing light material removal. Scratches are short in length and not axially aligned with the tube. Evidence was not conclusive, but damage may be due to mishandling. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2985397 | 2959920 | 2856436 | 2955842201300500 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-15 | 2013-02-13 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-11-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable of the pk dissecting forceps snapped. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch cable was broken at the distal clevis hub. Cable segment that contains the crimp was still installed in the clevis. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2985400 | 2959924 | 2775374 | 2955842201300502 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-13 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the jaws of the mega suturecut needle driver instrument would not close. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. Other cables at wrist are not damaged. Additional findings included scratches on both distal pulleys. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2985401 | 2959925 | 2839190 | 2955842201300503 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-13 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-04 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, it was noted that the pk dissecting forceps instrument's tip were unable to open and close. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent, causing side to side misalignment of grips. There was a 0. 027 offset at the tips. Grips did not meet together once the grips were closed. Evidence not conclusive, but damage was likely due to overloading at the tip or excessive force applied normal to the grip. Instrument passed electrical continuity test. Additional observation not reported by the site was a frayed pitch cable located at distal clevis hub. No damage was found at the clevis. The customer reported complaint does not itself constitute a reportable event; however, if the reported malfunction were to reoccur, it could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2985402 | 2959927 | 2854289 | 2955842201300504 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-13 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the jaws of the pk dissecting forceps instrument are not closing. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering could not confirm the customer complaint of jaws not closing. The instrument was placed on an in-house system and driven. Recognition and engagement passed. The grips opened and closed. Additional observation not reported by the site was a frayed pitch cable. The frayed pitch cable was found at the distal clevis hub. No damage was found at the clevis. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2985425 | 2959972 | 2854289 | 2955842201300506 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-13 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-28 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wire on the mega suturecut needle driver instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering evaluation also found that the same broken grip cable was derailed at the distal idler pulley. The cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. The fleet angle between the proximal and distal pulleys likely contributed to derailment. The surface of the distal pulley also exhibited scratches. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2982320 | 2957106 | 2853622 | 2955842201300469 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-28 | 2013-02-12 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-28 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-10-31 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff alleged that a cable was out from a maryland bipolar forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint that the instrument had a cable that was out. The instrument was found to have a frayed pitch cable at distal idler. No damage was found on the pulley and clevis. The instrument passed electrical continuity testing. Engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks exhibiting light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering also found a bent bipolar pin. Engineering concluded that the scratches and bent pin were likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported cable and tube issues if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2982345 | 2957109 | 2853625 | 2955842201300470 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-15 | 2013-02-12 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-15 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a wire broke on a fenestrated bipolar forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering observed the following: the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. No other damage was found with the investigation of the instrument. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2982365 | 2957111 | 2853688 | 2955842201300471 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-15 | 2013-02-12 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-23 | Malfunction |  | It was observed during central processing that the fenestrated bipolar forceps instrument had frayed wires. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of frayed wires. The instrument's conductor wires were found to be intact and undamaged. However, the instrument's drive cable was found to be frayed. The pitch up/down cable was observed to be frayed at the distal clevis hub. Frayed strands were sticking out at the instrument's wrist. Other cables at the wrist were reportedly not damaged. Engineering evaluation also found a bent grip. One grip was found to be bent, causing side to side misalignment of the grips. There was a. 020 offset at the tips. The bent grip did not exhibit separation of the yaw pulley and the pulley cover at the glue joint. Engineering concluded that the likely cause of the bending was overloading at the tip. Electrical continuity testing of the instrument passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported fayed wire issue if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2982267 | 2957122 | 2853641 | 2955842201300464 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-12 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-15 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a wire broke on a fenestrated bipolar forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Engineering observed the following: the pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. Engineering also observed the following: scratches were found on the instrument's main tube. The distal end of the instrument's main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2982370 | 2957149 | 2853688 | 2955842201300472 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-15 | 2013-02-12 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-07 | Malfunction |  | It was reported that after completion of a da vinci si surgical procedure, the surgical staff reported seeing a broken wire with pk dissecting forceps instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of a broken wire. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. Engineering evaluation also found bent tips on the instrument. One grip was found to be bent, causing side to side misalignment of the grips. There was a 0. 069 offset at the tips, indicating overloading at the tip. Another finding was a scratch on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damages might have be due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken wire and tube insulation issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2982371 | 2957150 | 2853689 | 2955842201300473 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-16 | 2013-02-12 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-01-17 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy with hysterectomy procedure, the surgical staff alleged that a cable broke on the tip of the mega needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was able to confirm the alleged complaint of a broken cable. One grip close cable was found to be broken at the distal idlers. The idler pulley was observed to spin freely and was not damaged. A cable segment was reportedly sticking out at the instrument's wrist. Other cables at the wrist were not damaged. Engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the scratches were likely due to mishandling of the instrument. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable and tube insulation issues if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2982384 | 2957156 | 2853672 | 2955842201300475 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-17 | 2013-02-12 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-11-16 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff alleged that the prograsp forceps instrument was not working properly. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer initially reported that the instrument would not work. Upon microscopic and visual inspection, engineering evaluation found the pitch cables to be loose and frayed at the grip assembly. Intuitive motion of the instrument was not smooth and precise. Upon housing removal found, the pitch cables were observed to be loose at the clamping pulley. However, the clamping pulley screw was not found to be loose. The customer reported complaint does not itself constitute a mdr reportable event; however the reported frayed cable issue if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2982387 | 2957158 | 2853674 | 2955842201300476 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-17 | 2013-02-12 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42018304 | 2012-11-06 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the surgical staff alleged that the permanent cautery hook instrument was not recognized by the da vinci s system. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation was unable to confirm the alleged complaint that the system would not recognize the instrument. The instrument was checked for recognition on an in-house is3000 system. The is3000 system successfully recognized the instrument. The pogo pins were not sticking or found to be contaminated. Dcp verification of the instrument passed. The recognition issue could not be replicated with the evaluation of the instrument. Engineering evaluation also found tube damage on the instrument and a cracked prox clevis. The distal end of the instrument's main tube had various scratch marks with light material removal. The scratches were. 075 -. 115 in length and not aligned with the tube axis. The prox clevis was found to be cracked on one side. The crack was in the axial direction, located halfway between the ears. The crack orientation was consistent with overloading the tip with the wrist pitched. Engineering concluded that the tube and clevis damages were most likely caused by mishandling/misuse. No other instrument damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported tube damage issue if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOKINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2982400 | 2957159 | 2853675 | 2955842201300477 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-17 | 2013-02-12 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-20 | Malfunction |  | It was reported that during a da vinci si nissen fundoplication procedure, the surgical staff alleged that the fenestrated bipolar forceps instrument would not cauterize. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation was unable to confirm the alleged complaint that the instrument would not cauterize. A visual inspection of the instrument showed no damage to conductor wires. Electrical continuity testing, which is an indicator of cautery function, passed. A functional test was performed on the system using a generator and bipolar cord. With the grips holding a wet towel, the cautery pedal was pressed and steam was generated from the towel. The functional test revealed successful cautery function. Engineering evaluation also found a frayed cable and a bent grip. The pitch down cable was frayed at the distal clevis hub. Frayed strands were observed to be sticking out at the instrument's wrist. Other cables at the wrist were not damaged. One grip was found to be bent, causing side to side misalignment of the grips. There was a 0. 025 offset at the tips. The bent grip did not exhibit separation of the yaw pulley and the pulley cover at the glue joint. Engineering concluded that the likely cause of the bent grip was overloading at the tip. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported frayed cable issue if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2982377 | 2957173 | 2817886 | 2955842201300474 | 1 | NAY | Manufacturer |  | Other | Initial | 2013-01-16 | 2013-02-12 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
42018306 | 2012-12-10 | Malfunction |  | It was reported that during a davinci si surgical procedure, the customer noted that a 'cable was out' on the permanent cautery hook instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2982465 | 2957205 | 2853724 | 2955842201300478 | 1 | NAY | Manufacturer | DistributorHealth Professional | Other | Initial | 2013-01-15 | 2013-02-12 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-11-30 | Malfunction |  | It was reported that during a davinci si surgical procedure, the customer noted that the 'cable was out' on the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage found. Additional observation not reported by the site was bent grips. Engineering also found one grip to be bent, causing side to side misalignment of grips. There was a. 07 offset at the tips. The bent grip exhibited separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. Engineering concluded that damage was likely due to mishandling. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2982572 | 2957305 | 2853824 | 2955842201300479 | 1 | NAY | Manufacturer | ForeignHealth Professional | Other | Initial | 2013-01-15 | 2013-02-12 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-22 | Malfunction |  | It was observed during central processing that the fenestrated bipolar forceps instrument had a broken cable. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint of a broken cable. The pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. Engineering evaluation also found tube damage. The distal end of the instrument's main tube had various scratch marks with light material removal. The scratches were. 090 -. 500 in length and not aligned with the tube axis. Engineering concluded that the scratches were likely due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable and main tube scratch issues if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2983190 | 2957668 | 2854184 | 2955842201300481 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-24 | 2013-02-12 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2013-01-14 | Injury | Other | It was reported that during a da vinci si transoral robotic surgery (tors), the patient suffered a lip burn. It was confirmed that the incorrect 3-pronged grounding pad was used. The site was physically removing the third pin in order to connect it to the cannula. This could be a possible failure mechanism as this would constitute an unvalidated use of the product. | The physician's assistants indicated it was believed that the lip burn was caused due to the incorrect grounding pad being used. The hospital staff has since been trained and the correct grounding pad is being used now. The 5mm cannulae were used during the procedure and no instruments were replaced during the procedure. The physician indicated the lip burn was noticed at the end of the case and no treatment other than observation was administered. The planned surgical procedure was completed using the da vinci system and the patient recovered as expected. Investigation has shown that it is possible for the electrosurgical unit (esu) current to pass through the patient side manipulator arms to ground, through the cannula accessory to the patient, if the patient is not properly grounded. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2983195 | 2957752 | 2854271 | 2955842201300480 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-14 | 2013-02-12 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2013-01-14 | No | No | Reuse | 2013
42022703 | 2013-01-24 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported seeing a broken cable with the pk dissecting forceps instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged complaint of a broken cable. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported broken cable issue if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2983162 | 2957764 | 2854283 | 2955842201300482 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-24 | 2013-02-12 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-04 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, while using the pk dissecting forceps instruments, the tips were not closing. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable located at distal clevis hub. There was no damage found at the clevis. Instrument passed electrical continuity test. Additional observation not reported by the site was a loose pitch cable. The pitch cable was found to be loose at distal clevis hub. Both cable crimps remained in the clevis, no damage was found on the cable. Upon housing removal engineering also found pitch cable loose at clamping pulley, but no loose clamping pulley screw was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2983207 | 2957773 | 2854289 | 2955842201300483 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-12 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-25 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a wire was observed to be sticking out from the pk dissecting forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of a wire sticking out. Both pitch cables of the instrument were found to be broken at the distal clevis hub. The cable segments that contain the crimp were still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2983305 | 2957862 | 2854378 | 2955842201300485 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-25 | 2013-02-12 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Initial | 2013
42022703 | 2013-01-24 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported seeing a broken cable with the pk dissecting forceps instrument. The instrument was reportedly not used in the planned surgical procedure. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation was able to confirm the alleged complaint of a broken cable. The instrument's pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still observed to be installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. Engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering evaluation concluded that the scratches might have been due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable and main tube scratch issues if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2983293 | 2957875 | 2854283 | 2955842201300484 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-24 | 2013-02-12 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
420183 | 2012-12-19 | Injury | Required Intervention | On (b)(4) 2013, medwatch uf/importer report (b)(4) was received: the patient is a (b)(6) year old female admitted to (b)(6) on (b)(6) 2012 for a scheduled surgery. The patient had been diagnosed with a retroperitoneal lymphadenopathy. The planned surgery was a robotic assisted peripancreatic lymph node excision. During the procedure to excise the lymph node the surgeon noticed that the cautery hook was malfunctioning. The malfunctioning device was removed. Upon removal of the cautery hook it was noted that part of the plastic casing around the pulley of the hook was broken off. The surgical team immediately began extensive search in the area of the lymph node dissection, underneath the liver and around the porta hepatis, but did not see any evidence of the broken plastic casing. Several attempts were made to locate the missing fragment including creating a hand port in the upper midabdomen. A flat plate kub was taken. The team was unable to locate the missing fragment. The patient remained stable throughout the surgery and at the conclusion of the procedure was taken back to recovery in stable condition. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The initial reporter returned pictures of the device. Which were evaluated by the failure analysis team: based on the pics, the failure mode is a broken distal clevis. One distal clevis ear is broken off at its base. Clevis ear on conductor wire side is broken off, resulting in dislodgment of yaw pulley. Additional observation is a missing conductor cap. Conductor cap likely detached as a result of clevis ear breakage. Evidence not conclusive, but breakage is likely due to overloading the tip. | PERMANENT CAUTERY HOOK | ENDOSCOPIC ELECTROSURGICAL INS | 2983368 | 2957935 | 2809959 | 2955842201300252 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-19 | 2013-02-12 | Yes | Yes | Health Professional | 420183 | No | No | 2012-12-19 | No | No | Unkown | 2013
42000306 | 2012-12-28 | Malfunction |  | It was reported that during a davinci si pulmonary lobectomy procedure, the clips would not stay fixed and kept falling out of the clip applier instrument. Upon inspection, tips noticed to be uneven. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that both grips were found to be broken. Missing pieces were not returned. Grips should had been straight. One grip to hold the clips was longer than the other. Damage was likely due to applying force normal to the grip while trying to install the clip. Evidence not conclusive, but damage was likely a result of misuse/user error. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Section 10 - small clip applier, caution: if excessive motion of the wrist or grips occurs, the clip should be removed and a new clip inserted. This is because if the wrist moves significantly, the grip action may be affected and the clip could become slightly compressed. This could result in the clip falling out of the grip and into the patient during insertion or engagement. If clip is not flush with the tip of the clip applier, the clip should be removed, grips opened and a new clip inserted. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2980163 | 2955062 | 2851578 | 2955842201300452 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-14 | 2013-02-11 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-03 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the prograsp forceps instrument was not recognized by the system. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument successfully passed recognition and engagement testing, showing that the instrument had 1 use remaining while installed on an is3000 in-house system. The pogo pins did not stick and were not contaminated. The instrument passed (b)(4) testing. Additional functional testing revealed that the instrument moved intuitively and the grips opened and closed properly. Engineering evaluation also found that the pitch cable at the distal idler is frayed. There was no damage observed to the pulley and clevis. The customer reported complaint does not itself constitute a mdr reportable event; however; the damage to the instrument's pitch cable found during failure analysis could likely cause or contribute to an adverse event if the malfunctions were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2980208 | 2955126 | 2851662 | 2955842201300457 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-16 | 2013-02-11 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
42018306 | 2013-01-02 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a 'frayed cable' on the permanent cautery hook instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a segment of the conductor wire was found to be dislodged from the conductor cap and it was exposed on the outside of the yaw pulley. The wire insulation was not damaged. The instrument passed electrical continuity test. The conductor cap was properly seated within distal clevis as hook moves in yaw. Additional observation not reported by site was main insulation tube damage. The distal end of the main tube exhibited various scratch marks with light material removal. The scratches were. 095 -. 165 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2980856 | 2955730 | 2852268 | 2955842201300458 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-14 | 2013-02-11 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-01-17 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported seeing a cable sticking out of the mega needle driver instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation results confirmed the issue of a cable sticking out of the instrument. The instrument was found to have one grip close cable broken near the proximal pulleys and proximal clevis cable hole. The idler pulley was able to spin freely and did not exhibit damage. The cable segment was observed to be sticking out at the instrument's wrist. No major wear was found on the proximal clevis cable hole. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | Unkown | 2980909 | 2955790 | 2852328 | 2955842201300461 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-02-25 | 2013-02-11 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-02-25 | Device Not Retu | No |  | 2013
42030901 | 2013-01-16 | Malfunction |  | It was reported that during a da vinci si surgical procedure, it was noted that the cable that controls movement on the megasuture-cut needle driver instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and did not exhibit damage. The cable segment stuck out at the wrist. The other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2980908 | 2955791 | 2852313 | 2955842201300460 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-16 | 2013-02-11 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-16 | Malfunction |  | It was reported that during a da vinci si surgical procedure, it was noted that the cable that controls movement on the megasuturecut needle driver instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was broken at the distal idlers pulley. The idler pulley spun freely; however it exhibited damage on the edge and on the surface. Engineering concluded that the damage was likely due to excess force contact at the surface and mishandling. The cable at the wrist also stuck out. The other cables at the wrist did not exhibit damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2980898 | 2955797 | 2852313 | 2955842201300459 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-16 | 2013-02-11 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
42020906 | 2013-01-22 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the surgical staff reported seeing a wire sticking out at the tip of the mega suturecut needle driver instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. Intuitive surgical contacted the initial reporter of this complaint on (b)(4) 2013, to obtain additional information concerning the reported event. The reporter confirmed that no fragments fell into a patient and no patients were harmed/injured because of the alleged issue. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2980915 | 2955807 | 2852323 | 2955842201300462 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-11 | No | Yes | Health Professional | 42020906 | No | No | 2013-01-22 | No | No | Unkown | 2013
42030901 | 2013-01-21 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported seeing a broken cable with the mega suturecut needle driver instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of a broken cable. One grip close cable was found to be broken at the distal idlers. The idler pulley was observed to spin freely and was not damaged. A cable segment was found to be sticking out at the instrument's wrist. Other cables at the wrist were not damaged. Engineering evaluation also found a derailed grip cable at the distal idler pulley. Engineering concluded that the cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. Engineering also concluded that the fleet angle between the proximal and distal pulleys might have contributed to the derailment. The customer reported complaint does not itself constitute a mdr reportable event; however the reported broken cable issue if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2980930 | 2955812 | 2852330 | 2955842201300463 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-22 | 2013-02-11 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-08 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the cable on the prograsp forceps instrument tore. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode with the following clarification that the instrument was returned with a frayed and derailed grip cable. The frayed cable sections were in various lengths and stuck out at the wrist. The grips can opened and closed; however the movement and functionality may not be precise. Engineering evaluation also found that the distal end of the main tube was missing material. The surface of the main tube appeared to be scraped off. Engineering concluded that the damage was likely due to misuse. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2981047 | 2955935 | 2852453 | 2955842201300466 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-17 | 2013-02-11 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-17 | Malfunction |  | It was reported that during reprocessing of the fenestrated bipolar forceps instrument, it was noted that the wire at the tip of the instrument was exposed. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis does not exhibit any damage or wear marks. The other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2981074 | 2955939 | 2852457 | 2955842201300467 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-17 | 2013-02-11 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-14 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable on the prograsp forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2981077 | 2955943 | 2852461 | 2955842201300468 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-18 | 2013-02-11 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-18 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-15 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the cable at the tip of the dissecting forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode, with the following clarification that the conductor wire was broken. The drive cable were found to be intact; however, the one conductor wire was broken at the side of the yaw pulley. The broken wire segments stuck out at the wrist. There was no evidence of arcing observed. The side of the yaw pulley had a small gouge adjacent to the wire breakage. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches were. 065 -. 220 in length and not aligned with the tube axis. Engineering evaluation concluded damage to the tube and wire was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2981090 | 2955960 | 2852476 | 2955842201300465 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-16 | 2013-02-11 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-31 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the mega suturecut needle driver instrument was noted to have frayed wires. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the one grip close cable was broken at the distal idler. The idler pulley spins freely and was not damaged. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. Engineering evaluation also found that there was an indentation on the edge of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2974605 | 2950299 | 2846835 | 2955842201300433 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-02 | 2013-02-08 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42017207 | 2013-01-02 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that the maryland bipolar forceps instrument would not rotate. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found the instrument to have a loose and frayed pitch cable at the distal clevis hub. The pitch cable was also found derailed at the idler block. Engineering also found chemical residues on the clamping pulley and cable. Engineering concluded that the residue was likely due to improper cleaning of the instrument. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2974678 | 2950328 | 2846870 | 2955842201300432 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-08 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-01-03 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff alleged that the monopolar cautery instrument was not properly responding to surgeon movement. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument to have a broken drive cable at the snake wrist. A broken segment was observed to be sticking out at the snake wrist. The wrist could not properly straighten and motion was not intuitive due to the cable breakage. Engineering evaluation also found the instrument's black tube insulation to be damaged at the distal end. The insulation was gouged on one side and a. 250 x. 090 piece was observed to be missing roughly 3. 25 above the snake wrist. No other instrument damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the damage to the instrument's grip cable and tube insulation found during failure analysis could likely cause or contribute to an adverse event if the malfunctions were to recur. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2974658 | 2950341 | 2846857 | 2955842201300431 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-08 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2013-01-11 | Other | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the harmonic instrument would not fire using the footrest. Per technical service engineer's (tse) instructions, the customer reseated the energy activation cable for the ethicon generator to the personality module energy device (pmed); however, there was no change. The customer did not want to troubleshoot any further at this time. The tse recommended the customer try rebooting the ethicon generator and then test the energy activation. Per site's point-of-contact during a phone follow-up, the generator reset was not performed; no spare cable for the instrument or pmed was available for testing, and only a reseat of pmed cable and instrument was done. Due to continued inability to fire the harmonic instrument, the case was then converted to traditional open surgical techniques. The pmed was later replaced by the field service engineer (fse) on site, and the energy is activating again. It was stated that a backup cable was ordered for the pmed and the instrument. There were no further issues reported. | The electrosurgical generator unit (esu) was returned and evaluated by engineering. Engineering was able to replicate the reported problem by exercising mono/bipolar cautery at all ports. Energy failed intermittently on ports 1 and 2. After checking the eas boards, it was found that eas numbers 1 and 2 were lifted up and did not make good contact as the instrument was firing. Energy activation switch (eas) 1 and 2 were reseated to correct the problem. The da vinci user manual explicitly states that environmental or equipment failures may cause the da vinci system to become unavailable. The surgical team should always have backup equipment and instrumentation available, and be prepared to convert to alternative surgical techniques. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. As of (b)(6) 2013, there have been no reported recurrences of this issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2975402 | 2950825 | 2847347 | 2955842201300446 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-11 | 2013-02-08 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | Yes | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-04 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a broken cord on the pk dissecting forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2975590 | 2951019 | 2847541 | 2955842201300445 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-14 | 2013-02-08 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-17 | Malfunction |  | It was reported that prior to starting a da vinci si sacrocolpopexy procedure, the surgical staff reported seeing broken wires with the mega suturecut needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of broken wires. The instrument's grip cable was found to be frayed at the distal idler pulley. Frayed strands were observed to be sticking out at the wrist. Other cables at wrist were not damaged. Engineering evaluation also found a derailed grip cable at the distal idler pulley. The grips were able to open and close. Engineering concluded that the cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. Engineering evaluation also noted that the fleet angle between the proximal and distal pulleys might have also contributed to the derailment. Engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the main tube damage might have been due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunctions if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2975641 | 2951074 | 2847590 | 2955842201300447 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-18 | 2013-02-08 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-01-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-17 | Malfunction |  | It was reported that during a da vinci si inguinal hernia repair procedure, the surgical staff alleged that a broken wire was observed with the mega suturecut needle driver instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged issue of a broken wire. The instrument was found to have a frayed grip cable at the distal idler pulley. Frayed strands were observed to be sticking out at the instrument's wrist. Other cables at wrist were reportedly not damaged. Engineering evaluation also found the same frayed grip cable to be derailed. The grip cable was derailed at the distal idler pulley. The grips were still able to open and close. Engineering concluded that the cable derailment was likely due to the cable losing contact with pulley during wrist articulation. Engineering also concluded that the fleet angle between the proximal and distal pulleys might have contributed to the derailment. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2975644 | 2951077 | 2847593 | 2955842201300448 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-18 | 2013-02-08 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-18 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-18 | Malfunction |  | It was observed during central processing that the fenestrated bipolar forceps instrument had a broken wire near the wrist. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of a broken wire. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. 48 - engineering evaluation also found a scratch on instrument's main tube. The distal end of the main tube had a scratch mark showing light material removal. The scratch was short in length and not axially aligned with the tube. Engineering concluded that the damage might have been due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunctions if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2975714 | 2951143 | 2847659 | 2955842201300449 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-18 | 2013-02-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-21 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the surgical staff reported seeing a wire sticking out of the mega suturecut needle driver instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of a wire sticking out. The instrument's grip cable was found to be frayed at the distal idler pulley. Frayed strands were observed to be sticking out at the instrument's wrist. Other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2975732 | 2951159 | 2847675 | 2955842201300450 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-21 | 2013-02-08 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42004906 | 2013-01-10 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported seeing a loose wire on the cadiere forceps instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. The reporter denied that the instrument was used on a patient. | The instrument was returned and evaluated. Engineering evaluation confirmed the alleged complaint of a loose wire. One grip close cable was found to be loose at the distal idlers. The idler pulley was observed to spin freely and was not damaged. The clamping pulley screws were found to be secured. The idler block showed no sign of damage. Engineering evaluation also found scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage might have been due to mishandling. Upon opening the instrument's housing, engineering evaluation also found corrosion and residue by the clamping pulley. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issue if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2975752 | 2951164 | 2843120 | 2955842201300451 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-21 | 2013-02-08 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-11-29 | Malfunction |  | It was reported that during a da vinci si surgical procedure the connecting rod of the wrist on the prograsp forceps instrument broke. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2973313 | 2949048 | 2845570 | 2955842201300434 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-01-15 | 2013-02-07 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-11 | Malfunction |  | It was reported that during reprocessing the fenestrated bipolar forceps instrument, it was noted that the wires on the instrument were broken. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2973336 | 2949050 | 2845572 | 2955842201300435 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-07 | No | Yes | Health Professional | 42020505 | No | No | 2013-01-11 | Device Not Retu | No | Unkown | 2013
42022703 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, that the pk dissecting forceps instrument was not functioning properly. Upon inspection of the instrument the site found that the wires at the tip of the instrument were loose. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the wires on the tip were loose. The pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp is missing. The clevis did not exhibit any damage or wear marks. The other cables at the wrist were not damaged. Engineering also found various scratch marks with light material removal at the distal end of the main tube. The scratches were. 100 -. 215 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2973370 | 2949081 | 2830461 | 2955842201300436 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-21 | 2013-02-07 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-08 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, wires were sticking out of the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a frayed grip cable. The frayed cable sections are in various lengths sticking out at the wrist. The idler pulley rim exhibited damage. The instrument has 1 use remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2973375 | 2949092 | 2845608 | 2955842201300437 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-07 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-01-11 | Malfunction |  | It was reported that during central processing, it was noted that wires were coming out of the small clip applier instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip cable is loose, but not visibly broken at the wrist. One grip input disc spins freely without corresponding output motion of grip. This suggests a failure at the back end. The housing was removed to find one grip cable broken near the back idler pulleys. The pulleys spin freely and do not exhibit mechanical damage. Other back end cables are not damaged. Additional observations not reported by site are pitting corrosion and tube damage. Aluminum backend components exhibit light pitting corrosion, but the severity is greater on the roll bushing than the clamping pulleys. The main tube exhibits wearing with light material removal and a rough surface finish throughout its length. Evidence not conclusive, but damage is likely due to improper cleaning. There were 77 fires left. No other damage was found. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The cleaning and sterilization user manual specifically states: do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. Prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. | SMALL CLIP APPLIER | ENDOSCOPIC INSTRUMENT | 2973404 | 2949131 | 2841119 | 2955842201300438 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-07 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42018406 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the wire on the permanent cautery spatula instrument was noted to be broken. No fragments were reported to have fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | PERMANENT CAUTERY SPATULA | ENDOSCOPIC ELECTROSURGICAL INS | 2973440 | 2949169 | 2845691 | 2955842201300439 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-07 | No | Yes | Health Professional | 42018406 | No | No | 2013-01-11 | Device Not Retu | No | Unkown | 2013
42020505 | 2013-01-11 | Malfunction |  | It was reported that prior to a da vinci si surgical procedure, a broken wire was noted on the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken pitch cable. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis does not exhibit any damage or wear marks. The instrument has 4 uses remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2973443 | 2949170 | 2845692 | 2955842201300440 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-07 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the maryland bipolar forceps instrument would only cauterize on the front half of the tip, not the entire tip. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering placed the instrument on an in-house system and it was driven. Recognition and engagement passed. A cauterizing test passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The instrument is fully functional. Electrical continuity passed. The instrument has 3 uses remaining. Additional findings revealed scratches on the main tube. The distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and not aligned with the tube axis, suggesting it may have been caused by instrument collisions or rough handling. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2973448 | 2949175 | 2805527 | 2955842201300441 | 1 | NAY | Manufacturer |  | Other | Initial | 2013-01-11 | 2013-02-07 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42017704 | 2012-12-18 | Malfunction |  | It was reported that during a da vinci si colon resection procedure, the jaws on the 5 mm bowel grasper instrument were not grasping. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. Functional testing of the instrument using an in-house is3000 system passed recognition and engagement testing. The instrument moved intuitively with a full range of motion in all directions. The grips opened and closed properly. Engineering evaluation also found that the black tube insulation was damaged at the distal end. The insulation was gouged on one side and has a. 140 x. 090 piece missing/sections lifted up roughly 8 above the snake wrist. There were also various scratches with no material removal below this damaged section. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage to the instrument's tube insulation found during failure analysis could likely cause or contribute to an adverse event if the malfunctions were to recur. | 5 MM BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2973607 | 2949303 | 2819990 | 2955842201300442 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-11 | 2013-02-07 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-14 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the fenestrated bipolar instrument failed to grasp during the procedure. The instrument was removed and replaced and the planned surgical procedure was completed with an instrument of the same type. There were no missing or fallen pieces reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument passed recognition and engagement testing while installed on an in-house is3000 system. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Engineering also observed that one conductor wire was broken at the side of the yaw pulley. The wire segments stuck out at the wrist. The instrument failed electrical continuity testing. No evidence of arcing was observed. The side of the yaw pulley is also broken adjacent to the wire breakage and a piece of the yaw is missing. Engineering concluded that the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the damage to the instrument's yaw pulley found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2973630 | 2949339 | 2845861 | 2955842201300443 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-07 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42029602 | 2012-12-14 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the large suturecut needledriver instrument would not grasp or cut. There were no missing or fallen pieces reported. The planned surgical procedure was completed with a replacement instrument and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering evaluation found that when the instrument was installed on an in-house is3000 system, the instrument failed a cut testing. The scissors did not cut cleanly through a 0-silk suture and the suture was smashed between the blades. Multiple attempts were made to cut completely through the suture. The blades exhibited indentation and the shearing contact between the blades was low. Engineering evaluation also found that the distal end of the main tube exhibits scratch marks with light material removal. Engineering concluded that the damage to the instrument's main tube was likely caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself does not constitute a mdr reportable event; however; the damage to the instrument's main tube found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur. | LARGE SUTURCUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2973642 | 2949351 | 2845873 | 2955842201300444 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-07 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-01-07 | Malfunction |  | It was reported that during central processing, broken wires at the tip of the mega needle driver instrument were noted. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken near the proximal pulleys and proximal clevis cable hole. The idler pulley spins freely and does not exhibit damage. The cable segment sticks out at the wrist. Additional observation not reported by the site is proximal clevis wearing. The proximal clevis cable hole exhibits wear on one edge. Other cables at the wrist are not damaged. Evidence not conclusive, but wear on hole may have contributed to cable breakage. There was 1 use left. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2969877 | 2945440 | 2841937 | 2955842201300415 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-07 | 2013-02-06 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci si surgical procedure, broken wires on the large needle driver instrument were noted. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a frayed grip cable. The frayed cable sections are in various lengths sticking out at the wrist. No obvious damage was found on the pulley. The instrument has 7 uses remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2969921 | 2945505 | 2842021 | 2955842201300416 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-07 | 2013-02-06 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-11 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the large needle driver instrument could not be recognized. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering checked for instrument recognition on an in-house system. The system successfully recognized instrument, showing 4 uses left. Pogo pins do not stick and are not contaminated. The dallas chip programmer (dcp) verification of instrument passed. The instrument was found with main tube broken at the mid point of the shaft. Evidence was not conclusive, but damage likely due to excess force applied normal to the main tube. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2969927 | 2945509 | 2842025 | 2955842201300417 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-07 | 2013-02-06 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-01-04 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tip of the small clip applier instrument broke in half. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one applier tip is broken in half and did not get returned with the instrument. The missing piece measured roughly about 0. 103 x 0. 049. Damage is a result of misuse. There were 22 fires left. No other damage was found. On (b)(4) 2013 a follow up call was made with the reporter of this complaint. It was stated that the broken tip did not fall into a patient. It was found broken on a surgical tray, and the broken piece was discarded. No further information was available. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | SMALL CLIP APPLIER1 | ENDOSCOPIC INSTRUMENT | 2969995 | 2945521 | 2842038 | 2955842201300418 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-08 | 2013-02-06 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-11-06 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff alleged that a cable was coming out of the articulation of the prograsp forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was reportedly completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint of a frayed cable was confirmed. The grip cable was found to be frayed at the distal idler pulley. Frayed strands were found to be sticking out at the wrist. Other cables at wrist were not damaged. An additional finding not reported by the customer was scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage might have been due to mishandling. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2970099 | 2945584 | 2842101 | 2955842201300380 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-08 | 2013-02-06 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42018406 | 2012-11-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff alleged that the permanent cautery spatula instrument was not recognized by the arm. There were no missing or fallen pieces reported. The reporter denied that the instrument involved with this complaint was used on a patient. No patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint that the instrument was not recognized by the arm was not confirmed. The instrument was checked for recognition on the system. The system successfully recognized the instrument. The instrument's pogo pins were not sticking or found to be contaminated. Engineering was unable to replicate the alleged recognition issue on multiple attempts. An additional observation not reported by the customer was a broken pitch cable. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULAINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2970103 | 2945622 | 2842138 | 2955842201300403 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-06 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42034301 | 2012-12-19 | Malfunction |  | It was reported that during central processing, the thoracic grasper instrument was noted to have a cracked base. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | The instrument was returned and evaluated. Engineering observed that based on visual investigation, there are no finding of cracks anywhere. However, there were deep scratches on the main tube. Black tube insulation is damaged at the distal end. The metal shaft is exposed as a result. The insulation is lifted but no missing piece. Evidence not conclusive, but damage is likely due to misuse. The instrument has 13 uses left. On (b)(4) 2013 a follow up call was made with the customer regarding the failure analysis findings. It was stated that no fragments were observed falling into the patient. The reported instrument issue was found during cleaning and sterilization. No further information was available. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | THORACIC GRASPER | ENDOSCOPIC INSTRUMENT | 2970200 | 2945676 | 2842192 | 2955842201300419 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-08 | 2013-02-06 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-01-10 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the orange sleeve on the monopolar curved scissors instrument is cracked. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering observed that the tube extension is cracked and broken with missing a piece at the proximal clevis interface. The missing piece measured roughly about 0. 254 x 0. 197. The clevis is dislodged from tube extension. The tube extension likely broke due to excessive side loading. Electrical continuity passed. The instrument has 3 uses left. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | HOT SHEARS TM MONOPOLARCURVED SCISSORS | Reuse | 2970318 | 2945739 | 2842255 | 2955842201300420 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-01-10 | 2013-02-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42017207 | 2013-01-10 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the wires of the maryland bipolar forceps instrument were exposed at the distal tip. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is missing and the other one still installed in clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity passed. The instrument has 3 uses remaining. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2970559 | 2945956 | 2842472 | 2955842201300421 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-06 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-12-11 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si prostatectomy procedure, the monopolar curved scissors instrument was dull. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering placed the instrument on an in-house system for the latex cut test. Scissors do not cut cleanly through. 006 latex. Latex gets snagged at scissor tips. Blade edges are not damaged, but edges exhibit wear at the tips, negatively affecting cut performance. Electrical continuity passed. The instrument has 3 uses remaining. Additional findings revealed that the orange print from the extension tube is removed. The missing print measured roughly about 0. 157 x 0. 127. Evidence not conclusive but damage most likely due to improper cleaning. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | HOT SHEARS TM MONOPOLARCURVED SCISSORS | Reuse | 2970741 | 2946088 | 2842605 | 2955842201300422 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-01-10 | 2013-02-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42020505 | 2012-12-28 | Malfunction |  | It was reported that during a da vinci si diagnostic laparoscopy procedure, the jaw on the fenestrated bipolar forceps instrument stopped working. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument passed recognition and engagement testing when installed on an in-house is3000 system. The instrument moved intuitively with a full range of motion in all directions and the grips opened and closed properly. Engineering evaluation also found that the pitch down cable was frayed at the distal clevis hub. The frayed strands stuck out at the wrist. The other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2971035 | 2946371 | 2842888 | 2955842201300423 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
41032204 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff alleged that a cable broke on a vessel sealer instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | On (b)(4) 2013, the reporter denied that any fragments fell or that a patient was harmed/injured because of the alleged issue. The reporter indicated that the alleged issue occurred during a hysterectomy procedure that was completed with no post-operative complications. The surgeon reportedly used a different unspecified instrument after the alleged issue was noticed. The initial reporter indicated the instrument was not used on the patient. The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | VESSEL SEALING INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2971167 | 2946460 | 2842976 | 2955842201300424 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-07 | 2013-02-06 | No | Yes | Health Professional | 41032204 | No | No | 2013-01-07 | No | No | Unkown | 2013
42004906 | 2013-01-10 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted broken wires on the cadiere forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Additional observation not reported by site was main insulation tube damage. The main insulation tube exhibited wears marks with light material removal. The instrument's shaft felt rough. Additional observation not reported by the site was a broken pitch cable inside the housing. The broken pitch cable exhibited corrosion on the clamping pulley. Engineering concluded that damage was likely due to improper cleaning. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2971321 | 2946604 | 2843120 | 2955842201300425 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-06 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-10 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted a frayed wire on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable located at the distal clevis hub. There was no damage found at the clevis. The frayed segment was approximately 0. 03 in length. Additional observation not reported by site was clamping pulley rusted located at the backend of instrument's housing. Engineering concluded that damage was likely due to improper cleaning. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2971368 | 2946665 | 2843182 | 2955842201300426 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-06 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-01-11 | Malfunction |  | Event Description: --/-- It was reported that during a davinci si colon resection procedure, a piece of wire was sticking out of bowel grasper instrument. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was found to be broken at the sneak wrist and sticking out. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the broken cable if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | BOWEL GRASPER 30MM JAW INSTRUMENT | Reuse | 2971397 | 2946710 | 2843227 | 2955842201300427 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-02-06 | 2013-02-06 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No |  | 2013
42017910 | 2013-01-04 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted that the monopolar curve scissors instrument's tip would not open and/or close. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. The instrument was placed on davinci in-house system and it was driven. Instrument passed recognition and engagement tests. The instrument was able to move intuitively with full range of motion in all directions. The grips opened and closed properly. Additional observations not reported by site were scratches on the main insulation tube and tube extension damage. The distal end of main insulation tube exhibited various scratch marks with light material removal. The scratches were short in length and were not aligned with the tube axis. The main insulation tube extension was also found to be broken and it was missing a piece at the proximal clevis interface. The missing piece measured roughly about 0. 236 x 0. 191. Engineering concluded that damages were likely due to mishandling. Instrument passed electrical continuity test. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | HOT SHEARS TM MONOPOLAR CURVE SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2971456 | 2946780 | 2843297 | 2955842201300428 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-02 | Malfunction |  | The customer fenestrated bipolar forceps instrument was returned by the customer without calling first to the customer service department to initiate a complaint. | The instrument was returned and evaluated. Engineering observed loose pitch cables at the distal clevis hub. The pitch cables were also observed to be frayed. The clevis did not exhibit any damage or wear marks. Both cable crimps remained in clevis. Upon removal of the housing the pitch cable was also found to be broken at the clamping pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2971460 | 2946799 | 2841181 | 2955842201300429 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-06 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-02-10 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a davinci s surgical procedure, the davinci system could not recognized the monopolar curved scissors instrument. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode. Instrument was checked for recognition using an in-house davinci robot system. The davinci robot system successfully recognized the monopolar curved scissors instrument, showing 5 uses left. The instrument's pogo pins did not appear to be stick and were not contaminated. Instrument passed electrical continuity test. The instrument was driven and it moved intuitively, the grips were able to open and close without any difficulty. Additional observation not reported by site was main insulation tube damage. The distal end of main insulation tube had various scratch marks with light material removal. The scratches were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action number 2955842-051613-005 to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | HOT SHEARS TM MONOPOLAR CURVED SCISSORS INSTRUMENT | Reuse | 2971550 | 2946834 | 2843350 | 2955842201300430 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2013-01-11 | 2013-02-06 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | Yes | 2013-04-27 | Device Not Retu | No |  | 2013
42000606 | 2013-01-07 | Malfunction |  | It was reported that during a davinci s hysterectomy procedure, the customer noted that the wires have been turn off, disrupted, in the working part of the large needle driver instrument, which lead to impossibility for movement and respectively use of it, the instrument was exchanged with a new one and the procedure was successfully completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment sticks out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2968418 | 2943929 | 2840445 | 2955842201300362 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-08 | 2013-02-05 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42805206 | 2013-01-04 | Malfunction |  | It was reported that after a da vinci si total benign cholecystectomy procedure, the surgical staff alleged that the shaft of the monopolar cautery instrument appeared to be damaged. The alleged issue was reportedly identified during central processing. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint that the instrument's shaft was damaged was confirmed. The instrument's main tube was found to be damaged at the tip. The insulation appeared to be damaged and a small piece was peeling off. No material was observed to have been removed. Engineering concluded that the damage was likely due to mishandling of the instrument. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | 5MM MONOPOLAR CAUTERYINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2968928 | 2944366 | 2840883 | 2955842201300363 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-07 | 2013-02-05 | No | Yes | Health Professional | 42805206 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-04 | Malfunction |  | It was reported that after completion of a da vinci si surgical procedure, the surgical staff alleged that frayed cables were observed on the prograsp forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint of frayed cables was confirmed. The pitch up cable was found to be frayed at the distal clevis hub. Frayed strands were observed to be sticking out at the instrument's wrist. Other cables at the wrist were reportedly not damaged. No other instrument damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2968955 | 2944437 | 2840953 | 2955842201300364 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-07 | 2013-02-05 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-11 | Malfunction |  | It was reported that during reprocessing the megasuturecut needle driver instrument, it was noted that the wires at the distal end of the instrument were broken. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the grip cable was broken. The broken cable section stuck out at the wrist. The idler pulley rim and face exhibited damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2968968 | 2944448 | 2840964 | 2955842201300365 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-13 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff alleged that the pk dissecting forceps instrument had a recognition issue with the system. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint of that the instrument was not recognized by the system could not be confirmed. The instrument was checked for recognition on an in-house is3000 system. The is3000 system successfully recognized the instrument. The pogo pins of the instrument were not found to be sticking and were not contaminated. Engineering was unable to replicate the alleged recognition issue. An additional observation not reported by the site was a broken cable. The pitch down cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the instrument's wrist were not damaged. No other instrument damage was observed. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2968995 | 2944459 | 2840976 | 2955842201300366 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-07 | 2013-02-05 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-08-31 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgical staff alleged that the maryland bipolar forceps instrument had a faulty wrist movement issue. The planned surgical procedure was completed using an alternative replacement instrument and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint that the instrument had a faulty wrist movement issue could not be confirmed. The instrument was placed on an is3000 system and driven. Recognition and engagement of the instrument passed. The instrument moved intuitively with full range of motion in all directions. The instrument grips opened and closed properly. Functional performance testing of the instrument did not find any issues. An additional observation not reported by the site was tube damage. The distal end of the instrument's main tube had various scratch marks with light material removal. The scratches were. 070 -. 120 in length and not aligned with the tube axis. Engineering concluded that the damage might have been caused by mishandling/misuse. No other instrument damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2968985 | 2944463 | 2840980 | 2955842201300369 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-07 | 2013-02-05 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-09 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted frayed cable, wires sticking out at the end of the prograsp forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Engineering also observed scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2969024 | 2944471 | 2841028 | 2955842201300370 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-09 | 2013-02-05 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-09 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wire on the megasuturecut needle driver instrument was observed to be frayed at the tip. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the wire at the tip of the instrument was frayed. The one grip close cable was broken at the distal idler pulley. The idler pulley spins freely but has damage on the edge and on the surface. Engineering concluded that the damage was likely due to mishandling. The cable segment stuck out at the wrist. The other cables at the wrist were not damaged. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969001 | 2944478 | 2839190 | 2955842201300368 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-14 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff alleged that a cable popped out of a large needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint that a cable popped out of the instrument was confirmed. One grip close cable was found to be broken at the distal idlers. The idler pulley was observed to spin freely and was not damaged. A cable segment was found to be sticking out at the instrument's wrist. Other cables at the wrist were not damaged. An additional finding not reported by the site were scratches and indentations on the distal pulleys. Both pulleys were found to have scratches on the surfaces and indentations on the edges. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2968982 | 2944482 | 2840998 | 2955842201300367 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-07 | 2013-02-05 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
420227 | 2012-06-01 | Malfunction |  | On (b)(4) 2013, isi received a user facility medwatch form from the fda. According to the information reported to the fda by the initial reporter, 'broken wire noted in davinci instrument after cleaning - no problem noted during surgery. '. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969044 | 2944537 | 2841053 | 2955842201300371 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 420227 | No | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-10 | Malfunction |  | Event Description: --/-- It was reported that during set-up of a da vinci si surgical procedure, the cables on the pk dissecting forceps instrument were observed to be frayed. | Manufacturer Narrative: --/-- On (b)(4) 2013, isi contacted the initial report concerning the reported event. The initial reporter indicated that there has been no report by the surgical staff that any piece from the reported affected instrument fell into a patient during a surgical procedure. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering evaluation found that the damaged instrument component was a broken pitch cable. Based on this clarification, the customer reported complaint is confirmed. The pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is remained installed in the clevis. The clevis does not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and are not aligned with the tube axis, suggesting that the damage was caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | Reuse | 2969096 | 2944549 | 2841065 | 2955842201300373 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2013-02-11 | 2013-02-05 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-02-11 | Device Not Retu | No |  | 2013
42027804 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci si colon resection procedure the wires at the wrist of the grasping retractor instrument broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit any damage or wear marks. Engineering also found that both of the distal pulleys exhibited scratches on the surface. The distal end of the main tube had various scratch marks with light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969100 | 2944553 | 2841115 | 2955842201300374 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-02 | Malfunction |  | It was reported that during set-up of a da vinci si surgical procedure, the tips on the fenestrated bipolar forceps instrument were observed to be bent. There were no missing or fallen pieces reported. | The instrument was returned and evaluated Per the customer reported complaint engineering observed that the instrument was returned with one bent grip which caused side to side misalignment of the grips There was a 075 offset at the tips indicating that overloading at the tip occurred The instrument also passed electrical continuity testing The instruments and accessories user manual specifically states general precautions and warnings handle instruments with care Avoid mechanical shock or stress that can cause damage to the instruments Do not use an instrument to clean debris from another instrument intraoperatively This may result in damage to the instruments or other unintended consequences such as disconnection of the instrument tip Engineering evaluation also found that the pitch cable at the distal clevis hub was broken The cable segment that contains the crimp was remained installed in the clevis The clevis did not exhibit any damage or wear marks The surface of the distal pulley also exhibited scratches The customer reported complaint does not itself constitute a mdr reportable event however the damage to the instruments pitch cable found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969076 | 2944557 | 2841073 | 2955842201300372 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-09 | 2013-02-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-09 | Device Not Retu | No | Reuse | 2013
42027804 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci si colon resection procedure the wires at the wrist of the grasping retractor instrument broke. No missing or fallen pieces were reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint engineering found that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis did not exhibit any damage or wear marks. Engineering evaluation found that the distal end of the main tube had various scratch marks with light material removal. The scratches were short in length and not aligned with the tube axis. Engineering concluded that damage to the main tube was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969136 | 2944598 | 2841115 | 2955842201300375 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-13 | Malfunction |  | It was reported that after completion of a da vinci si prostatectomy procedure, the surgical staff alleged that a cable from a maryland bipolar forceps instrument appeared to be damaged. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The customer initially reported a damaged cable. However for clarification, the nonconformance was an exposed conductor wire. Based on this, the complaint was confirmed. A small section of the insulation had been removed off allowing the conductor wire to be seen. Electrical continuity testing of the instrument passed. No other instrument damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969139 | 2944600 | 2841117 | 2955842201300377 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-08 | 2013-02-05 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that the prograsp forceps instrument had a broken wire at the distal end. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported a broken wire issue. However for clarification, the nonconformance was a broken pitch cable. Based on this, the complaint was confirmed. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969142 | 2944601 | 2811073 | 2955842201300378 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-08 | 2013-02-05 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42000306 | 2013-01-08 | Malfunction |  | It was reported prior to starting a da vinci s surgical procedure the small clip applier instrument tip broke off while loading clips. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed both clip applier grips are broken. The missing pieces measured roughly about 0. 116' x 0. 050. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969145 | 2944602 | 2841119 | 2955842201300384 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-05 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-14 | Malfunction |  | It was reported that during central processing the fenestrated bipolar forceps instrument was noted to have frayed cables. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the drive cable is frayed. The pitch up cable is frayed at the distal clevis hub. The frayed strands stick out at the wrist. The other cables at the wrist are not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969147 | 2944603 | 2841120 | 2955842201300388 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-14 | 2013-02-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-28 | Malfunction |  | It was reported that during a da vinci si nephrectomy procedure, the surgical staff alleged that a wire came off the track on the top of the fenestrated bipolar forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint that a wire came off a track was confirmed. One grip close cable was found to be derailed at the distal idler pulley. The instrument grips were still able to open and close. Engineering concluded that the cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. Engineering also concluded that possibly the fleet angle between the proximal and distal pulleys might have contributed to derailment. Electrical continuity testing of the instrument passed. An additional finding not reported by the site was scratches on the instrument's main tube. The distal end of the main tube had various scratch marks with light material removal. The scratches were short in length and not aligned with the tube axis. Another finding not reported by the site was a bent bipolar pin. The bottom pin was found to be bent upward. Engineering concluded that the observed damages might have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969172 | 2944605 | 2841122 | 2955842201300391 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-09 | 2013-02-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-09 | Device Not Retu | No | Reuse | 2013
42031802 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci si proctocolectomy procedure, the wires on the small grasping retractor instrument broke and a fragment from the instrument fell into the patient. The broken fragment was retrieved and the planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable at the distal clevis hub is broken and the cable segment that contains the crimp was missing from the clevis. The clevis does not exhibit excessive damage. The broken cable segment was also returned. On (b)(4) 2013 isi contacted the initial reporter of this incident. The initial reporter indicated that the surgeon reported that he was not grasping tissue with the small grasper retractor instrument as the instrument was not responding. The initial reporter indicated the surgeon used another instrument to aid in the removal of the small grasper retractor instrument. The initial reporter indicated that to the best of their knowledge, the patient has not returned to the hospital due to any post surgical complications as a result of the reported event. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969157 | 2944606 | 2841123 | 2955842201300392 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-09 | 2013-02-05 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | Yes | 2013-01-09 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-09 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a broken wire on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with frayed cable sections located at various lengths, the frayed cables were sticking out on the wrist. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969173 | 2944607 | 2841124 | 2955842201300394 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the surgical staff alleged that the pk dissecting forceps instrument had a broken wire issue. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint of a broken wire was confirmed. The instrument's pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. An additional finding not reported by the site was scratches on distal pulley. Scratches on the surface of both distal pulleys were observed. Another finding was scratches on the instrument's main tube. The distal end of the main tube had various scratch marks with light material removal. The scratches were short in length and did not align with the tube axis. Engineering concluded that the damages might have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969161 | 2944608 | 2841125 | 2955842201300398 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-11 | 2013-02-05 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-09 | Malfunction |  | It was reported that during set-up of a da vinci si surgical procedure, it was noted that the cable on the large needle driver instrument was broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with one grip open/close cable frayed at the distal idler pulley. The frayed strands stick out at the wrist. The other cables at the wrist are not damaged. Engineering evaluation also found that the distal idler pulley on which the broken cable was seated had a section that is bent inwards. Engineering concluded that the damage to the pulley is likely due to mishandling and that the damage likely contributed to the cable breakage. The distal end of the main tube has various scratch marks with light material removal. The scratches are. 090 -. 220 in length and were aligned with the tube axis. Engineering concluded that the damage to the main tube was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969180 | 2944610 | 2841127 | 2955842201300400 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-09 | 2013-02-05 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-09 | Device Not Retu | No | Reuse | 2013
42020405 | 2013-01-03 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported seeing a broken wire on a atrial retractor instrument. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The alleged complaint that the instrument had a broken wire was confirmed. The instrument was found to have frayed pitch cable at the distal clevis hub. No damage was found at the clevis. A frayed segment was observed to be approximately 0. 02 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | ATRIAL RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969088 | 2944613 | 2841129 | 2955842201300376 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2013-01-08 | 2013-02-05 | No | Yes | Health Professional | 42020405 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-10 | Malfunction |  | It was reported that during set-up of a da vinci si surgical procedure, the cables on the megasuturecut needle driver instrument were found to frayed. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was frayed at the distal idler pulley and the frayed strands stuck out at the wrist. The other cables at wrist did not exhibit damage. Engineering evaluation also found that the same frayed grip cable is derailed at the distal idler pulley. The grips still opened and closed; however the movement may not be precise. Engineering concluded that derailment of the cable was likely due to losing contact with the pulley during wrist articulation. The fleet angle between the proximal and distal pulleys may have contributed to the cable derailment. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969141 | 2944618 | 2841134 | 2955842201300379 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42017910 | 2013-01-09 | Malfunction |  | It was reported that during set-up of a da vinci si surgical procedure, it was noted that the blade on the monopolar curved scissors instrument had a chip. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the blade on the instrument has a chip. Engineering concluded that the damage is likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969152 | 2944620 | 2841136 | 2955842201300381 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-09 | 2013-02-05 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-01-09 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-12-26 | Malfunction |  | Event Description: --/-- It was observed during central processing that the tip cover of the monopolar curved scissors instrument was hard to remove. Due to the difficulty of removing the tip cover, the site thought the instrument's jaws were damaged. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. The alleged complaint of the tip cover being hard to remove and that the instrument's jaws were damaged could not be confirmed. The instrument was returned with the tip cover installed. The tip cover was able to be removed easily with a twisting motion. Friction between the tip cover and tube extension was not excessive. A visual inspection of the instrument revealed no damage to the blades. The instrument blades were able to open/close fine when inputs were manually rotated. Am additional observation not reported by the site was tube damage. The distal end of the instrument's main tube had various scratch marks with light material removal. The scratches were. 070 -. 090 in length and not aligned with the tube axis. Engineering concluded that the damage might have been caused by mishandling/misuse. No other instrument damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- This mdr is being submitted for retrospective activity performed relating to field action (b)(4) to investigate micro-cracks on the monopolar curved scissors instrument. These types of micro-cracks will not lead to mechanical failure of the instrument; however, there is a potential for insulation failure after reprocessing, resulting in a pathway for electrosurgical energy to leak to tissue and potentially cause unintended injuries. The location of theses micro-cracks is confined to 2 cm of the distal end of the instrument shaft. This instrument was inspected as part of this retrospective activity and found to contain cracks on the instrument main tube. | HOT SHEARS MONOPOLARCURVED SCISSORS INSTRUMENT | Reuse | 2969153 | 2944621 | 2841137 | 2955842201300382 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-01-08 | 2013-02-05 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-04-27 | Device Not Retu | No |  | 2013
42020706 | 2012-12-24 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the cable on the tenaculum forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the pitch cable broken at the distal clevis hub. The cable segment that contains the crimp was remained installed in the clevis. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969144 | 2944622 | 2820582 | 2955842201300383 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-09 | 2013-02-05 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-01-09 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-01-14 | Malfunction |  | It was reported during central processing that the mega needle driver instrument was noted to have broken wires. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken grip cable. The frayed cable sections are in various lengths sticking out at the wrist. The idler pulley exhibited damage to the face and rim. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969146 | 2944624 | 2841140 | 2955842201300385 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-14 | 2013-02-05 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-01-10 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, it was noted broken wires on the mega needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a broken grip cable. The broken cable section sticks out at the wrist. The idler pulley face also exhibited damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969148 | 2944625 | 2841141 | 2955842201300387 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-12-19 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the surgical staff alleged that the tip cover part of the monopolar curved scissors instrument was broken. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint the tip cover part of the instrument was broken off was confirmed. The instrument's tube extension was found to be dislodged from the main tube. The entire tube extension wall was circumferentially broken at the base of the axial keys. As a result, the tube extension was able to be rotated 360 degrees. No pieces were observed to be missing. Engineering concluded that the damage was possibly due to excessive side loading or another type of mishandling. Electrical continuity testing of the instrument passed. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Multiple attempts by intuitive surgical inc. To contact the site for follow-up information have been unsuccessful at this time. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | HOT SHEARS MONOPOLARCURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969170 | 2944626 | 2841142 | 2955842201300386 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-08 | 2013-02-05 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-14 | Malfunction |  | It was reported that during central processing the fenestrated bipolar forceps instrument were noted to have frayed cables. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the drive cable is frayed. The pitch down cable is frayed at the distal clevis hub. The frayed strands stick out at the wrist. The other cables at wrist are not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969155 | 2944627 | 2841120 | 2955842201300389 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-14 | 2013-02-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42017704 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci si colectomy surgical procedure the bowel grasper instrument was noted to have a wire sticking out of the tip of the instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed one drive cable is broken at the snake wrist. The broken cable segment sticks out of the snake wrist. The wrist motion is not intuitive as a result of the cable breakage. The other cables at the wrist are not damaged. Engineering also observed scratches on the maintube. The insulation is gouged in various places and has a. 036 x 0. 032 piece of insulation missing. The metal shaft is exposed as a result. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | BOWEL GRASPER 30MM JAW INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969156 | 2944628 | 2841144 | 2955842201300390 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-05 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-09 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was noted that the wire on the fenestrated bipolar forceps instrument was broken. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable at the distal clevis hub was broken. The cable segment that contains the crimp remained installed in the clevis. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969158 | 2944630 | 2841146 | 2955842201300393 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-09 | 2013-02-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-09 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-01-07 | Malfunction |  | It was reported that during suturing of da vinci si total benign hysterectomy procedure, the surgical staff alleged that a wire snapped on the large suturecut needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint that an instrument wire snapped was confirmed. The grip cable was found to be broken at the distal idlers. The idler pulley was observed to spin freely and was not damaged. A cable segment was found to be sticking out at the instrument's wrist. Other cables at the wrist did not appear to be damaged. No other instrument damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969159 | 2944631 | 2841147 | 2955842201300395 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-09 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff alleged that wires were exposed on the prograsp forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The alleged complaint that wires were exposed was confirmed. The grip cable was found to be frayed at the distal idler pulley. The frayed strands were observed to be sticking out at the wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969160 | 2944632 | 2841148 | 2955842201300396 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-05-25 | Malfunction |  | It was reported during a da vinci si colectomy procedure, the surgeon was unable to move the fenestrated bipolar forceps instrument smoothly and it was noted that the wire at the distal end of the instrument was visible; however, the wired did not appear to be broken. Reportedly, no instrument collisions were noted to have occurred. The instrument was replaced and the planned surgical procedure was completed. No missing or fallen pieces were reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering found that the pitch cable at the distal clevis hub was broken. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit any damage or wear marks. The instrument passed electrical continuity testing. Engineering evaluation also found that the distal end of the main tube had various scratch marks with light material removal. The scratches were short in length and not align with the tube axis. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969174 | 2944633 | 2841149 | 2955842201300397 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-01-18 | 2013-02-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-18 | Device Not Retu | No | Reuse | 2013
40027202 | 2013-01-11 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a harmonic ace insert fell into a patient. The insert was retrieved and the surgical procedure was completed. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | HARMONIC ACE CURVED SHEARSINSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2969177 | 2944635 | 2841151 | 2955842201300399 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-11 | 2013-02-05 | No | Yes | Health Professional | 40027202 | No | Yes | 2013-01-11 | No | Yes | Unkown | 2013
40027202 | 2013-01-11 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a harmonic ace curved shears insert used in conjunction with the harmonic ace curved shears instrument fell into a patient. The insert was retrieved and the surgical procedure was completed. There was no allegation of any harm, injury, or adverse outcome to a patient. | The accessory has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | HARMONIC ACE CURVED SHEARSINSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2969178 | 2944637 | 2841151 | 2955842201300401 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-11 | 2013-02-05 | No | Yes | Health Professional | 40027202 | No | Yes | 2013-01-11 | No | No | Unkown | 2013
42009308 | 2013-01-10 | Malfunction |  | It was reported that prior to starting a davinci s surgical procedure, the customer noted a broken wire on the prograsp forceps instrument. The instrument was not used in the patient. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip open cable was found to be frayed at the distal idler pulley. The frayed strands stick out at the wrist. Other cables at wrist were not damaged. Additional observation not reported by site was main insulation tube damage. Distal end of main tube had various scratch marks with light material removal. The scratches were 0. 100 - 0. 180 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969165 | 2944643 | 2841148 | 2955842201300402 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-04 | Malfunction |  | It was observed during central processing that the fenestrated bipolar instrument had a broken cord. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The alleged complaint that the instrument had a broken cord was confirmed. The instrument's pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969212 | 2944664 | 2841181 | 2955842201300405 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-14 | 2013-02-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-11 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff allegedly observed broken/loose cables on the maryland bipolar forceps instrument. The surgical staff also claimed that the instrument was not articulating correctly. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The alleged complaint that the instrument had a broken cable was confirmed. The instrument's pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing of the instrument passed. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969206 | 2944668 | 2841195 | 2955842201300406 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2013-01-15 | 2013-02-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-14 | Malfunction |  | It was reported that during a da vinci si pyeloplasty procedure, the surgical staff alleged that a wire was observed to be hanging on the side of a large needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported a wire hanging issue. However for clarification, the nonconformance was a broken grip cable. Based on this information, the alleged complaint was confirmed. Frayed cable sections were found in various lengths sticking out at the instrument's wrist. An additional finding not reported by the customer were several indentation and ridges observed on the idler pulley. Engineering concluded that the damage might have been a result of misuse. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969203 | 2944678 | 2841194 | 2955842201300404 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-14 | 2013-02-05 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-11 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff allegedly observed broken/loose cables on the fenestrated bipolar forceps instrument. The surgical staff also alleged that the instrument was not articulating correctly. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The alleged complaint that the instrument had a broken cable was confirmed. One conductor wire was found to be broken at the yaw pulley exit. The wire was detached from its connection at the grip. Electrical continuity testing of the instrument failed. The yaw pulley showed no signs of arcing. An additional finding not reported by the customer was a bent grip on the instrument. One grip was bent, causing side to side misalignment of grips. There was a 0. 067 offset at the tips, indicating overloading at the tip. Another finding not reported by the site was a bent bipolar pin. The bottom pin was bent upward and both pins were loose. Engineering concluded that the instrument was likely mishandled. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969214 | 2944679 | 2841195 | 2955842201300407 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-05 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-15 | Malfunction |  | It was observed during central processing that the mega suturecut needle driver instrument had a broken wire. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The customer initially reported a broken wire issue. However, for clarification, the nonconformance was a broken grip cable. Based on this, the alleged complaint was confirmed. The broken cable section was found to be sticking out at the instrument's wrist. An additional finding not reported by the site was a damaged idler pulley. The idler pulley exhibited rim and face damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969232 | 2944687 | 2841210 | 2955842201300408 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-15 | Malfunction |  | It was observed during central processing that the mega suturecut needle driver instrument had a broken wire. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The customer initially reported a broken wire issue. However for clarification, the nonconformance was a broken grip cable. Based on this, the alleged complaint was confirmed. The broken cable section was found to be sticking out at the instrument's wrist. An additional finding not reported by the site was a damaged idler pulley. The face of the idler pulley exhibited damage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969221 | 2944693 | 2841210 | 2955842201300409 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-15 | 2013-02-05 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42020906 | 2013-01-09 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a cable broke on the mega suturecut needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The alleged complaint of a broken cable was confirmed. One grip close cable was found to be broken at the distal idlers. The idler pulley was observed to spin freely and was not damaged. A cable segment was found to be sticking out at the instrument's wrist. Other cables at the wrist were not damaged. An additional finding not reported by the site was scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Engineering concluded that the damage might have been due to mishandling. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969228 | 2944701 | 2841218 | 2955842201300410 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-15 | 2013-02-05 | No | Yes | Health Professional | 42020906 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wire at the end of the pk dissecting forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969255 | 2944705 | 2841222 | 2955842201300414 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable on the tenaculum forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969252 | 2944733 | 2841249 | 2955842201300411 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42020706 | No | No | 2013-01-10 | No | No | Reuse | 2013
42000606 | 2013-01-10 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable on the large needle driver instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with a frayed grip cable. The frayed cable sections were in various lengths and stuck out at the wrist. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2969253 | 2944734 | 2841250 | 2955842201300412 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-09 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wire at the end of the pk dissecting forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch down cable was broken at the distal clevis hub. The cable segment that contains the crimp remained installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. Engineering evaluation also found that one yaw pulley cap exhibited a char mark at the grip base. The char mark was only present on one yaw pulley, and was not located between the two grips. The conductor wires did not exhibit visible damage and the instrument passed electrical continuity testing. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2969271 | 2944736 | 2841222 | 2955842201300413 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-10 | 2013-02-05 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | Yes | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-04 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a broken hinge on the prograsp forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found to be broken at the proximal clevis to the main tube interface. The main tube was cracked at the interface. The clevis was found to be dislodged from the main tube as a result. No missing pieces. The breakage was likely due to overloading at the tip. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2966816 | 2942421 | 2838937 | 2955842201300344 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-08 | 2013-02-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42011006 | 2012-10-29 | Malfunction |  | It was observed during central processing that the tip of a precise bipolar forceps instrument had a broken wire. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The complaint of a broken wire at the tip of instrument was confirmed. The pitch up cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Other cables at the wrist were not damaged. An additional observation not reported by the site was tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 075 -. 125 in length and not aligned with the tube axis. Evidence is not conclusive, but the damage may have been caused by mishandling/misuse. No other damage was found. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PRECISE BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2967025 | 2942615 | 2839131 | 2955842201300345 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-08 | 2013-02-04 | No | Yes | Health Professional | 42011006 | Device Returned To Manufacturer | Yes | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-12 | Malfunction |  | It was reported during a da vinci si surgical procedure the pk dissecting forceps instrument read as expired on screen. The drape was reseated and the instrument reinserted. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the instrument was checked for recognition on a da vinci si system. The da vinci si system successfully recognized the instrument on multiple attempts. The pogo pins do not stick and are not contaminated. The instrument registered with two uses left. Instrument is not expired. The dcp verification of instrument passed, indicating two uses left as well. Engineering also observed a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2967078 | 2942624 | 2839140 | 2955842201300346 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-01-23 | 2013-02-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-01-23 | Device Not Retu | No | Reuse | 2013
42019408 | 2013-01-23 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the mega needle driver instrument was noted to have wires sticking out at the distal end. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967079 | 2942625 | 2839141 | 2955842201300347 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-04 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci si surgical procedure the prograsp forceps instrument was allegedly locking repeatedly in open position while in use on the system. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the instrument was placed on system and driven. Recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. Engineering also observed the pitch cable frayed at the distal clevis hub. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967080 | 2942626 | 2839142 | 2955842201300349 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-23 | 2013-02-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-02 | Malfunction |  | It was reported that a megasuture cut needle driver instrument had a frayed grip cable. Per the reporter, the alleged issue was not identified during a procedure. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. However, the customer reportedly returned the subject instrument without a complaint. Therefore, an alleged complaint cannot be confirmed or denied. However, an observation not reported by the customer was a frayed grip cable. The frayed grip cable was located at the distal idler pulley. Frayed cable sections are in various lengths sticking out at the wrist. The frayed grip cable is also derailed. The idler pulley exhibited rim and face damage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2967066 | 2942655 | 2839171 | 2955842201300329 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-07 | 2013-02-04 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42020706 | 2013-01-23 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the tenaculum forceps instrument was noted to have broken wires at the distal end. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967090 | 2942660 | 2839176 | 2955842201300348 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-04 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-16 | Malfunction |  | It was reported that during a da vinci si prostatectomy surgical procedure the prograsp forcep instrument was not responding to the system. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the cables are derailed. The one grip open cable is derailed from the force amp pulley. The grips are yawed over to one side and unable to properly straighten due to the derailment. Engineering also observed a derailed backend cable and tube damage. The pitch cables are loose at the wrist, but not visibly broken. The housing was removed to find one pitch cable derailed from the back idler pulley. The cable is wedged between the two pulleys and has lost tension. The clamping pulley screws are still tight. The distal end of main tube has various scratch marks with light material removal. The scratches are. 075 -. 165 in length and are not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967092 | 2942661 | 2839177 | 2955842201300350 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-24 | 2013-02-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No | Reuse | 2013
42029602 | 2013-01-16 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure the large suturecut needle driver instrument jaws would not open or close. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the instrument would not open or close. The intuitive motion was not being experienced because the grip cable was broken. The cable segment sticks out at the wrist. The other cables at wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967093 | 2942662 | 2839178 | 2955842201300351 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-25 | 2013-02-04 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-01-25 | Device Not Retu | No | Reuse | 2013
42004906 | 2012-12-10 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the cadiere forceps instrument tip did not respond properly to the system. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967095 | 2942663 | 2839179 | 2955842201300352 | 1 | NAY | Manufacturer | ForeignDistributor | Other | Initial | 2013-01-29 | 2013-02-04 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-01-29 | Device Not Retu | No | Reuse | 2013
42022703 | 2013-01-15 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the pk dissecting forceps instrument was noted to have one of the cables broken. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch up cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The other cables at the wrist are not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2967096 | 2942664 | 2839180 | 2955842201300354 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-16 | 2013-02-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
| 2012-11-08 | Malfunction |  | On (b)(6) 2013, intuitive surgical inc. (isi) received (b)(4) from the fda via mail. Based on the report, the site alleged that a mega needle driver instrument malfunctioned. On (b)(6) 2013, a follow-up contact was made between an isi regulatory complaints analyst and the site. The reporter clarified that the instrument had a broken cable at the distal end and the wires were all frayed. The reporter denied that there was any harm, injury, or adverse outcome to a patient because of the alleged issue. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967083 | 2942665 | 2839181 | 2955842201300353 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-18 | 2013-02-04 | No | Yes | Health Professional |  | No | No | 2013-01-18 | No | No | Unkown | 2013
42019408 | 2013-01-11 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the mega needle driver instrument was noted to have frayed wires. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken grip cable. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967085 | 2942666 | 2839182 | 2955842201300355 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-17 | 2013-02-04 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-08 | Malfunction |  | It was reported during a da vinci s hysterectomy procedure the mega suturecut needle driver instrument tips would not respond to the system. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the tips would not open or close. The grips would not open because the grip cable was broken which supports the open and close function. The broken cable segment sticks out at the wrist. The idler pulley spins freely and exhibited damage to the pulley face and rim. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967106 | 2942674 | 2839190 | 2955842201300358 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-23 | 2013-02-04 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-23 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-11-29 | Malfunction |  | It was reported that during a da vinci si surgical procedure the prograsp forceps instrument stopped opening and closing, and had a broken cable. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch up cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The other cables at the wrist are not damaged. Engineering also observed tube damage. The distal end of the main tube has various scratch marks with light material removal. The scratches are. 205 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. On (b)(4) 2013 a follow-up call was made with the customer regarding the reported complaint. It was stated that no fragments were observed falling into the patient. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967105 | 2942693 | 2839209 | 2955842201300357 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-22 | 2013-02-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-01-22 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-11 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure it was noted that the wires on the tip of the pk dissecting forceps instrument were loose. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the wires on the tip were loose. The pitch up cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The other cables at the wrist are not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2967107 | 2942696 | 2839212 | 2955842201300359 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-21 | 2013-02-04 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-21 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-28 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff alleged that one of the 4 circles was not turning on the maryland bipolar forceps instrument. In addition, the surgical staff claimed that possibly a cable was off a track. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The complaint that one of the four circles was not turning is confirmed. The pitch cable was found to be loose and frayed at the distal clevis hub. The actual broken location was inside the housing by the pulley that creates the free movement of the disc. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Electrical continuity testing passed. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2967109 | 2942700 | 2839216 | 2955842201300356 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-08 | 2013-02-04 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff alleged that the fenestrated bipolar forceps instrument had non-real fluid movement. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint of non-real fluid movement was confirmed. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear mark. Additional finding: scratches were found on the instrument's main tube. The distal end of the main tube had various scratch marks with light material removal. The scratches were short in length and not aligned with the tube axis, suggesting the damage might have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2967131 | 2942712 | 2839228 | 2955842201300360 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-08 | 2013-02-04 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42009308 | 2013-01-07 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the surgical staff reported seeing a broken cable. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported a broken cable. However for clarification, the nonconformance was a frayed pitch cable. Based on this, the complaint was confirmed. A frayed pitch cable was found at the distal clevis hub. A frayed segment was found to be sticking out. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2967142 | 2942722 | 2839238 | 2955842201300361 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-09 | 2013-02-04 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-09 | Device Not Retu | No | Reuse | 2013
42018305 | 2012-12-27 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff reported seeing a burned tip on the permanent cautery hook instrument nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of instrument tip was burned was confirmed. Instrument was found with charred mark on the proximal clevis near the main tube connection. There was a hairline crack found on the clevis near the char marked. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2965314 | 2940423 | 2836939 | 2955842201300343 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-04 | 2013-02-03 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2013-01-04 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-26 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that the tip of the fenestrated bipolar forceps instrument had frayed wires. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported a frayed wire at the tip of the instrument. However for clarification, the nonconformance observed through engineering was a broken pitch cable. Based on this observation, the alleged complaint was confirmed. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2965335 | 2940437 | 2836954 | 2955842201300342 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-04 | 2013-02-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-04 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-04 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a broken wire was discovered on the fenestrated bipolar forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported a broken wire. However for clarification, the nonconformance was a frayed grip cable. Based on this observation, the alleged complaint was confirmed. Frayed cable sections were observed sticking out at the wrist. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2965334 | 2940438 | 2836953 | 2955842201300341 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-04 | 2013-02-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-04 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2013-01-02 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the surgeon had articulation issues with one of the patient side manipulators (psm). The staff reportedly re-seated the instrument and eventually replaced the drape on that psm. However, the issue persisted intermittently. It was stated that the arm jumped, causing the instrument held in that psm to stab an artery, which led to some bleeding that the surgeon was able to control without converting the procedure. The scheduled procedure was completed successfully. | The field service engineer (fse) performed a test drive at the customer site with the alleged psm arm, and observed no non-intuitive or unexpected motion. A brake test on the patient side manipulator (psm) and set-up joint (suj) was performed by the fse, and both passed according to isi standards. The fse exchanged the psm as a precaution. The system is working according to isi specifications. The psm was returned and evaluated. Engineering performed a visual inspection of the psm. The arm was configured to the system and it powered up in maintenance mode. The check-sensors, friction, and brake tests as well as an hour long sine cycle were run. The psm was also restarted on the in-house system in normal mode. The arm was driven with an instrument for several minutes in several different positions. Engineering tried bumping the psm in several spots to attempt and get it to move on its own; however, the reported problem was not able to be recreated or verified. No repairs are required for this issue. On (b)(6) 2013, the reporter of this complaint was contacted for additional information. It was stated that the patient was doing well, and was seen for a regular check-up recently. The injury to the artery was fixed by the surgeon as soon as it was discovered during the procedure, and the surgery did not have to be converted as a result. The site reportedly did not experience any recurrence of this event since the reported complaint. This was the only time that this has happened without the surgeon moving or repositioning/controlling the arm movement, according to the reporter. There were no malfunctions of the instruments/ accessories reported. No further information was provided. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2963792 | 2939247 | 2835765 | 2955842201300332 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-02 | 2013-02-01 | Yes | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
IS2000 A51P7 | 2013-01-04 | Injury | Required Intervention | It was reported that during a da vinci si hysterectomy procedure, the patient's ureter was damaged. During repair of the patient's ureter by the site's urologist, the urologist noted that the jaw on the endowrist instrument installed on the patient side manipulator (psm) arm was not articulating correctly. Reportedly, the site replaced the instruments multiple times and replaced the drape; however, the issue persisted. The site reported that during installation of the sterile adapter onto the psm, it was noted that the sterile adapter did not appear to engage fully with the psm. Unable to resolve the issue, the surgeons made the decision to repair patient's ureter and complete the surgical procedure using open surgical techniques. | The investigation conducted by the field service engineer (fse) found that the customer reported failure mode was related to the sterile adapter release tab on the patient side manipulator arm 1, which prevented seating of the sterile adapter onto the psm. The psm is an instrument arm located on the patient side cart, which provides the sterile interface for the endowrist instruments. The sterile adapter release tab is located on the carriage of the psm arm and holds the sterile adapter securely in place on the psm. The sterile adapter provides the isolation barriers to electrical current. The system was repaired by replacing the affected psm. The psm was returned to isi for failure analysis investigation. Engineering confirmed the customer reported problem and found that the wrist output assembly post heights of axis 4 through axis 7 on the psm were out of specification and the link 6 on the release tab was broken, thus causing the issue experienced by the site. The psm was repaired by replacing the affected output assembly posts and release tab. Engineering evaluation also found that four of the inner potentiometers were out of specification for jitter, but were otherwise functioning correctly. The affected potentiometers were also replaced while repairing the psm. On (b)(4) 2013, isi contacted the isi clinical sales representative (csr) concerning the reported event. The csr advised that the site indicated to him that the ureter injury sustained by the patient during the da vinci s hysterectomy procedure was unrelated to any malfunction of the da vinci s surgical system, instrument and/or accessories, but rather an error by the surgeon occurred. The csr also indicated that there was no report by the site that any malfunction of the da vinci s surgical system, instrument and/or accessories occurred during the hysterectomy surgical procedure that caused or contributed to the injury sustained by the patient. No other information concerning the reported event was provided. As of (b)(4) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2964607 | 2939795 | 2836313 | 2955842201300333 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2013-01-04 | 2013-02-01 | Yes | Yes | Health Professional | IS2000 A51P7 | Device Returned To Manufacturer | No | 2013-01-04 | Device Not Retu | No | Reuse | 2013
IS 2000 A51P8 | 2013-01-07 | Other | Other | It was reported that during a da vinci s hysterectomy procedure, the surgeon experienced articulation issues while manipulating an instrument installed on the 2 patient side manipulator (psm) arm. The site contacted isi for technical support assistance. The site indicated to the technical support engineer (tse) that they reseated the sterile adapter; however, the issue persisted. Prior to contacting isi for technical support assistance, the surgeon made the decision to complete the planned surgical procedure using open surgical techniques. No patient harm was reported. | The investigation conducted by the field service engineer (fse) was unable to replicate the customer reported failure mode. Functional testing of the 2 patient side manipulator (psm) arm in conjunction with the sterile adapter used during the surgical procedure found no issues with the sterile adapter. Functional testing of the left master tool manipulator (mtm) in conjunction with the psm found that the left mtm functioned properly. Inspection of the psm's cable tensions and internal gears found that they were also within specification. The patient side manipulator is an instrument arm located on the patient side cart, which provides the sterile interface for the endowrist instruments. The master tool manipulator refers to the master controllers which provide the means for the surgeon to control the instruments and endoscope inside the patient from the surgeon's side console. One mtm is assigned to the surgeon's left hand (mtml) and one to his right (mtmr). The sterile adapter provides the isolation barriers to electrical current. While the fse was unable to replicate the issue experienced by the site, psm 3 and psm 4 were swapped and the left mtm was re-calibrated as a precaution. A sine cycle performed on psm 2, psm 3, psm 4 and the left mtm found that the site's da vinci s surgical system functioned within specification. A sine cycle is a test procedure in which the psm arm is automatically driven back and forth repeatedly while monitoring sensor data and error states. On (b)(6) 2013 isi contacted the initial reporter of this incident. The initial reporter indicated that the surgeon's decision to complete the planned surgical procedure using open surgical techniques was due to a difficult patient anatomy and the articulation issue experienced by the surgeon. The initial reporter indicated that the patient tolerated the open procedure well, there was no harm to the patient and the patient did not experience any surgical complications as a result of the reported event. As of (b)(6) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2964662 | 2939813 | 2836313 | 2955842201300334 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-07 | 2013-02-01 | Yes | Yes | Health Professional | IS 2000 A51P8 | Yes | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-02 | Malfunction |  | It was reported that during a da vinci si surgical procedure, it was noted that the wire of the large needle driver instrument was sticking out at the pivot point. No missing or fallen fragments have been reported by the site. There was no patient harm, injury or adverse outcome reported. | The instrument was returned and evaluated. Engineering observed that the grip cable is frayed at the distal idler pulley. The frayed strands stick out at the wrist. Other cables at wrist are not damaged. The instrument has 0 use left. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2964753 | 2939862 | 2836380 | 2955842201300335 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-03 | 2013-02-01 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42018306 | 2012-12-21 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tip of the permanent cautery hook instrument was cracked. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the instrument was returned with a large section of the ceramic sleeve broken off. The entire distal end of the sleeve is broken off and missing. The section of sleeve that is intact was cracked. Additional observation not reported by the site is a bent hook. The hook is bent near to where it comes out of the yaw pulley. The direction of applied load is perpendicular to side of the hook. Evidence was not conclusive, but damage to the sleeve and hook is likely due to mishandling/misuse. There were 7 uses left. No other damage was found. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK | ENDOSCOPIC ELECTROSURGICAL INS | 2964755 | 2939863 | 2836381 | 2955842201300336 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-03 | 2013-02-01 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-27 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the wire on the maryland bipolar forceps instrument was observed to be broken. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode. However for clarification, the nonconformance was a frayed pitch cable. The frayed segments are in various lengths and stick out at the wrist. No damage or wear marks were found at the clevis. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2964744 | 2939865 | 2836383 | 2955842201300338 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-04 | 2013-02-01 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2013-01-04 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-02 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the fenestrated bipolar forceps instrument was not recognized by the system. There were no missing or fallen pieces reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode, as the instrument passed recognition testing when installed on an in-house isi is3000 system. Engineering evaluation also found that the instrument's pitch was broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis and the clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2964749 | 2939870 | 2836408 | 2955842201300339 | 1 | NAY | Manufacturer | User facility | Health Professional | Initial | 2013-01-07 | 2013-02-01 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-13 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, a frayed wire was noted on the pk dissecting forceps instrument. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a frayed cable on the instrument. The instrument was found with a frayed pitch cable at distal idler. No damage was found on the pulley and clevis. Electrical continuity passed. The instrument has 4 use remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2964747 | 2939879 | 2836395 | 2955842201300337 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-04 | 2013-02-01 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-04 | Device Not Retu | No | Reuse | 2013
42004906 | 2013-01-04 | Malfunction |  | It was reported that during reprocessing of the cadiere forceps instrument, a wire at the distal end of the instrument was observed to be broken. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The customer initially reported a broken wire; however, for clarification, the damaged instrument component was a frayed grip cable. Based on this clarification, the customer reported complaint is confirmed. The grip cable is frayed at the distal idler pulley and frayed cable strands stick out at the wrist. The other cables at the wrist are not damaged. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and are not aligned with the tube axis, suggesting that the scratches may have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2964771 | 2939892 | 2836428 | 2955842201300340 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-04 | 2013-02-01 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-01-04 | Device Not Retu | No | Reuse | 2013
42018406 | 2012-12-31 | Malfunction |  | Event Description: --/-- It was observed during central processing that the permanent cautery spatula instrument had a cracked insulation issue. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. On (b)(4) 2013, the isi regulatory complaints analyst contacted the site and was able to obtain/verify the following information: the reporter stated that no fragments fell and that no patient was harmed/injured because of the alleged issue. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the cracked insulation was confirmed. The distal yaw pulley was missing a piece from the ceramic insulation, which was close to the grip tip. There was no sign of charring on the tip. There was no information given if missing piece fell into patient or not. Engineering concluded that the damage was likely due to mishandling. Engineering also found the conductor wire was dislodged. The conductor wire was sticking out from the conductor wire cap around the distal area. The evidence was not conclusive, but the wire probably slipped out while instrument was driven on system. The length of the wire sticking from distal end at approximately. 1160. The electrical continuity testing passed. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULAINSTRUMENT | Reuse | 2961940 | 2938056 | 2832689 | 2955842201300325 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2013-01-02 | 2013-01-31 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2013-01-02 | Device Not Retu | No |  | 2013
42020505 | 2013-01-02 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a cable broke on a fenestrated bipolar forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found the damaged instrument component was a frayed pitch cable. The instrument was found with a frayed pitch cable at the distal idler. No damage was found on the pulley and the clevis. The electrical continuity testing passed. Engineering also found various scratch marks on the distal end of the main tube showing light material removal. The scratches were short in length and not axially aligned with the tube. Engineering concluded that damage was likely due to mishandling. The bottom bipolar pin was also bent towards the top pin. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | Reuse | 2961970 | 2938061 | 2834577 | 2955842201300326 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2013-03-20 | 2013-01-31 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-03-20 | Device Not Retu | No |  | 2013
42020504 | 2013-01-01 | Malfunction |  | It was observed during central processing that the fenestrated bipolar forceps instrument had a broken wire. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The alleged complaint of a broken wire was confirmed. The grip cable was found to be loose at the distal idler and the actual breakage was observed inside the housing by the clamping pulley. The clamping pulley showed some residue collecting around where the wires were wrapped and around the breakage area. Evidence is not conclusive but the damage was most likely due to improper cleaning. Electrical continuity testing passed. 48 - an additional observation not reported by the site was a loose conductor wire. The conductor wire was found to be loose in the same area as the grip cable but no breakage was observed. Another finding was scratches on the instrument's main tube. The distal end of the main tube had various scratch marks showing light material removal and a rough surface finish. The scratches were short in length and not axially aligned with the tube. Evidence is not conclusive, but the damage might have been due to mishandling. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2961989 | 2938066 | 2834582 | 2955842201300327 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-02 | 2013-01-31 | No | Yes | Health Professional | 42020504 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci si gynecological surgical procedure, the surgical staff allegedly found a broken frayed cable on the prograsp forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint of a broken frayed cable was confirmed. The pitch cable was found to be broken at the distal clevis hub. The cable segment that contained the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. An additional finding not reported by the site was scratches on the instrument's main tube. The distal end of the main tube had various scratch marks with light material removal. The scratches were short in length and not aligned with the tube axis; suggesting the damage might have been caused by instrument collisions or rough handling. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2962114 | 2938144 | 2834660 | 2955842201300328 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-02 | 2013-01-31 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-02 | Malfunction |  | It was observed during central processing that the fenestrated bipolar forceps instrument had a broken wire at the wrist. There were no missing or fallen pieces reported. There was no allegation of any harm, injury, or adverse outcome to a patient. | The instrument was returned and evaluated. The customer initially reported a broken wire issue. However for clarification, there was a broken pitch cable. Based on this, the customer's alleged complaint was confirmed. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2962150 | 2938180 | 2834696 | 2955842201300330 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-02 | 2013-01-31 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42034301 | 2012-12-31 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si thoracotomy procedure, the insulation of the thoracic grasper instrument was noted to be burned/compromised. No adverse event, patient harm or injury was reported. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. | Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering evaluation results confirmed the alleged issue of compromised insulation. The main tube insulation exhibited damage at the distal end. The insulation was found with a gouge mark with material missing. The mark was not axially aligned with the tube. In addition, the main tube also exhibited a couple of different damaged tube insulation sections with deep scratches but no material removed. Engineering evaluation concluded that the damage might have been due to misuse or mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported main tube scratch issues if to recur could cause or contribute to an adverse event. | THORACIC GRASPER | Unkown | 2962526 | 2938402 | 2834919 | 2955842201300331 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | InitialFollowup | 2013-02-05 | 2013-01-31 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | Yes | 2013-02-05 | Device Not Retu | No |  | 2013
410322 | 2013-01-30 | Injury | Required Intervention | The blade portion of the endo wrist one vessel sealer device was unable to divide the tissue. The sealer portion of the device worked fine. |  | ENDO WRIST ONE VESSEL SEALER | VESSEL SEALER | 2974004 | 2948859 | 2845375 | MW5028845 | 1 | NAY | Voluntary |  | Nurse | Initial | 2013-01-31 | 2013-01-31 | No | Yes | Health Professional | 410322 | Yes | Yes |  |  |  |  | 2013
42030901 | 2013-01-02 | Malfunction |  | It was reported that during set up of a da vinci si surgical procedure, it was noted that the cable on the megasuturecut needle driver instrument was broken. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering that the instrument was returned with a frayed grip cable located at the distal idlers. The frayed cable sections are in various lengths sticking out. The grip cable also became derailed as well. The idler pulley showed pulley face and rim damage. Engineering evaluation also found that the insulation at the distal end of the main tube exhibits missing material that appears to have been scraped off. On (b)(4) 2013, isi contacted the o. R. Coordinator at the hospital concerning the condition of the returned instrument. The o. R. Coordinator indicated that there was no report by the surgical staff that any piece from the returned affected instrument fell into a patient during a surgical procedure. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2959624 | 2935649 | 2832167 | 2955842201300314 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-02 | 2013-01-30 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42014204 | 2013-01-02 | Malfunction |  | It was reported that during reprocessing of the monopolar cautery instrument it was noted that the insulation on the instrument was peeling off. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with gouge marks and material removal at the distal end of the main tube. Engineering concluded that the damage was likely due to excess force contact on the main tube surface. 62 - the instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2959761 | 2935698 | 2832214 | 2955842201300315 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-02 | 2013-01-30 | No | Yes | Health Professional | 42014204 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
420177 | 2012-12-10 | Malfunction |  | It was reported that during a davinci si colectomy procedure, the customer noted loss of insulation, manipulation of tip on the bowel grasper instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | BOWEL GRASPER | ENDOSCOPIC INSTRUMENT | 2959868 | 2935834 | 2832350 | 2955842201300316 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-04 | 2013-01-30 | No | Yes | Health Professional | 420177 | No | No | 2013-01-04 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-27 | Malfunction |  | It was reported that during a davinci s thyroidectomy procedure, the customer noted a broken wire on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable at distal clevis hub. There was no damage found at the clevis. The frayed segment was approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2960015 | 2935915 | 2832433 | 2955842201300317 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-04 | 2013-01-30 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2013-01-04 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-27 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s surgical procedure, the wires at the tip of the megasuturecut needle driver instrument were observed to be broken and frayed. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idler pulley. The idler pulley spun freelyp; however, it had damages on the edge and on the surface. The cable segment was sticking out at the wrist. The other cables at the wrist were not damaged. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | Reuse | 2960378 | 2936120 | 2832638 | 2955842201300318 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-01-31 | 2013-01-30 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-31 | Yes | No |  | 2013
42018406 | 2012-12-31 | Malfunction |  | Event Description: --/-- It was reported that during central processing the permanent cautery spatula instrument was noted to have cracked insulation. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the nonconformance was a broken ceramic sleeve. An uneven piece of the ceramic sleeve was missing, exposing the shaft of the spatula. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | Initial | 2960408 | 2936171 | 2832689 | 2955842201300319 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | FollowupInitial | 2013-02-05 | 2013-01-30 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2013-02-05 | Device Not Retu | No |  | 2013
42022703 | 2012-12-31 | Malfunction |  | It was reported that during a da vinci si procedure the pk dissecting forceps instrument was noted to have a wire sticking out at the tip of the instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed using an instrument of the same type. No patient harm adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a wire sticking out at the tip of the instrument. The pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp remained in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2960444 | 2936202 | 2832700 | 2955842201300320 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-03 | 2013-01-30 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-17 | Malfunction |  | It was reported during a da vinci s hysterectomy surgical procedure that the fenestrated bipolar forceps instrument tips were not aligned. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the one grip is bent, causing side to side misalignment of grips. There is a 0. 029 offset at the tips. The grips do not meet together once the grips are closed. Electrical continuity passed. Engineering concluded that damage was likely due to mishandling. Engineering also observed a frayed pitch cable. The frayed pitch cable was found at the distal clevis hub. No damage was found at the clevis. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2960515 | 2936231 | 2832747 | 2955842201300321 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-16 | 2013-01-30 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-16 | Device Not Retu | No | Reuse | 2013
42000606 | 2013-01-04 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted frayed wires on the large needle driver instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable located at distal clevis hub. There was no damage was found at the clevis. No other cable damage was found. Additional observation not reported by site was main insulation tube damage. The distal end of main tube exhibited various scratch marks with light material removal. The scratches were short in light and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2960531 | 2936270 | 2832808 | 2955842201300322 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-04 | 2013-01-30 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-04 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-01-03 | Malfunction |  | It was reported that during reprocessing of the thoracic grasper instrument, the insulation on the instrument was observed to be torn. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering found that the instrument was returned with the black tube insulation damaged at the distal end. A 1. 4 long section at the distal end of the tube had a gouge in the insulation with a. 075 x. 040 piece missing and two larger sections of the insulation was lifted up. Engineering evaluation also found that the tube is bent roughly 2 above the snake wrist. The distal end roll motion is not aligned with proximal end of tube due to bending. Engineering concluded that the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2960621 | 2936343 | 2832861 | 2955842201300323 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-03 | 2013-01-30 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-04 | Malfunction |  | It was reported that during reprocessing of the megasuturecut needle driver instrument the cable on the instrument snapped. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with one grip close cable broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment stuck out at the wrist. The other cables at wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2960626 | 2936393 | 2832909 | 2955842201300324 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-04 | 2013-01-30 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-04 | Device Not Retu | No | Reuse | 2013
420309 | 2013-01-07 | Injury | Required Intervention | Robotic instrument was noted to have broken during laparoscopic robotic assisted surgery, wires from the instrument were seen frayed inside pt and a small piece of metal wire was removed by surgeon. An x-ray was performed at the end of the case to confirm no other foreign material was left inside. No foreign item seen. |  | DAVINCI INTUITIVE SURGICAL | MEGA SUTURECUT NEEDLE DRIVER | 2972387 | 2945961 | 2842477 | 2945961 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-01-25 | 2013-01-30 | No | Yes | Health Professional | 420309 | Yes | Yes |  |  |  |  | 2013
IS3000 A60P6 | 2012-04-02 | Injury | Other | On (b)(6) 2012, a patient posted the following information on isi's social network site via (b)(6): "i had the davinci surgery 8 months ago. I still have some numbness in one area and i still have some pain. Do you know why i would still have pain after this long? i will be seeing my doctor on (b)(6). I have asked her before why it still hurt; she told me since i had problems with my left side before my surgery that is why it still hurts. It does not make sense at all to me. And also about a week ago i started to spot. It lasted for a day and a half. ". | Isi has contacted the operating surgeon at this facility, and she indicated that she does not recall any malfunction of the da vinci system, instruments and/or accessories occurring during the surgical procedure. The surgeon indicated that the patient was positioned in the trendelenburg position. The surgeon also indicated that she has not received any report that the patient experienced any post surgical complications despite the patient's report that she was in the hospital for a week. The doctor also stated that she is unable to provide any other details concerning the patient without written consent from the patient. Furthermore, on (b)(6) 2013, the patient was contacted and provided more insight on the reported complaint. It was stated that prior to having the scheduled hysterectomy procedure, she experienced heavy menstrual cycles. At the present time, she feels numbness on her left side where the surgical incision was. During her last visit with the surgeon, it was stated that it could be an infection or it could be the patient's bowels. The patient then went to her primary care physician (pcp), who performed an ultrasound. The ultrasound results came back fine. The pcp prescribed some pain medication and recommended more fiber in her diet. No further information was provided. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2958285 | 2933637 | 2830153 | 2955842201300309 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-31 | 2013-01-29 | Yes | No | Health Professional | IS3000 A60P6 | No | No | 2012-12-31 | No | No | Reuse | 2013
42022703 | 2012-12-31 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff allegedly found a frayed cable on the pk dissecting forceps instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint of a frayed instrument cable was confirmed. For clarification, the instrument was found to have a frayed and broken cable at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2958401 | 2933688 | 2830206 | 2955842201300310 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-02 | 2013-01-29 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-05 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the fenestrated bipolar forceps instrument was not grasping properly. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode based on the condition of the returned instrument. Engineering evaluation found that the main tube is broken and remains attached one and 3/4 inches from the proximal end of the instrument by the internal wiring. The angle of the damage on the instrument prevented the grip from grasping. Engineering concluded that the damage to the instrument likely occurred outside of a cannula and that the damage likely occurred as a result of mishandling. The instrument also passed electrical continuity testing. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported failure mode does not constitute a mdr reportable event; however; the damage to the instrument's main tube found during failure analysis, could likely cause or contribute to an adverse event if the malfunction were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2958658 | 2933956 | 2830454 | 2955842201300311 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-02 | 2013-01-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the jaws on the pk dissecting forceps instrument would not open. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as functional testing of the instrument showed that the grips opened and closed properly. The instrument passed electrical continuity testing. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and are not aligned with the tube axis, suggesting the damage may have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported failure mode does not constitute a mdr reportable event; however, damage to the instrument's main tube found during failure analysis, could likely cause or contribute to an adverse event if the malfunction were to recur. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2958672 | 2933963 | 2830461 | 2955842201300312 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-02 | 2013-01-29 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42020505 | 2013-01-02 | Malfunction |  | It was reported that during reprocessing of the fenestrated bipolar forceps instrument the cable on the instrument was observed to be broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the pitch up cable broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The other cables at the wrist are not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2958691 | 2933970 | 2830488 | 2955842201300313 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2013-01-02 | 2013-01-29 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
| 2013-01-14 | Malfunction |  | The robotic grasper would not let go of the tissue. We used the emergency allen wrench to release it, but it still would not release. We had to do a total system shutdown to get the grasper to open its jaws. There was no patient injury. What was the original intended procedure?robotically assisted colectomy and rectopexy. |  | DA VINCI ENDOWRIST GRAPTOR | SYSTEM SURGICAL COMPUTER | 2963272 | 2938981 | 2835497 | 2938981 | 1 | NAY | User Facility |  | Nurse | Initial | 2013-01-29 | 2013-01-29 | No | Yes | Physician |  | Yes | Yes |  |  |  |  | 2013
42020706 | 2013-01-07 | Malfunction |  | It was reported that after completion of a da vinci si gynecological surgical procedure, it was observed that wires on the tenaculum forceps instrument were sticking out. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode; however, for clarification, the damaged instrument component was a broken grip cable. The grip cable at the distal idler pulley is broken. The idler pulley was observed to spin freely and cable segments were found to be sticking out at the wrist. Engineering evaluation also found that the distal end of the main tube exhibits insulation damage. Material was found to be missing. The surface of the main tube appeared to be scraped off. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2956963 | 2931915 | 2828433 | 2955842201300306 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-08 | 2013-01-28 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-04 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the antennas on the maryland bipolar forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering confirmed the customer reported failure mode as the top bipolar pin is pushed into the back end. Engineering concluded that the damage is likely due to mishandling. The instrument passed electrical continuity testing. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and are not aligned with the tube axis, suggesting the damage was caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The failure analysis finding does not constitute a mdr reportable event; however; the damage to the instrument's main tube, could likely cause or contribute to an adverse event if the malfunction were to recur. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2957112 | 2932047 | 2792389 | 2955842201300307 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-02 | 2013-01-28 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42018406 | 2012-12-05 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the plastic tip of the cover on the permanent cautery spatula seemed to be loose. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument's ceramic sleeve was found broken at the base of the spatula. Engineering concluded that the damage was likely due to excessive force contact on the sleeve. The instrument also passed electrical continuity testing. No other physical damage was found. The instruments and accessories user manual specifically states: general precautions and warnings, handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2957130 | 2932085 | 2820000 | 2955842201300308 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-02 | 2013-01-28 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
VER04 | 2013-01-18 | Injury | Required Intervention | During a robotic-assisted right middle lobectomy, the graphite housing of robotic forceps instrument split and the instrument was stuck laterally and the trocar had to be removed from the pt to get the instrument out of the trocar. No injury to pt. |  | CADIERE FORCEPS | ROBOTIC FORCEPS | 2968529 | 2941470 | 2837986 | MW5028812 | 1 | GCJ | Voluntary |  | RISK MANAGER | Initial | 2013-01-28 | 2013-01-28 | No | Yes | Health Professional | VER04 | Yes | Yes |  |  |  |  | 2013
42034301 | 2012-11-30 | Malfunction |  | It was reported that after a da vinci si surgical procedure, the surgical staff reported seeing shavings from the shaft of the thoracic grasper instrument, when trying to remove it from the cannula. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of shavings from shaft is confirmed. Black tube insulation is damaged at the distal end. Insulation is missing a 6 long section all around the shaft starting directly above the snake wrist. Various strips of clear plastic that covers the metal shaft are sticking out at the distal end. Insulation has various circular shaped sections missing that are all axially aligned located above the midpoint of the tube. Evidence not conclusive, but damage is likely due to mishandling/misuse. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2955320 | 2930195 | 2826711 | 2955842201300291 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-27 | 2013-01-25 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
40027202 | 2012-12-11 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported the tip of the harmonic ace curved shears instrument used in conjunction with the harmonic ace curved shears insert accessory had a dent. No patient harm, adverse outcome or injury was reported. | The insert accessory was returned and evaluated. The harmonic ace curved shears instrument is used in conjunction with the harmonic ace curved shears insert accessory. Investigation by failure analysis showed a missing teflon pad and broken round rack feature. Insert was returned installed on the harmonic ace curved shears instrument. Insert is stuck in instrument and cannot be removed. Visual inspection shows the entire teflon pad is missing from the clamp arm. No damage observed to the clamp arm groove that holds pad. At the proximal end of the insert, the round rack feature is broken around the retaining ring groove. Retaining ring is detached from insert, making insert removal from instrument difficult. Evidence not conclusive, but damage to insert is likely due to mishandling/misuse. No other damage found. The harmonic ace curved shears instrument was returned and evaluated. Customer initially reported tip has a dent. However, for clarification, the nonconformance was a damaged shaft. Shaft of the harmonic ace curved shear instrument exhibited numerous dents on the shaft. Instrument was returned with the ace insert also still installed with instrument. Ace insert cannot be removed from instrument. Insert was likely installed incorrectly with excessive force, causing binding of the insert and instrument. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The da vinci harmonic ace curved shears instructions for use specifically states: inspection prior to use examine the device prior to use. In particular, examine the following components for cracks or flaws: the clamp arm and blade on the insert - the shaft of the insert - the reusable instrument housing caution: if cracks or flaws are observed, do not use the device. Note: scratches on the blade may lead to premature blade failure. Avoid accidental contact with other instruments during use. Do not use any means other than the blade wrench to attach or detach the instrument from the hand piece. General precautions and warnings - do not attempt to bend, sharpen, or otherwise alter the shape of the blade. Doing so may cause blade failure and user or patient injury. Avoid contact with any and all metal or plastic instruments or objects when the device is activated. Contact with staples, clips, or other instruments while the device is activated may result in cracked or broken blades, which may be identified by a generator solid tone or an instrument error. Care should be taken not to apply pressure between the blade and clamp arm without having tissue between them. This can result in possible damage to the instrument. Both conditions may cause a system failure signaled by a continuous beep when either of the foot pedals is depressed. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2955322 | 2930196 | 2826712 | 2955842201300289 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-27 | 2013-01-25 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Initial | 2013
42020906 | 2012-11-28 | Malfunction |  | It was reported that a mega suturecut needle driver instrument had loose wiring. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer initially reported a loose wire issue. However, for clarification the damaged instrument component was a broken grip cable. Based on this clarification, the customer reported complaint is confirmed. One grip close cable was found to be broken at the distal idlers. The idler pulley spun freely and was not damaged. A cable segment was found to be sticking out at the wrist. Other cables at the wrist were not damaged. Additional finding: the instrument was found to have broken housing snaps. Four of four housing snap tabs and three adjacent housing pins were broken. The housing was able to be removed as a result of the damage. Instrument may have been mishandled. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2955325 | 2930199 | 2826715 | 2955842201300294 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-28 | 2013-01-25 | No | Yes | Health Professional | 42020906 | Device Returned To Manufacturer | No | 2012-12-28 | Device Not Retu | No | Reuse | 2013
42018406 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff reported seeing energy leaking from the permanent cautery spatula instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaining of energy is leaking was confirmed. Instrument was found with char mark at the base of yaw pulley within the distal clevis. Melted material was found on the pulley. No sign of arc track outside of the clevis ear. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2955326 | 2930200 | 2826716 | 2955842201300293 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-28 | 2013-01-25 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-12-28 | Device Not Retu | No | Reuse | 2013
42000305 | 2012-11-29 | Malfunction |  | It was reported that during a da vinci si thymectomy procedure, the surgical staff reported, that between loads while the instrument was not in the patient, the ends of the small clip applier instrument broke off. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The complaint that the ends of the instrument broke off was confirmed. Both clip applier grips were broken at the midpoint. Evidence is not conclusive; however, the damage was likely due to excess force applied normal to the grip while trying to install a clip. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2955341 | 2930201 | 2826717 | 2955842201300295 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-29 | 2013-01-25 | No | Yes | Health Professional | 42000305 | Device Returned To Manufacturer | No | 2012-12-29 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-02 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff reported seeing a broken cable. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer initially reported broken wire; however, for clarification the damaged instrument component was a frayed grip cable. Based on this clarification, the customer's reported complaint is confirmed. The grip cable was frayed at the distal idler pulley. Frayed strands were observed to be sticking out at the wrist. Other cables at the wrist were not damaged. Additional finding: the grip cable was also found to be derailed at the distal idler pulley. The cable derailment is likely due to cable losing contact with pulley during wrist articulation. The fleet angle between proximal and distal pulleys may have contributed to the derailment. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2955328 | 2930205 | 2826721 | 2955842201300298 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-03 | 2013-01-25 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-24 | Malfunction |  | It was observed during central processing that the mega suture cut needle driver instrument had a frayed wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer initially reported a frayed wire, however, for clarification the damaged instrument component was a broken grip cable. Based on this clarification, the customer reported complaint is confirmed. One grip close cable is broken at the distal idlers. Idler pulley spins freely and was not damaged. Cable segment sticks out at wrist. Other cables at wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2955313 | 2930212 | 2826728 | 2955842201300286 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-25 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Initial | 2013
42019408 | 2012-12-27 | Malfunction |  | On (b)(4) 2012, the mega needle driver instrument was returned without a complaint filed against it. The failure analysis team determined the instrument had a broken cable. No information was provided. | The instrument was returned and evaluated. No reason for return was provided. A complaint cannot be confirmed or denied. Additional observation not reported by site is a broken cable. One grip close cable is broken near the proximal pulleys and prox clevis cable hole. Idler pulley spins freely and does not exhibit damage. Cable segment sticks out at wrist. Prox clevis cable hole exhibits wear on one edge. Other cables at wrist are not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2955289 | 2930213 | 2826729 | 2955842201300287 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-27 | 2013-01-25 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-26 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the surgical staff reported seeing the wires on the mega suturecut needle driver instrument snap when placed through the cuff. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of instrument wires snapped was confirmed. One grip close cable is broken at the distal idlers pulley. Idler pulley spins freely but has damages on the edge and on the surface. Cable segment sticks out at wrist. The other grip cable was derailed due to lost of tension, but not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2955314 | 2930214 | 2826730 | 2955842201300285 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-25 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-27 | Malfunction |  | On (b)(4) 2012 the pk dissecting forceps instrument was returned without a complaint filed against it. The failure analysis team determined the instrument had a broken cable. No information was provided. | The instrument was returned and evaluated. No reason for return was provided. A complaint cannot be confirmed or denied. Additional observation not reported by site is a broken cable. The pitch down cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Other cables at wrist are not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2955318 | 2930219 | 2826735 | 2955842201300288 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-27 | 2013-01-25 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
42000306 | 2012-12-10 | Malfunction |  | It was observed during central processing that the small clip applier instrument had a loose wire. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer initially reported loose wire, however, for clarification the damaged instrument component was a broken grip cable. Based on this clarification, the customer reported complaint is confirmed. Grip cable is broken at the distal idlers. Idler pulley spins freely and was not damaged. Cable segment sticks out at wrist loose. The location where broken wire occured is inside the housing the clamping pulley is at. Other cables at wrist are not damaged. Additional finding: scratches and indentations on distal pulleys. Both pulleys have indentations on the edges and some small scratches on the surface. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2955334 | 2930221 | 2826737 | 2955842201300290 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-27 | 2013-01-25 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-28 | Malfunction |  | It was observed during central processing that the fenestrated bipolar instrument had frayed wires on the tip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint wires were frayed is confirmed. Instrument was found with frayed pitch cable at distal clevis hub. No damage was found at the clevis. No other cable damage was found. Electrical continuity passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2955337 | 2930225 | 2826741 | 2955842201300292 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-28 | 2013-01-25 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-28 | Device Not Retu | No | Reuse | 2013
42034301 | 2013-01-03 | Malfunction |  | It was observed during central processing that the insulation of the thoracic grasper instrument was torn. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The complaint of the torn insulation was confirmed. The black tube insulation was damaged at the distal end. The insulation was gouged on one side and had a. 060 x. 045 piece missing roughly 1. 8 above the snake wrist. The insulation also had various scratch marks with no material missing below the damaged section. Additional observation not reported by site was a bent tube. The tube was slightly bent roughly 6 above the snake wrist. The distal end roll motion was not aligned with proximal end of tube due to bending. Evidence is not conclusive, however, the damage is likely due to mishandling/misuse. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2955344 | 2930227 | 2826743 | 2955842201300297 | 1 | GEI | Manufacturer | User facility | Other | Initial | 2013-01-03 | 2013-01-25 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42030901 | 2013-01-02 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff reported seeing a broken cable. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The alleged complaint of a broken cable was confirmed. One grip close cable was found to be broken at the distal idlers. The idler pulley was observed to spin freely and did not exhibit damage. The cable segment was found to be sticking out at the wrist. Other cables at the wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2955347 | 2930229 | 2826721 | 2955842201300299 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-03 | 2013-01-25 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42004906 | 2012-12-31 | Malfunction |  | It was observed during a da vinci si surgical procedure that a cable snapped on a cadiere forceps instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The complaint that the cable snapped is confirmed. One grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. A cable segment sticks out at wrist. Other cables at the wrist are not damaged. Additional finding: deep scratches are observed on the maintube. The distal end of main tube has various scratch marks with light material removal. The scratches are short in length and not aligned with the tube axis, suggesting the may have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS | ENDOSCOPIC INSTRUMENT | 2955438 | 2930298 | 2826814 | 2955842201300296 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-03 | 2013-01-25 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-19 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff reported seeing frayed wires at the distal end of a megasuture needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer initially reported a frayed wire; however, for clarification the damaged instrument component was a broken grip cable. Based on this clarification, the customer reported complaint is confirmed. One grip close cable was broken at the distal idlers. The idler pulley was found to spin freely and was not damaged. A cable segment was observed to be sticking out at the wrist. Other cables at the wrist were not damaged. Additional finding: scratches were observed on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2955442 | 2930302 | 2826818 | 2955842201300300 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-03 | 2013-01-25 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42019408 | 2012-12-20 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure, the surgical staff alleged that a cable broke on a mega needle driver instrument. The surgical staff claimed that a wire was frayed and sticking out. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no pt harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The customer initially reported a frayed wire; however, for clarification the damaged instrument component was a broken grip cable. Based on this clarification, the customer's reported complaint is confirmed. One grip close cable was found to be broken at the distal idlers. The idler pulley was observed as being able to spin freely and was not damaged. A cable segment was found to be sticking out at the wrist. Other cables at the wrist were not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 3008512 | 2956865 | 2853383 | 2956865 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2012-12-28 | 2013-01-25 | No | Yes | Health Professional | 42019408 | Yes | No | 2012-12-28 | Yes | No | Reuse | 2013
42014204 | 2012-12-10 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff alleged that the "casing" of a monopolar cautery instrument was torn. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no pt harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint of instrument's casing being torn was confirmed. The instrument's main tube insulation appeared to have scratches and gouge marks with material removed at the distal end. Evidence is not conclusive, but the damage was likely caused by excess force contact on the main tube surface. The instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Per the customer, at the time of the reported event, nothing fell into a pt during a surgical procedure. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3013323 | 2960685 | 2857201 | 2955842201300305 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-28 | 2013-01-25 | No | Yes | Health Professional | 42014204 | Yes | Yes | 2012-12-28 | Yes | No | Reuse | 2013
42034301 | 2012-12-05 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff alleged that the thoracic grasper instrument was not grasping. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no pt harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint of the instrument not grasping could not be confirmed. The instrument was placed on an is3000 system and driven. Recognition and engagement of the instrument passed. The instrument moved intuitively with full range of motion in all directions. The instrument grips opened and closed properly. Functional performance testing did not find any issues. An additional observation not reported by the site was tube damage. The black tube insulation was found to be damaged at the distal end. The tube insulation was gouged on one side and had a. 100" x. 030" piece missing roughly 3. 35" above the snake wrist. There were also other areas at the distal end where the insulation was scratched with no material removal. Evidence is not conclusive, but the damage was likely due to mishandling/misuse. No other damage was found. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT GEI | 3013252 | 2960686 | 2857202 | 2955842201300301 | 1 | GEI | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-27 | 2013-01-25 | No | Yes | Health Professional | 42034301 | Yes | Yes | 2012-12-27 | Yes | No | Reuse | 2013
40020708 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci standard total benign hysterectomy procedure, the surgical staff alleged that the tip of a tenaculum forceps instrument would not straighten. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no pt harm, adverse outcome, or injury was reported. | The instrument was returned and evaluated. The alleged complaint that the instrument's tip would not straighten could not be confirmed. The drive cables at the instrument's wrist were intact and undamaged. The wrist was straight when input disc dogs were all aligned horizontally. The grips were able to open/close fine when input discs were manually rotated. An additional observation that was not reported by the site was tube damage. The distal end of the main tube had various scratch marks with light material removal. The scratches were. 070" -. 300" in length and not aligned with the tube axis. Evidence is not conclusive, but the damage might have been caused by mishandling/misuse. Per the customer, at the time of the reported event, nothing fell into a pt during a surgical procedure. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 3013300 | 2960715 | 2857234 | 2955842201300304 | 1 | NAY | Manufacturer | User facility | UNKNOWN | Initial | 2012-12-28 | 2013-01-25 | No | Yes | Health Professional | 40020708 | Yes | No | 2012-12-28 | Yes | No | Reuse | 2013
42019408 | 2012-12-14 | Malfunction |  | It was reported that during a da vinci s total benign hysterectomy procedure, the surgical staff alleged that a cable snapped on a mega needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no pt harm, adverse outcome, or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT NAY | 3013278 | 2960716 | 2857235 | 2955842201300302 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-28 | 2013-01-25 | No | Yes | Health Professional | 42019408 | No | Yes | 2012-12-28 | Device Not Retu | No | Unkown | 2013
42030901 | 2012-12-31 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a broken cable on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley spins freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional finding not reported by the site was scratches on the surface and the edge of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2952697 | 2927637 | 2824153 | 2955842201300275 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-02 | 2013-01-24 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42029602 | 2012-11-30 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a cable snapped on a large suturecut needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The complaint that the cable snapped was confirmed. One grip close cable was broken at the distal idlers. The idler pulley was observed to spin freely. A cable segment was found to be sticking out at the wrist. Other cables at the wrist were not damaged. An additional observation not reported by the site was a damaged pulley and idler pin. The distal idler pulley on which the broken cable was seated had various scratch marks on the outer surface. The idler pin exhibited wearing on one side. Evidence is not conclusive, but the grip close cable may have become derailed and this likely contributed to cable breakage. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2953351 | 2928381 | 2824897 | 2955842201300276 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-03 | 2013-01-24 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | Yes | 2013-01-10 | Device Not Retu | No | Reuse | 2013
42029602 | 2012-12-05 | Malfunction |  | It was reported that during a da vinci si total benign hysterectomy procedure, the surgical staff alleged that a cable snapped on a large suturecut needle driver instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The complaint that the cable snapped is confirmed. One grip close cable was broken at the distal idlers. The idler pulley was observed to spin freely. A cable segment was found to be sticking out at the wrist. Other cables at the wrist were not damaged. An additional observation not reported by the site was a damaged pulley and idler pin. The distal idler pulley on which the broken cable was seated had various scratch marks on the outer surface. The idler pin exhibited wearing on one side. Evidence was not conclusive; but the grip close cable may have become derailed and this likely contributed to the cable breakage. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2953633 | 2928549 | 2825065 | 2955842201300277 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-03 | 2013-01-24 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | Yes | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42029602 | 2012-12-28 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the cable snapped on the large suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken near the proximal pulleys and proximal clevis cable opening. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment stick out at wrist. Additional observation not reported by site was proximal clevis wearing. The proximal clevis cable opening exhibited wear on one edge. Other cables at wrist were not damaged. Evidence not conclusive, but wear on hole could had contributed to cable breakage. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLEDRIVER | ENDOSCOPIC INSTRUMENT | 2953933 | 2928861 | 2825377 | 2955842201300278 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-03 | 2013-01-24 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | Yes | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-21 | Malfunction |  | It was reported prior to starting a da vinci si hysterectomy procedure that the pk dissecting forceps instrument cable was broken immediately upon placement in robot arm. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2953993 | 2928911 | 2825427 | 2955842201300279 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-28 | 2013-01-24 | No | Yes | Health Professional | 42022703 | No | No | 2012-12-28 | No | No | Unkown | 2013
42019408 | 2012-11-18 | Malfunction |  | Event Description: --/-- It was reported that during a davinci si hysterectomy procedure, the customer noted a broken cable on the mega needle driver instrument. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment stuck out at wrist. Other cables at wrist were not damaged. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. This report does not admit that the report or information submitted under this report constitutes an admission that the device, intuitive surgical or intuitive surgical employees, caused or contributed to the reportable event. | MEGA NEEDLE DRIVER | Reuse | 2954033 | 2928934 | 2825450 | 2955842201300280 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-02-06 | 2013-01-24 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2013-02-06 | Device Not Retu | No |  | 2013
42030901 | 2012-12-26 | Malfunction |  | It was reported during central processing the mega suturecut needle driver instrument was noted to have a broken wire. There were no missing or fallen pieces reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. Engineering also observed scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2954088 | 2929014 | 2825530 | 2955842201300282 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-24 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-26 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the mega suturecut needle driver instrument was noted to have wires coming out of the end of the instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the one grip close cable is frayed at the distal idler pulley. The frayed strands stick out at the wrist. The other cables at the wrist are not damaged. Engineering also observed a derailed cable. The same cable that is frayed is also derailed from the distal idler pulley. The one grip close cable is derailed at the distal idler pulley. The grips can still open and close, but movement may not be precise. Engineering concluded cable derailment is likely due to cable losing contact with pulley during wrist articulation. The cable derailment likely contributed to the cable fraying. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2954124 | 2929019 | 2822383 | 2955842201300283 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-24 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-26 | Malfunction |  | The fenestrated bipolar forceps instrument was returned without any allegations of malfunction. There were no missing or fallen pieces reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken cable. The pitch up cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2954119 | 2929035 | 2825551 | 2955842201300281 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-24 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42027802 | 2012-12-19 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the grasping retractor instrument was not recognized by the system. There were no missing or fallen pieces reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken pitch cable and scratches on the main tube. The pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in clevis. The clevis does not exhibit any damage or wear marks. The distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the analysis findings; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2954205 | 2929116 | 2825632 | 2955842201300284 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-24 | No | Yes | Health Professional | 42027802 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
420227 | 2013-01-14 | Malfunction |  | Pk dissector was inspected prior to use per policy. After using the robotic instrument for over an hour, the surgeon noticed that it had frayed. Surgeon tried to remove the instrument from the port but the instrument malfunctioned and could not be removed from the port. The port and instrument had to be removed from the pt (with instrument stuck in port and tip exposed). Once outside the pt's body, the instrument had to be forcefully removed from the port. Upon inspection, it was noted that one of the cables at the tip of the instrument was broken. The trocar was also removed from pt. The trocar was flushed and found to have a piece of cable inside the trocar. Per policy, an intra-op x-ray was taken and was negative. A new port and instrument were used thereafter. |  | DAVINCIINTUITIVE SURGICAL | PK DISSECTOR | 2964191 | 2938071 | 2834587 | 2938071 | 1 | NAY | User Facility |  | RISK MANAGER | Initial | 2013-01-17 | 2013-01-24 | No | Yes | Health Professional | 420227 | Yes | Yes |  | No Answer Provi |  | Invalid Data | 2013
42009308 | 2012-12-25 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the surgeon could not open/close the grips on the prograsp forceps instrument. The surgical staff removed the prograsp forceps instrument and inspection of the instrument found that it had a broken wire. The planned surgical procedure was completed with a replacement instrument. There were no missing or fallen pieces reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode, with the following clarification that the instrument was returned with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment is approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2951839 | 2926969 | 2823485 | 2955842201300266 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2012-12-27 | 2013-01-23 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2012-12-28 | Injury | Other | It was initially reported that during a da vinci si total benign hysterectomy procedure, the bipolar instrument fired unintentionally during a procedure according to the information provided to the clinical sales representative (csr), the customer alleged that while they were installing the bipolar instrument to the davinci robot system, as they connected the bipolar instrument cable to the valleylab generator, the electrosurgical unit (esu) made a firing tone and the bipolar instrument fired. The instrument tip did touch the patient's uterus (which was eventually removed). According to the information provided to the csr, the customer was able to successfully complete the case plus two additional cases using the same generator with no further incidents. | The clinical sales representative (csr) visited the site to verify that the customer was using the correct electrosurgical unit (esu), valley lab force, with the correct energy activation cable (blue valley lab cable). The csr confirmed that the equipment was compatible. Based on the additional information provided to the csr by the operating room nurse, no further issues had occurred after the cautery cable was replaced. The cautery cable used during the case was not returned for evaluation. As of (b)(6) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2952168 | 2927223 | 2823739 | 2955842201300267 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | Initial | 2012-12-28 | 2013-01-23 | Yes | Yes | Health Professional | IS3000 A60P7 | No | No | 2012-12-28 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-25 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the wire on the large needle driver instrument broke. The instrument was removed and the planned surgical procedure was completed with a replacement instrument. There were no missing or fallen pieces reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idler pulley. Idler pulley spins freely and exhibits damages on the edge and on the surface. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2952240 | 2927292 | 2823808 | 2955842201300268 | 1 | NAY | Manufacturer | ForeignCompany Representative | Other | Initial | 2012-12-27 | 2013-01-23 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-09-25 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the insulation at the tip of the maryland bipolar instrument was broken. The cables had to be cut in order to remove the instrument. A fragment fell into the patient, however, was removed during the surgical procedure. The intended surgery was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The instrument was found with proximal clevis completely dislodged from the main tube. The distal end of the main tube was broken and missing material approximately 0. 25 x 0. 30 in size. Evidence is not conclusive, but damage likely due to excess force applied normal to the clevis. The instrument has 7 uses remaining. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if additional information is received. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2952410 | 2927377 | 2823893 | 2955842201300269 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-11-20 | 2013-01-23 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-11-20 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-27 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the tips on the fenestrated bipolar forceps instrument were observed to be bent. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip is bent, causing side to side misalignment of grips. There is a. 066" offset at the tips, indicating overloading at the tip. The instrument passed electrical continuity testing. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering evaluation also found that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The surface of the distal pulley also exhibits some scratches. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage found to the instrument's pitch cable found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2952491 | 2927435 | 2823951 | 2955842201300270 | 1 | NAY | Manufacturer | Other | Other | Initial | 2013-01-08 | 2013-01-23 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42018907 | 2012-12-01 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a broken wire on the double fenestrated grasper instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | DOUBLE FENESTRATED GRASPER | ENDOSCOPIC INSTRUMENT | 2952549 | 2927462 | 2823978 | 2955842201300271 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2013-01-03 | 2013-01-23 | No | Yes | Health Professional | 42018907 | Device Returned To Manufacturer | Yes | 2013-01-03 | Device Not Retu | No | Reuse | 2013
42029602 | 2012-12-21 | Malfunction |  | It was reported that during set-up of a da vinci si surgical procedure, the blades on the suturecut needledriver instrument were dull. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode. Instrument will not cut suture in its current condition, as the blades do not open/close properly. There is a gap between the blades when grips are in the closed position. Engineering evaluation also found that one grip is bent backwards, the force applied is perpendicular to the blade cutting surface. The grips have a. 050 side to side offset at the tips as a result of bending. The visual inspection shows that the bent grip is cracked at the edge of the crimp pocket. The cross sectional area of the grip is at its smallest at the crack location. Engineering concluded that the cracking likely allowed the grip to bend backwards and that the damage is likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the damage found to the instrument's grip found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur. | LARGE SUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2952589 | 2927534 | 2824050 | 2955842201300272 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-28 | 2013-01-23 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2012-12-28 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-18 | Malfunction |  | It was reported that during reprocessing of the prograsp forceps instrument it was noted that the tips are not bent. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode with the following clarification that the instrument was returned with a frayed grip cable at the distal idler pulley. Engineering evaluation also found that the idler pulley on the clevis exhibits physical damage. Engineering concluded that the damage is likely due to excess force contact. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the damage found to the instrument's grip cable found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2952615 | 2927542 | 2824058 | 2955842201300273 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-27 | 2013-01-23 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
42034301 | 2012-12-04 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the thoracic grasper instrument was not grasping. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode, however, engineering evaluation found that the main tube insulation appears to have scratches and gouge marks. Engineering concluded that the damage is likely caused by external excess force contact on the main tube surface. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the damage to the instrument's main tube found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2952629 | 2927573 | 2824089 | 2955842201300274 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-27 | 2013-01-23 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Initial | 2013
07 | 2013-01-14 | Malfunction |  | During davinci hysterectomy near start of procedure maryland bipolar forcep in arm 2 was not moving correctly. Physician could not get it to move where he wanted. Assistant tried removing instrument and resetting the cassette. That did not work. Instrument removed and a small piece of wire near tip of instrument was noted to be exposed. Instrument had 3 lives left. What was the original intended procedure?hysterectomy. |  | DAVINCI ENDOWRIST MARYLAND BIPOLAR FORCEP | SYSTEMSURGICALCOMPUTER CONTROL | 2968264 | 2943840 | 2840356 | 2943840 | 1 | NAY | User Facility |  |  | Initial | 2013-01-23 | 2013-01-23 | No | Yes | Physician | 7 | Device Returned To Manufacturer | No Answer Provided |  |  |  |  | 2013
42009308 | 2012-09-04 | Malfunction |  | It was reported that during a da vinci si cystectomy procedure, it was observed that the steel wire on the prograsp forceps instrument was broken. There were no missing or fallen pieces reported. The instrument was removed from the patient and the planned surgical procedure was completed with an instrument of the same type. There was no patient harm, adverse outcome or injury reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2947732 | 2923334 | 2819850 | 2955842201300237 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-20 | 2013-01-22 | No | Yes | Health Professional | 42009308 | No | No | 2012-12-20 | No | No | Reuse | 2013
42017207 | 2012-10-09 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, it was observed that the steel wire on the maryland bipolar forceps instrument was broken. There were no missing or fallen pieces reported. The instrument was removed from the patient and the planned surgical procedure was completed with an instrument of the same type. There was no patient harm, adverse outcome or injury reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2947744 | 2923344 | 2819860 | 2955842201300238 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-20 | 2013-01-22 | No | Yes | Health Professional | 42017207 | No | No | 2012-12-20 | No | No | Reuse | 2013
42019408 | 2012-12-20 | Malfunction |  | It was reported that during central processing, a broken cable on the mega needle driver instrument was noted. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. Other cables at the wrist are not damaged. The instrument has 3 uses left. Additional findings revealed scratches and indentations on distal pulleys. There are scratches on the surfaces of both pulleys and indentation on the edges of both pulleys as well. On (b)(4) 2013 a follow up call was made with the customer regarding the failure analysis findings. It was stated that no fragments were observed falling into the patient. The reported instrument issue was found during cleaning and sterilization. No further information was available. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2947766 | 2923347 | 2819863 | 2955842201300239 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-21 | 2013-01-22 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-12-21 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2012-12-21 | Death | Death | It was reported that during a da vinci si hysterectomy procedure, an emergency conversion to open surgery due to lack of visualization had occurred. The patient expired during the open procedure. There were no reports of any system, instrument(s) and/or accessory(ies) malfunction during the course of the procedure. | During a follow up call on (b)(4) 2012 with the company representative who was present during this event, it was indicated that during this da vinci si hysterectomy procedure the surgeon inadvertently damaged a vessel while taking down adhesions, causing significant bleeding into the patient's abdomen. Reportedly, the surgeon was unable to control the patient's bleeding and the patient expired. It is unknown which exact vessel was damaged. Allegedly, the surgeon was aware that the patient had significant adhesions prior to starting the procedure and the surgeon was possibly going to complete the planned procedure using open surgical techniques. No malfunction of the da vinci surgical system, instruments and/or accessories occurred during the surgery. During a secondary follow-up on (b)(4) 2013 with the nurse, it was stated that the hospital's risk management department has advised not to provide any further information, as this event was not related to any isi device malfunction. No further information was obtained. Should additional information become available, a follow-up mdr will be submitted. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2947773 | 2923387 | 2819903 | 2955842201300240 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-12-21 | 2013-01-22 | Yes | No | Health Professional | IS3000 A60P7 | No | No | 2012-12-21 | No | No | Reuse | 2013
42020505 | 2012-12-27 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a frayed cables at the distal end of the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2947837 | 2923424 | 2819940 | 2955842201300242 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-27 | 2013-01-22 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-20 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, while using the fenestrated bipolar instrument, the customer noted that the cable snapped during the case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. Instrument passed electrical continuity test. The clevis did not exhibit any damage or wear marks. Additional finding not reported by the site was scratches on the surface of both pulleys. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2947854 | 2923427 | 2819943 | 2955842201300241 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-27 | 2013-01-22 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
40027202 | 2012-12-21 | Malfunction |  | It was reported that prior to starting a da vinci si hysterectomy procedure, during set up, the surgical staff reported seeing a broke tip on the harmonic ace curved shears insert accessory. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of broken tip cannot be confirmed. Insert does not have any broken pieces. Insert was returned installed on ace instrument under rma 1045754. Instrument was installed on in-house is3000 system and driven. Clamp arm opened/closed fine. Additional observation not reported by site is a functional test failure. Harmonic handpiece and generator were used to perform a functional energy test. Generator producted an error message during test of harmonic energy (both max and min footswitch pedals). Functional test failure suggests a defect in the insert, but visual inspection shows no obvious damage to curved blade. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | HARMONIC ACE CURVED SHEARS INSERT | ENDOSCOPIC ELECTROSURGICAL INS | 2947875 | 2923463 | 2819979 | 2955842201300248 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-21 | 2013-01-22 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2012-12-21 | Device Not Retu | No | Reuse | 2013
42000306 | 2012-12-21 | Malfunction |  | It was observed during central processing that the small clip applier instrument had a wire coming out at the tip of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint wires coming out of tip is confirmed. One grip close cable is broken at the distal idlers. Idler pulley spins freely and was not damaged. Cable segment sticks out at wrist. Other cables at wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2947860 | 2923465 | 2819981 | 2955842201300246 | 1 | NAY | Manufacturer |  | Other | Initial | 2012-12-21 | 2013-01-22 | No | Yes | Health Professional | 42000306 | Device Returned To Manufacturer | No | 2012-12-21 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-03-15 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff reported seeing a broken tip on the monopolar curved scissors instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2947876 | 2923466 | 2819982 | 2955842201300250 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-20 | 2013-01-22 | No | Yes | Health Professional | 42017910 | No | No | 2012-12-20 | No | No | Unkown | 2013
42017704 | 2012-12-20 | Malfunction |  | It was observed during central processing that the bowel grasper instrument had splintering at the shaft. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of instrument shaft is splintering was confirmed. Instrument main tube was found damaged at multiple location, with material removed. Evidence not conclusive, but damage is likely caused by excess force contact on the main tube surface. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | BOWEL GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2947862 | 2923474 | 2819990 | 2955842201300247 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-20 | 2013-01-22 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2012-12-20 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-18 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff reported seeing frayed wires on the fenestrated bipolar forceps. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2947864 | 2923477 | 2819993 | 2955842201300245 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-20 | 2013-01-22 | No | Yes | Health Professional | 42020505 | No | Yes | 2012-12-20 | No | No | Unkown | 2013
42029602 | 2012-12-21 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, during set up, the surgical staff reported seeing a broken cable on the large suturecut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of one of the cables broke is confirmed with the following clarification: cable is frayed. One grip close cable is frayed at the distal idler pulley. Frayed strands stick out at the wrist. Other cables at wrist are not damaged. Additional observation not reported by site is a derailed cable. Same cable that is frayed is also derailed from the distal idler pulley. Grips can still open and close, but movement may not be precise. Evidence not conclusive, but cable derailment is likely due to cable losing contact with pulley during wrist articulation. Derailment likely contributed to cable fraying. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2947896 | 2923482 | 2819998 | 2955842201300244 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-21 | 2013-01-22 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2012-12-21 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-20 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the surgical staff reported seeing a broken cable on the mega suturecut needle driver instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint broken cable is confirmed. The pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Additional finding: scratches on distal pulley. There are scratches on the surface and edges on the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2947900 | 2923483 | 2819999 | 2955842201300249 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-20 | 2013-01-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-20 | Device Not Retu | No | Reuse | 2013
42018406 | 2012-12-03 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, during set up, the surgical staff reported seeing broken housing on the permanent cautery spatula instrument. Nothing was reported having fallen into the patient. | The instrument was returned and evaluated. Complaint of broken housing cannot be confirmed. Housing snaps are intact and undamaged. Housing cannot be lifted up from chassis. No cracking damage observed on housing. Additional observations not reported by site are a broken distal clevis, missing conductor cap, and broken banana plug. Instrument was returned with one distal clevis ear broken off at its base. Clevis ear on side opposite conductor wire is broken off, resulting in dislodgment of yaw pulley. Broken piece was not returned, but is roughly. 285 x. 220 in size. Conductor cap likely detached as a result of clevis ear breakage. Yaw pulley boss feature on conductor wire side that mates with cap is broken off. Evidence not conclusive, but breakage is likely due to overloading the tip. Banana plug is cracked at the tip and the spring features are no longer formed correctly as a result. A monopolar cord cannot be installed on plug due to damage. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2947868 | 2923484 | 2820000 | 2955842201300251 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-20 | 2013-01-22 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-12-20 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, it was noted that the cable on the tenaculum forceps instrument was hanging out. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken near the proximal pulleys and proximal clevis cable hole. The idler pulley spins freely and does not exhibit damage. The cable segment sticks out at the wrist. No major wear was found on proximal clevis cable hole. No other damage was found. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2948498 | 2924066 | 2820582 | 2955842201300253 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-22 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-20 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the cable snapped on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be derailed at the idler pulley. The grips were still able to open and close with clicking snap sounds, but their movement could not be precise. The cable derailment was likely due to the cable losing contact with the pulley during wrist articulation. Fleet angle between proximal and distal pulleys could had contributed to the cable derailment. Additional observation not reported by the site was main tube insulation damage. The distal end of the main insulation tube exhibited various scratch marks with light material removal, suggesting that they may have been caused by instrument collisions or rough handling. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2948577 | 2924100 | 2820616 | 2955842201300254 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-27 | 2013-01-22 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-12-05 | Malfunction |  | It was reported that during a da vinci si surgical procedure the monopolar curved scissors instrument were noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. Additionally engineering also observed scratches on the maintube. The distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MONOPOLAR CURVED SCISSORSINSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2948814 | 2924306 | 2820822 | 2955842201300255 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2013-01-07 | 2013-01-22 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2013-01-07 | Device Not Retu | No | Reuse | 2013
42805206 | 2012-12-20 | Malfunction |  | It was reported that during a davinci si surgical procedure, the customer noted a broken tip on the monopolar cautery single site instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the distal end component was found to be bent. Visual inspection showed no broken components; however, the electrical contact located within the monopolar adapter was damaged. One leg of the contact was bent downwards. The bent leg created interference that prevented full installation of the disposable cautery hook. Instrument passed electrical continuity test. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MONOPOLAR CAUTERY SINGLESITE | ENDOSCOPIC ELECTROSURGICAL INS | 2948817 | 2924308 | 2820824 | 2955842201300259 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-27 | 2013-01-22 | No | Yes | Health Professional | 42805206 | Device Returned To Manufacturer | No | 2012-12-27 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-14 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure a spring came off of the megasuturecut needle driver instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken grip cable. The grip cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. Other cables are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2948795 | 2924320 | 2820836 | 2955842201300258 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-28 | 2013-01-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-28 | Device Not Retu | No | Reuse | 2013
42004906 | 2012-12-28 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure the cadiere forceps instrument was noted to have a broken cable. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment is approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | CADIERE FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2948833 | 2924343 | 2820859 | 2955842201300257 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-28 | 2013-01-22 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2012-12-28 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the pk dissecting forceps instrument had a loose wire looped at the top. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2949068 | 2924576 | 2821092 | 2955842201300256 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-04 | 2013-01-22 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-04 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-14 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted that the wheels on the back of the fenestrated bipolar forceps instrument were not spinning properly, identified during set up, not used in case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. Additional finding not reported by the site was a char mark on yaw pulley. The char mark on one side of the pulley melting part of the pulley. Engineering also found scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2949729 | 2925158 | 2821674 | 2955842201300260 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-28 | 2013-01-22 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-28 | Device Not Retu | No | Reuse | 2013
IS3000 | 2009-02-09 | Injury | Required Intervention | Per legal summons received on (b)(4) 2012, it was alleged that the patient had a da vinci s hysterectomy surgery on (b)(6) 2009 to remove ovarian fibroid tumors. Five weeks after surgery, the patient began feeling feverish and nauseous and had diarrhea. It was alleged the patient had not engaged in sexual intercourse since the surgery. At that time, it was alleged the patient engaged in sexual intercourse with her husband and experienced light bleeding. When the patient used the bathroom thereafter, the patient heard a pop and allegedly saw approximately an inch of intestines coming out of her vagina. The patient was admitted at that time to (b)(6). A different surgeon performed emergency open laparotomy surgery. While operating, a massive infection with cysts and adhesions encapsulating the fallopian tube and ovary was found and the right ovary had to be removed as a result. Additionally, the infection deteriorated the vaginal cuff tissue that had been sewn during the da vinci s hysterectomy causing the cuff to tear leading to a small bowel evisceration. The infected tissue was removed and vaginal cuff re-sewn. The patient stayed in the hospital five days. | On (b)(4) 2013, the surgeon was contacted to request additional information regarding the reported complaint. It was indicated the surgery occurred on (b)(6) 2009 at 11 am and no malfunctions occurred with the da vinci s system, instruments or accessories; no complications were noted during the surgery. No arcing occurred during the procedure, while using bipolar energy. The physician informed that a recommendation had been made to wait eight weeks before resuming sexual intercourse. The patient engaged in sexual intercourse prior to recommended eight weeks, the patient was admitted for evisceration of the bowel on (b)(6) 2009 at (b)(6) where care was provide by a different surgeon. A review of system logs for the system used in the procedure did not produce any data for the given system at the time period indicated. The surgery was not found in the system logs and could not be checked. | DA VINCI S | ENDOSCOPIC INSTRUMENT CONTROL | 2949800 | 2925169 | 2821685 | 2955842201300243 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-17 | 2013-01-22 | Yes | No | Health Professional | IS3000 | No | No |  | No | No | Reuse | 2013
40018012 | 2012-12-23 | Injury | Other | It was reported that during a da vinci si axillary lymph node dissection procedure, arcing from the mcs tip cover accessory installed on the monopolar curved scissors instrument occurred, causing damage to the patient's iliac vein. The affected vessel was repaired and the planned surgical procedure was completed. | The tip cover accessory and monopolar curved scissors instrument were not returned to intuitive surgical for evaluation; however, a video recording of the surgical procedure was received and reviewed by intuitive surgical. Review of the procedure video showed that the back side of the mcs tip cover accessory installed on the mcs instrument collided and pushed against the wrist area of the cadiere grasper instrument that was installed on the 3 patient side manipulator (psm) arm located on the patient side cart. Review of the video recording showed that there were several scrape marks on the back side of the tip cover, due to colliding and rubbing against the pulley/cables on the cadiere grasper instrument. The video recording showed that arcing from the tip cover accessory occured twice during the surgical procedure; however, the first instance of arcing was not observed by the surgeon. The second instance of arcing was observed and the site removed and replaced the tip cover accessory. Based on the video recording findings, intuitive surgical has determined that the arcing was the result of the tip cover making contact with the cadiere grasper instrument during the surgical procedure. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. On (b)(6) 2013, the isi clinical sales representative met with the surgeon who performed the surgical procedure. The surgeon indicated to the csr, that the patient recovered fine and that the surgeon indicated that the patient had not experienced any post surgical complications as a result of the reported event. | 8MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2950640 | 2925864 | 2822380 | 2955842201300261 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-12-23 | 2013-01-22 | Yes | Yes | Health Professional | 40018012 | No | No | 2013-12-23 | No | No | Unkown | 2013
42030901 | 2012-12-26 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wire on the megasuturecut needle driver instrument was observed to be coming out at the end of the instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and does not exhibit damage. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. Engineering evaluation also found that the idler pulley on which the broken cable was seated exhibits various light scratch marks on the outer face. Engineering concluded that the scratches may have been associated with the cable breakage. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLEDRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2950606 | 2925867 | 2822383 | 2955842201300262 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-22 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-28 | Malfunction |  | It was reported that during a da vinci s procedure the spring came off of the megasuturecut needle driver instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2950655 | 2925873 | 2822389 | 2955842201300263 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-28 | 2013-01-22 | No | Yes | Health Professional | 42030901 | No | No | 2012-12-28 | No | No | Unkown | 2013
42022703 | 2012-12-26 | Malfunction |  | It was reported that during reprocessing of the pk dissecting forceps instrument it was noted that the wire on the instrument was broken. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a frayed pitch cable at distal clevis hub. No excessive damage was found at the clevis. The frayed segment is approximately 0. 03 in length. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2950666 | 2925903 | 2822419 | 2955842201300264 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-22 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-20 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable on the large needle driver instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the drive cables are not broken or frayed at the wrist. The grips are found to open/close properly when the input discs are manually rotated. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches are. 080 -. 200 in length and not aligned with the tube axis. Engineering concluded that the damage is likely due to mishandling/misuse. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2950701 | 2925921 | 2822437 | 2955842201300265 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-22 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-05-08 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, it was observed that the steel wire on the large needle driver instrument was broken. There were no missing or fallen pieces reported. The instrument was removed from the patient and the planned surgical procedure was completed with an instrument of the same type. There was no patient harm, adverse outcome or injury reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2947362 | 2923121 | 2819637 | 2955842201300235 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-12-20 | 2013-01-18 | No | Yes | Health Professional | 42000606 | No | No | 2012-12-20 | No | No | Reuse | 2013
42017207 | 2012-12-03 | Malfunction |  | It was reported that during a davinci s prostatectomy procedure, the bipolar forceps instrument was not responding, the customer observed that the cable sticks at distal level. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub, and proximal clevis exhibited various scratch marks. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2947511 | 2923244 | 2819760 | 2955842201300236 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-12-26 | 2013-01-18 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the jaws on the fenestrated bipolar instrument would not open. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the grips were able to open and close fine; however, evaluation found that one grip is bent, causing side to side misalignment of grips. There is a. 085 offset at the tips, indicating overloading at the tip. The instrument passed electrical continuity testing. Engineering evaluation also found the distal end of the main tube has various scratch marks with light material removal, suggesting that the damage may have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The failure analysis finding does not constitute a mdr reportable event; however; the damage to the instrument's main tube, could likely cause or contribute to an adverse event if the malfunction were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2944297 | 2920076 | 2811079 | 2955842201300218 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-02 | 2013-01-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-02 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-10 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the pk dissecting forceps instrument was not rotating or moving correctly. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer reported complaint could not be confirmed. The instrument was placed on an in-house system and driven. Recognition and engagement passed. The instrument moved intuitively with full range of motion in all directions. The grips opened and closed properly. The instrument is fully functional. Electrical continuity passed. The instrument has 1 use remaining. However, additional findings revealed scratches on the main tube. The distal end of the main tube has various scratch marks with light material removal. Scratches are short in length and not aligned with the tube axis, suggesting it may have been caused by instrument collisions or rough handling. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. The instruments and accessories user manual specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2944396 | 2920087 | 2816603 | 2955842201300222 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-17 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-18 | Malfunction |  | The customer returned a maryland bipolar instrument with no failure or complaint statement. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. No reason for return was provided. Engineering observed tube damage at the distal end during visual inspection. The distal end of the main tube has various scratch marks with light material removal. Scratches are. 085 -. 140 in length and not aligned with the tube axis. Evidence was not conclusive, but damage may have been caused by mishandling/misuse. There was 1 use left. No other damage was found. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2944357 | 2920127 | 2816643 | 2955842201300219 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-18 | 2013-01-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-18 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-12-07 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the tips on the monopolar curved scissors instrument were loose. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer reported complaint of tips being loose cannot be confirmed. The scissors are not damaged and drive cables are not loose at the wrist upon visual inspection. Additional observations not reported by the site are a cracked tube extension and tube damage. The tube extension is cracked at the proximal clevis interface. The tube extension fractured next to one of the keys that mates with the prox clevis. The evidence was not conclusive, but the tube extension likely cracked due to excessive side loading or other mishandling/misuse. The distal end of the main tube has a 1 long section with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. There was 1 use left. No other damage was found. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. Based on this, no additional follow up is required. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HOT SHEARS TM MONOPOLARCURVED SCISSORS | ENDOSCOPIC ELECTROSURGICAL INS | 2944365 | 2920135 | 2816651 | 2955842201300220 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-18 | 2013-01-17 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2012-12-18 | Device Not Retu | No | Reuse | 2013
42027804 | 2012-12-12 | Malfunction |  | Event Description: --/-- On (b)(6) 2013 we received user facility med watch report number 1300060000-2012-8027 which stated: prior to a procedure a cable on the grasping retractor was found to be broken. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). --/-- Additional Event Description: --/-- It was reported that prior to starting a da vinci s surgical procedure, the orange sleeve at the distal end of the grasping retractor instrument was separating. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The customer reported complaint of the orange sleeve at the distal end separating cannot be confirmed, as this instrument does not have an orange sleeve (tube extension). There is nothing observed to be separating at the distal end. Additional observations not reported by the site are a derailed cable and tube damage. The pitch up cable is derailed from the pitch pulley located below the distal clevis hub. The cable is wedged between two pulleys. The cable tension is decreased and movement may not be as precise due to the derailment. The distal end of the main tube has a couple of 1 long sections with material removed on one side of the tube. The damaged area is parallel to tube axis and has a rough surface finish. There was 1 use left. No other damage was found. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. Based on this, no additional follow up is required. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. --/-- Additional Manufacturer Narrative: --/-- | GRASPING RETRACTOR | Reuse | 2944390 | 2920138 | 2805531 | 2955842201300221 | 1 | NAY | Manufacturer | OtherUser facility | Other | InitialFollowup | 2012-01-21 | 2013-01-17 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2012-01-21 | Device Not Retu | No |  | 2013
42030901 | 2012-12-26 | Malfunction |  | The mega suture cut needle driver instrument was returned without an initial complaint notification. As received, instrument shows a broken grip cable. | The instrument was returned and evaluated. No reason for return was provided. A complaint cannot be confirmed or denied. Engineering observed a broken cable. One grip close cable was found to be broken near the proximal pulleys and proximal clevis cable opening was found. The idler pulley spins freely and did not exhibit any damage. The cable segment sticks out at wrist. Proximal clevis cable opening exhibited wear on one edge. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2944756 | 2920490 | 2817006 | 2955842201300223 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-21 | Malfunction |  | It was reported that during a davinci s surgical procedure, the customer noted a broken cable on the fenestrated bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2944768 | 2920517 | 2817033 | 2955842201300224 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-26 | 2013-01-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42019004 | 2012-12-18 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the shaft of the cobra grasper instrument shredded while being removed from the patient, causing fragments from the cobra grasper instrument to fall into the patient's abdomen. The fragments were removed laparoscopically. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | On (b)(4) 2013, isi contacted the isi clinical sales representative (csr) that was present during the surgical procedure. The csr indicated that all of the fragments were removed from the patient and there were no intra-operative or post-surgical complications experienced by the patient. The csr indicated that on (b)(4) 2013, the surgeon reported that a post-surgical follow-up with the patient found that the patient is recovery very well and the patient has not reported experiencing any post-surgical complications as a result of the reported event. The csr indicated that a video recording of the surgical procedure is not available and that the cannula used during the surgical procedure was also inspected by site and the site found no damage to the cannula. The csr indicated that the surgeon used the cobra grasper instrument for grasping the patient's uterus. The csr indicated that the cobra grasper instrument is not being returned to isi for evaluation, as the site would like to retain the instrument for their records. Therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | COBRA GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2945394 | 2921098 | 2817614 | 2955842201300225 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-12-18 | 2013-01-17 | No | Yes | Health Professional | 42019004 | No | No | 2012-12-18 | No | No | Reuse | 2013
42020505 | 2012-12-18 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the fenestrated bipolar forceps instrument lost mobility due to a broken wire. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode of lost mobility however, evaluation found that the pitch cable at the distal idler is frayed. The instrument also passed electrical continuity testing. Engineering evaluation also found that the surface of the distal pulley exhibits scratches and that one grip is bent, causing side to side misalignment of the grips. There is a. 068 offset at the tips, indicating overloading at the tip. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2945449 | 2921158 | 2817674 | 2955842201300226 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-18 | 2013-01-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-18 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-06 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, the customer noted that the maryland bipolar forceps instrument's tips were misaligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent, causing side to side misalignment of grips. There was a 0. 127 offset at the tips, indicating overloading at the tip. Instrument passed electrical continuity test. Additional observation not reported by site was main insulation tube damage. The distal end of main tube had various scratch marks with light material removal. The scratches were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2945452 | 2921162 | 2817678 | 2955842201300227 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-26 | 2013-01-17 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-18 | Malfunction |  | It was reported that during reprocessing of the megasuturecut needle driver instrument the wires were exposed. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2945459 | 2921172 | 2817688 | 2955842201300228 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-18 | 2013-01-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-18 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-06 | Malfunction |  | It was reported that during a davinci si surgical procedure, the prograsp forceps instrument did not work anymore. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a frayed grip cable was found at distal idler pulley. The frayed grip cable also became derailed at the distal idler pulley. The idler pulley rim exhibited damage. Grips were still able to open and close, but their movement and functionality could not be precise. Additional observation not reported by site was main insulation tube damage. The distal end of main tube had various scratch marks with light material removal. Engineering concluded that damage was likely due to mishandling. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2945490 | 2921182 | 2817698 | 2955842201300229 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-12-18 | 2013-01-17 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-19 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, it was noted that the cables at the scissor end on the megasuturecut needle driver instrument were coming out. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. Additionally there is an indentation on the edge of the distal pulley that was observed though not reported by the site. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2945634 | 2921311 | 2817827 | 2955842201300230 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-11-28 | Malfunction |  | It was reported during a da vinci si hysterectomy procedure, one of the wires at the distal end of the fenestrated bipolar forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed using an instrument of the same type. No patient harm adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2945653 | 2921317 | 2779043 | 2955842201300232 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-18 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, while using the pk dissecting forceps instrument, the customer noted a broken cable at the tip. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2945648 | 2921344 | 2817860 | 2955842201300231 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-17 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-18 | Malfunction |  | Event Description: --/-- It was reported during a da vinci si hysterectomy procedure, it was noted that wires on the fenestrated bipolar forceps instrument were broken. There were no missing or fallen pieces reported. The planned surgical procedure was completed using an instrument of the same type. No patient harm adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity passed. Engineering evaluation also observed scratches on the surface of both distal pulleys. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity passed. Engineering evaluation also observed scratches on the surface of both distal pulleys. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | Reuse | 2945672 | 2921370 | 2817886 | 2955842201300233 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-17 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No |  | 2013
42030901 | 2012-11-12 | Malfunction |  | It was reported that after completion of a successful da vinci si hysterectomy procedure, it was noted that the wires at the distal tip of the megasuturecut needle driver instrument were loose. There were no missing or fallen pieces reported. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode, however, for clarification, the damaged instrument component was a broken grip cable. Based on this clarification, the customer reported complaint is confirmed. One grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. Engineering evaluation also found the distal pulley exhibit damage. One pulley has scratches on the surface and the other has an indentation at the edge. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2945694 | 2921391 | 2817907 | 2955842201300234 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-19 | 2013-01-17 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-14 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, a frayed cable was noted on the maryland bipolar forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the conductor wires are intact and undamaged, but the drive cable is frayed. One grip close cable is frayed at the distal idler pulley. Frayed strands stick out at the wrist. Other cables at the wrist are not damaged. There were 3 uses left. No other damage were found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2941808 | 2918000 | 2814516 | 2955842201300200 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-17 | 2013-01-16 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-17 | Device Not Retu | No | Reuse | 2013
40030901 | 2012-12-14 | Malfunction |  | It was reported that during a da vinci standard hysterectomy procedure frayed wires were noted on the mega suturecut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the damaged instrument component was a broken grip cable. One grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. Other cables at the wrist are not damaged. The instrument has 7 uses left. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2942319 | 2918380 | 2802255 | 2955842201300201 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-17 | 2013-01-16 | No | Yes | Health Professional | 40030901 | Device Returned To Manufacturer | No | 2012-12-17 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-19 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, there were wires noted coming out from the distal end of the tenaculum forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in clevis. The clevis does not exhibit any damage or wear marks. The instrument has 6 uses remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS | ENDOSCOPIC INSTRUMENT | 2942343 | 2918399 | 2814915 | 2955842201300203 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-16 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42018907 | 2012-12-18 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, the cable of the double fenestrated grasper instrument was noted to be frayed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is frayed near the distal idler pulley. The frayed strands stick out at the wrist. Other cables at the wrist are not damaged. There were 6 uses left. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | DOUBLE FENESTRATED GRASPER | ENDOSCOPIC INSTRUMENT | 2942266 | 2918424 | 2814940 | 2955842201300202 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-18 | 2013-01-16 | No | Yes | Health Professional | 42018907 | Device Returned To Manufacturer | Yes | 2012-12-18 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-13 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure a frayed cable was noted on the prograsp forceps instrument. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch up cable is frayed at the distal clevis hub. The frayed strands stick out at the wrist. Other cables at the wrist are not damaged. Additional observation not reported by the site is a derailed cable at the back end. The pitch cables are loose at the wrist. The housing was removed to find one pitch cable derailed at the back idler pulleys. The cable is wedged between two pulleys and has lost tension. The clamping pulley screws are still tight. There were 4 uses left. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2942407 | 2918494 | 2815010 | 2955842201300204 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-19 | 2013-01-16 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a frayed wire at the distal end of the fenestrated bipolar forceps instrument was noted. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken pitch cable. The pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity passed. The instrument has 5 uses remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2942470 | 2918512 | 2815028 | 2955842201300205 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-19 | 2013-01-16 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-05 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the maryland bipolar instrument was not moving properly. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that intuitive motion was not being experienced due to both pitch cables being broken at the distal clevis hub. Both cable crimps remained in the clevis. The instrument has 5 uses remaining. Additional observation not reported was insulation damage to the main tube. Material was missing, as the surface of main tube appeared to be scraped off. Evidence was not conclusive, but damage may have been a result of misuse. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. Based on this, no additional follow up is required. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2942467 | 2918534 | 2815050 | 2955842201300206 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-20 | 2013-01-16 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-20 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-18 | Malfunction |  | It was reported that during set-up of a da vinci si surgical procedure it was noted that the tip on the fenestrated bipolar forceps instrument was broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still remained in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2942734 | 2918756 | 2815272 | 2955842201300207 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-18 | 2013-01-16 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-18 | Device Not Retu | No | Initial | 2013
42030901 | 2012-12-21 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, while the surgeon was suturing the mega suturecut needle driver instrument broke exposing the wire and stopped working. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. Other cables at the wrist are not damaged. The instrument has 3 uses left. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2942756 | 2918797 | 2815313 | 2955842201300208 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-21 | 2013-01-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-21 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable on the fenestrated bipolar forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2942794 | 2918816 | 2815332 | 2955842201300209 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-16 | No | Yes | Health Professional | 42020505 | No | No | 2012-12-19 | No | No | Unkown | 2013
42019408 | 2012-12-19 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted that the wire snapped at wrist of mega needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley spins freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional finding not reported by the site was scratches on the surface of distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2942764 | 2918831 | 2815347 | 2955842201300210 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-26 | 2013-01-16 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-19 | Malfunction |  | It was reported that during reprocessing of the tenaculum forceps instrument, it was noted that the cable on the instrument was broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is missing. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2942855 | 2918848 | 2815364 | 2955842201300211 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-16 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-11-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wire on the large needle driver instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at wrist. The other cables at the wrist are not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2942901 | 2918899 | 2815415 | 2955842201300212 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-16 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Initial | 2013
42030901 | 2012-12-21 | Malfunction |  | It was reported that during a davinci s surgical procedure, the customer noted a broken wire on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken near the proximal pulleys. The idler pulley was able to spin freely and it did not exhibit any damage. The cable segment sticks out at the wrist. Additional observation not reported by site was pulley damage. The idler pulley where the broken cable was seated exhibited light scratch marks on the outer face. Evidence not conclusive, but scratches could be associated with cable breakage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2943021 | 2919008 | 2815524 | 2955842201300213 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-26 | 2013-01-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-26 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-19 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci s surgical procedure, the wire on the pk dissecting forceps instrument was observed to be sticking out. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as microscopic and visual inspection of the instrument found no wire or cable broken or damaged within the jaw assembly. Engineering evaluation also found that one grip was bent, causing side to side misalignment of grips. There was a 0. 02 offset at the tips. The grips did not meet together once the grips were closed. Engineering concluded that the damage was likely due to overloading at the tip excessive force applied normal to the grip. One cautery pin was also found bent at the pk riser and the cautery plug could not be inserted flush with the pin. | PK DISSECTING FORCEPS INSTRUMENT | Reuse | 2943272 | 2919232 | 2815748 | 2955842201300214 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2012-12-19 | 2013-01-16 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2013-01-17 | Yes | No |  | 2013
42030901 | 2012-12-18 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wire at the distal tip of the megasuturecut needle driver instrument broke off and fell into the patient. The broken fragment was retrieved and the planned surgical procedure was completed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was broken. Initially reported wire at distal tip broke off. The idler pulley spins freely and exhibit damage to the face and rim. The cable segment sticks out at the wrist. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2943280 | 2919237 | 2815753 | 2955842201300215 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-19 | 2013-01-16 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-11-19 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the wire on the fenestrated bipolar forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable is broken at the distal clevis hub. The broken strands stick out at the wrist. The cable segment that contains crimp is still installed in the clevis. The clevis does not exhibit damage. The pitch down cable is frayed at the distal clevis hub. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTURMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2943291 | 2919253 | 2815769 | 2955842201300216 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-16 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-17 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the tips on the maryland bipolar forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the tips are perfectly aligned and functional; however, engineering evaluation found that three of the four housing snap tabs and three adjacent housing pins are broken. The housing can be removed as a result of the damage. Engineering concluded that the damage is likely due to mishandling. The instrument passed electrical continuity testing. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and are not aligned with the tube axis, suggesting the damage was caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2943322 | 2919283 | 2815799 | 2955842201300217 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-16 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
IS2000 A51P8 | 2012-12-17 | Other | Other | It was reported that prior to starting a da vinci s prostatectomy procedure, during startup of the da vinci s system, system error codes 23013 and 23008 occurred. The surgeon had placed the ports in the patient, but did not start the surgery. According to the information provided by the customer, the surgical staff attempted to restart the system several times, but the error re-occurred. The patient was under anesthesia when the surgeon made the decision to convert to traditional open surgical techniques to complete the planned surgical procedure. No patient harm was reported. | The investigation conducted by the field service engineer concluded that the system error codes experienced by the site were associated with a patient side manipulator (psm) arm. The psm is an instrument arm located on the patient side cart and provides the sterile interface for the endowrist instruments. The system was repaired by replacing the affected psm. The system alarm (system generated fault code) functioned as designed and there was no injury to the patient. System error codes 23008 and 23013 appear when the da vinci s safety system determines that the angular position of one or more robotic joints on the specified manipulator, as measured by that joint's primary control sensor (encoder) and the secondary sensor (potentiometer), were out of specified tolerance for agreement. Upon determining this condition, the safety system puts da vinci in a recoverable safe state. The psm was returned to intuitive surgical for failure analysis investigation. Engineering was able to replicate the reported failure mode during testing and determined that an axis potentiometer and the upper pulley assemblies had malfunctioned, thus generating the system error codes experienced by the customer. As of (b)(6) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2940324 | 2916352 | 2812868 | 2955842201300184 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-12-17 | 2013-01-15 | Yes | Yes | Health Professional | IS2000 A51P8 | Device Returned To Manufacturer | Yes | 2012-12-17 | Device Not Retu | No | Reuse | 2013
42031802 | 2012-12-19 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a broken wire on the small grasping retractor instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with frayed pitch cable at distal idler. No damage found on pulley and clevis. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | SMALL GRASPING RETRACTOR | ENDOSCOPIC INSTRUMENT | 2940485 | 2916385 | 2812901 | 2955842201300185 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-19 | 2013-01-15 | No | Yes | Health Professional | 42031802 | Device Returned To Manufacturer | Yes | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42017704 | 2012-12-10 | Malfunction |  | It was reported that during a davinci si gastric bypass procedure, the customer was unable to open the bowel grasper instrument's tips. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode of instrument's tips could not be open. The grips were not stuck and when input disc was manually rotated, the instrument's tips were able to open/close. Additional observation not reported by site was a broken cable. One of the snake wrist cables was found to be broken at the distal most snake wrist disc. The snake wrist could not be properly straighten and the wrist motion was non-intuitive as a result of damage. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | BOWEL GRASPER 30MM JAW | ENDOSCOPIC INSTRUMENT | 2940587 | 2916464 | 2812980 | 2955842201300186 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-15 | No | Yes | Health Professional | 42017704 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-13 | Malfunction |  | It was reported that during a davinci s myomectomy procedure, the customer noted that the tips were not functioning nor moving in all directions on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2940998 | 2916919 | 2813435 | 2955842201300187 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-19 | 2013-01-15 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-19 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-11 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, while using the fenestrated bipolar instrument, it was noted that the jaws were misaligned, identified during set up, not used in case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode of misaligned jaws. The instrument's grips were not bent and the instrument's teeth mesh properly when grip were in the closed position. There was no offset located at the tips. Additional observations not reported by site were broken cable, missing crimp and damaged wire insulation. The pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was missing. The clevis did not exhibit any damage or wear marks. Other cables at wrist were not damaged. One conductor wire exhibited insulation damage near the proximal pulley. A section of bare wire was exposed. No char marks were observed on the distal end components. Instrument passed electrical continuity test. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2941033 | 2916940 | 2813456 | 2955842201300188 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-20 | 2013-01-15 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-20 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-06-19 | Malfunction |  | It was reported that during a davinci si prostatectomy procedure, the customer noted the prograsp forceps instrument was found faulty. The surgeon used a cobra grasper instrument to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument intuitive motion was not functioning due to loose pitch cable found at the distal clevis hub. The loose pitch cable was also found to be frayed. The clevis did not exhibit any damage or wear marks. Upon housing removal found a pitch cable loose at clamping pulley, but no loose clamping pulley screw was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2941045 | 2916993 | 2813509 | 2955842201300189 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-12-20 | 2013-01-15 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2012-12-20 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-18 | Malfunction |  | It was reported that during a davinci si surgical procedure, the customer noted that a wire popped out on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley spins freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional finding not reported by the site was scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2941067 | 2916998 | 2813514 | 2955842201300190 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-21 | 2013-01-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-21 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-17 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the cable on the maryland bipolar forceps instrument was observed to be frayed. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in clevis. The clevis does not exhibit any damage or wear marks. The instrument passed electrical continuity testing. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2941183 | 2917099 | 2813615 | 2955842201300191 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-18 | 2013-01-15 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-18 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-18 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the wire on the megasuturecut needle driver instrument was observed to be hanging from the end of the instrument. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2941209 | 2917144 | 2813660 | 2955842201300192 | 1 | NAY | Manufacturer | User facilityOther | Nurse | InitialFollowup | 2013-03-13 | 2013-01-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2013-03-13 | Yes | No | Reuse | 2013
42020706 | 2012-12-17 | Malfunction |  | On (b)(4) 2012 the tenaculum forceps instrument was returned without a receiving any prior complaint information. The customer was contacted to gain additional information, however no further information has been provided. | No reason for return was provided. A complaint cannot be confirmed or denied. Additional observation not reported by site is a broken cable. The pitch up cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Other cables at wrist are not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2941248 | 2917170 | 2813686 | 2955842201300199 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-17 | 2013-01-15 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-12-17 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-17 | Malfunction |  | On (b)(4) 2012 the mega suturecut needle driver instrument was returned without a receiving any prior complaint information. The customer was contacted to gain additional information, however no further information has been provided. | No reason for return was provided. A complaint cannot be confirmed or denied. Additional observations not reported by site are a broken cable and pulley damage. One grip close cable is broken at the distal idlers. Idler pulley spins freely. Cable segment sticks out at wrist. Other cables at wrist are not damaged. Distal idler pulley on same side as broken cable and top of idler pin exhibit various scratch marks on exposed outer surfaces. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2941262 | 2917171 | 2813687 | 2955842201300198 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-17 | 2013-01-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-17 | Device Not Retu | No | Reuse | 2013
42034301 | 2012-12-17 | Malfunction |  | It was observed during central processing that the 5mm thoracic grasper instrument had continued sloughing. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2941263 | 2917180 | 2811039 | 2955842201300196 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-17 | 2013-01-15 | No | Yes | Health Professional | 42034301 | No | Yes | 2012-12-17 | No | No | Unkown | 2013
42030901 | 2012-12-17 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff reported seeing a broken wire on the mega suturecut needle driver instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of wire broke is confirmed. One grip close cable is broken at the distal idlers. Idler pulley spins freely and does not exhibit damage. Cable segment sticks out at wrist. Other cables at wrist are not damaged. Additional observations not reported by site are a broken housing and corroded clamping pulleys. Housing has one rear snap tab and adjacent locating pin broken off. Rear corner of housing can be lifted up as a result of damage. Housing was removed to find clamping pulleys exhibiting light pitting corrosion. Clamping pulleys are not cracked and screws are still tight. Evidence not conclusive, but clamping pulley damage is likely due to improper cleaning. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. The cleaning and sterilization user manual specifically states: when scrubbing the tip of cautery instruments, take care not to damage the insulation. Do not expose instruments to hydrogen peroxide (h2o2), bleach, or alkaline-based cleaning agents, as this may result in instrument damage. O prolonged exposure to either ultrasonic cleaning or cleaning agents may result in instrument damage. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2941264 | 2917181 | 2813697 | 2955842201300197 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-17 | 2013-01-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-17 | Device Not Retu | No | Reuse | 2013
42029602 | 2012-11-01 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the surgical staff reported seeing a frayed cable at the tip of the large suturecut needle driver. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | ENDOSCOPIC INSTRUMENT | 2941265 | 2917182 | 2813698 | 2955842201300195 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-17 | 2013-01-15 | No | Yes | Health Professional | 42029602 | No | No | 2012-12-17 | No | No | Unkown | 2013
42034301 | 2012-12-17 | Malfunction |  | It was observed during central processing that the 5mm thoracic grasper instrument had continued sloughing. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2941270 | 2917183 | 2811039 | 2955842201300182 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-17 | 2013-01-15 | No | Yes | Health Professional | 42034301 | No | Yes | 2012-12-17 | No | No | Unkown | 2013
42009308 | 2012-08-10 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the surgical staff reported seeing a twisted cable at the tip of the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2941267 | 2917184 | 2813700 | 2955842201300194 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-17 | 2013-01-15 | No | Yes | Health Professional | 42009308 | No | No | 2012-12-17 | No | No | Unkown | 2013
42034301 | 2012-12-17 | Malfunction |  | It was observed during central processing that the 5mm thoracic grasper instrument had continued sloughing. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2941249 | 2917192 | 2811039 | 2955842201300183 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-17 | 2013-01-15 | No | Yes | Health Professional | 42034301 | No | Yes | 2012-12-17 | No | No | Unkown | 2013
42030901 | 2012-12-17 | Malfunction |  | It was observed during central processing that the mega suturecut needle driver instrument had a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint broken cable is confirmed. One grip close cable is broken at the distal idlers. Idler pulley spins freely and was not damaged. Cable segment sticks out at wrist. Other cables at wrist are not damaged. Additional finding: indentation with some scratches on distal pulley. There are some scratches on the surface and indentation on the edge of the distal pulley. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2941266 | 2917193 | 2813709 | 2955842201300193 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-17 | 2013-01-15 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-17 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-11-29 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the surgical staff reported seeing a broken wire on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint broken wire is confirmed. The pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Additional finding: scratches on maintube. Distal end of main tube has various scratch marks with light material removal. Scratches are short in length and not aligned with the tube axis, suggesting the may have been caused by instrument collisions or rough handling. More finding: indentation on the distal pulley. There are some indentations on the edge of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2937746 | 2913748 | 2780652 | 2955842201300158 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-13 | 2013-01-14 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-04 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the tips on the fenestrated bipolar forceps instrument did not meet. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. Late submission of this report is due to a reporting oversight by the responsible complaint handler. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip is bent, causing side to side misalignment of the grips. There is a. 092 offset at the tips, indicating overloading at the tip. The instrument passed electrical continuity testing. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering evaluation also found that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in clevis. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the damage found to the instrument's pitch cable found during failure analysis could likely cause or contribute to an adverse event if the malfunction were to recur. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2938155 | 2913990 | 2810506 | 2955842201300169 | 1 | NAY | Manufacturer | Other | Nurse | Initial | 2012-12-14 | 2013-01-14 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2012-12-14 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-04-25 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, the cable on the maryland bipolar forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed using an instrument of a different type and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one conductor wire is broken at the yaw pulley exit. The wire is detached from its connection at the grip. The instrument failed electrical continuity testing. The yaw exhibited no evidence of arcing. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and not aligned with the tube axis, suggesting they may have been caused by instrument collisions or rough handling. The bottom bipolar pin is also bent. Engineering concluded that the damage is likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2938277 | 2914072 | 2810588 | 2955842201300170 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-17 | 2013-01-14 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-17 | Device Not Retu | No | Initial | 2013
42000606 | 2012-08-10 | Malfunction |  | Event Description: --/-- It was reported that during cleaning and sterilization, the customer noted a frayed cable on the large needle driver instrument. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with a frayed pitch cable located at the distal idler. There was no damage found on the pulley and/or clevis. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | LARGE NEEDLE DRIVER | Reuse | 2938587 | 2914302 | 2810818 | 2955842201300171 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | InitialFollowup | 2013-01-15 | 2013-01-14 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No |  | 2013
42018406 | 2012-12-17 | Malfunction |  | It was reported that after successful completion of a da vinci si salpingectomy procedure the wire at the tip of the permanent cautery hook instrument was sticking out. There were no missing or fallen pieces reported. There was no patient harm, adverse outcome or injury. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still remained in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2938709 | 2914403 | 2810919 | 2955842201300172 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-17 | 2013-01-14 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | Yes | 2012-12-17 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-12 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a broken cable on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was found to be frayed at the distal idler pulley. The frayed strands stick out at the wrist. Other cables at wrist were not damaged. Engineering observed that the same grip cable was also found to be derailed at the distal idler pulley. The grips were able to open and close, but their movement could not be precise. The cable derailment was likely due to cable losing contact with pulley during wrist articulation. Fleet angle between proximal and distal pulleys may have contributed to the cable derailment. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2938808 | 2914478 | 2810994 | 2955842201300173 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-17 | 2013-01-14 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-17 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-11-27 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, customer noted that fenestrated bipolar instrument's tips were not aligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip was found to be bent, causing side to side misalignment of grips. There was a. 107 offset at the tips, indicating overloading at the tip. Instrument passed electrical continuity test. Engineering also observed a frayed pitch cable. Instrument was found with frayed pitch cable at distal idler. No damage found on the pulley or the clevis. Additional observation not reported by site was main insulation tube damage. The distal end of main tube had various scratch marks with light material removal. The scratches were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2938899 | 2914520 | 2811036 | 2955842201300175 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-17 | 2013-01-14 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-17 | Device Not Retu | No | Reuse | 2013
42034301 | 2012-12-17 | Malfunction |  | It was reported that during a da vinci si lobectomy procedure, it was noted that the coating on main tube of the thoracic grasper was peeling. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2012, isi contacted the initial reporter of this report. The initial reporter indicated that there was no report by the surgical staff that any pieces from the instrument fell into the patient. The initial reporter indicated that to the best of her knowledge it is unknown if the patient has returned to the hospital due to experiencing any post surgical complications related to retaining any foreign body. The initial reporter indicated that it is unknown if the cannula was inspected prior to use. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2938913 | 2914523 | 2811039 | 2955842201300176 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-17 | 2013-01-14 | No | Yes | Health Professional | 42034301 | No | No | 2012-12-17 | No | No | Reuse | 2013
42009308 | 2012-12-18 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was noted that one of the wires at end was broken on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2938917 | 2914557 | 2811073 | 2955842201300177 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-18 | 2013-01-14 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-18 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-18 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, it was noted a frayed cable on the fenestrated bipolar instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2938935 | 2914563 | 2811079 | 2955842201300178 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-18 | 2013-01-14 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-18 | Device Not Retu | No | Reuse | 2013
42018907 | 2012-12-18 | Malfunction |  | It was reported that prior to starting a davinci si surgical procedure, it was noted that the wires were sticking out at the distal end of the double fenestrated grasper instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | DOUBLE FENESTRATED GRASPER | ENDOSCOPIC INSTRUMENT | 2938995 | 2914632 | 2811148 | 2955842201300179 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-18 | 2013-01-14 | No | Yes | Health Professional | 42018907 | Device Returned To Manufacturer | No | 2012-12-18 | Device Not Retu | No | Reuse | 2013
42029602 | 2012-12-01 | Malfunction |  | Event Description: --/-- It was observed during central processing that the large suturecut needle driver instrument had a broken cable at the instrument tip. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was broken at the distal idlers. The idler pulley spun freely and was not damaged. The cable segment was sticking out at the wrist. The other cables at the wrist were not damaged. | LARGE SUTURECUT NEEDLE DRIVERINSTRUMENT | Reuse | 2939085 | 2914720 | 2811236 | 2955842201300180 | 1 | NAY | Manufacturer | User facilityOther | Other | FollowupInitial | 2013-01-15 | 2013-01-14 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2013-01-15 | Device Not Retu | No |  | 2013
42030901 | 2012-12-12 | Malfunction |  | It was reported that after a da vinci s hysterectomy procedure, the surgical staff reported seeing broken and loose wires at the distal tip of the mega suturecut needle driver. No patient harm, adverse outcome or injury was reported. | Complaint broken and loose wire is confirmed. One grip close cable is broken at the distal idlers. Idler pulley spins freely and was not damaged. Cable segment sticks out at wrist. Other cables at wrist are not damaged. Additional finding: indentation and scratch on distal pulley. There is an indentation on the edge and couple of scratches on the surface of one pulley. Other distal pulley is fine. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2939086 | 2914721 | 2811237 | 2955842201300181 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-17 | 2013-01-14 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-17 | Device Not Retu | No | Unkown | 2013
42034301 | 2012-12-17 | Malfunction |  | It was reported that during a da vinci si lobectomy procedure, it was noted that the coating on main tube of the thoracic grasper was peeling. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2012, isi contacted the initial reporter of this report. The initial reporter indicated that there was no report by the surgical staff that any pieces from the instrument fell into the patient. The initial reporter indicated that to the best of her knowledge it is unknown if the patient has returned to the hospital due to experiencing any post surgical complications related to retaining any foreign body. The initial reporter indicated that it is unknown if the cannula was inspected prior to use. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | THORACIC GRASPER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2939087 | 2914722 | 2811039 | 2955842201300174 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-17 | 2013-01-14 | No | Yes | Health Professional | 42034301 | No | No | 2012-12-17 | No | No | Reuse | 2013
42009308 | 2012-12-14 | Malfunction |  | It was reported that prior to starting a da vinci si hysterectomy procedure, during set up, the surgical staff reported seeing frayed cables at the distal end of the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint frayed cable is confirmed. Instrument was found with frayed pitch cable at distal idler. No damage found on pulley and clevis. Additional finding: scratches on maintube. Distal end of main tube has various scratch marks with light material removal. Scratches are short in length and not aligned with the tube axis, suggesting the may have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2936771 | 2912572 | 2809088 | 2955842201300168 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-13 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff reported seeing a bent tip on the pk dissecting forceps instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated Complaint tips are bent is confirmed One grip is bent causing side to side misalignment of grips There is a 082 offset at the tips indicating overloading at the tip Electrical continuity passed 48 additional finding is scratches on maintube Distal end of main tube has various scratch marks with light material removal Suggesting the may have been caused by instrument collisions or rough handling The instruments and accessories user manual specifically states general precautions and warnings handle instruments with care Avoid mechanical shock or stress that can cause damage to the instruments Do not use an instrument to clean debris from another instrument intraoperatively This may result in damage to the instruments or other unintended consequences such as disconnection of the instrument tip | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2936776 | 2912574 | 2809090 | 2955842201300167 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-13 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
42014304 | 2012-11-22 | Malfunction |  | It was reported that after a davinci si surgery procedure, when the surgical staff tried to take the maryland dissector instrument out of the cannula, they noticed shavings from shaft. Customer indicated that nothing fell inside the patient and there was no injury to the patient. | The instrument was returned and evaluated. Engineering was able to confirm the reported failure mode. The main insulation tube was found to be damaged at the distal end. The main insulation tube showed gouges on one side and exhibited a 0. 195 x 0. 065 piece missing roughly 6 above the snake wrist. There were a couple of areas below this section with scratches on the main insulation tube. The cannula accessory used with this instrument was returned and investigated under mdr 2955842-2013-00156. During the failure analysis, the accessory cannula was found with dents on the distal tip of the cannula. A gage pin could not be passed through the cannula tip. Engineering was able to confirm that the main insulation damage found on the instrument was caused by the cannula accessory. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MARYLAND DISSECTOR | ENDOSCOPIC INSTRUMENT | 2936574 | 2912378 | 2808894 | 2955842201300166 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-13 | 2013-01-12 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-14 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci s procedure the wire came out from the tip on the large needle driver instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering observed the one grip close cable is loose at the distal idlers. The idler pulley spins freely and was not damaged. The broken wire is inside housing by clamping pulley. The clamping pulley screws were found to be secured. The idler block shows no sign of damaged. The instrument was returned and evaluated. Engineering observed scratches on the maintube. The distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | Reuse | 2934705 | 2910487 | 2807002 | 2955842201300144 | 1 | NAY | Manufacturer | OtherUser facility | Other | FollowupInitial | 2013-01-14 | 2013-01-11 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-14 | Device Not Retu | No |  | 2013
A60P7 | 2012-01-29 | Injury | Other | It was reported that approximately two weeks after completion of a da vinci si hysterectomy procedure, it was discovered during a post-surgical check-up that the patient developed erythema at two of the port site incisions. The affected port sites were located at the upper left and upper right quadrants. Reportedly, a pk dissector instrument was installed in the upper left quadrant port site and a monopolar curved shears (mcs) instrument was installed in the upper right quadrant during the surgical procedure. | Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. Additionally, the sterile reprocessing point-of-contact has been contacted for further information. It was stated that the hospital follows isi's cleaning and sterilization protocol step by step, and that additional questions would be answered as soon as approval is obtained from the risk management office. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2935912 | 2911706 | 2808221 | 2955842201300145 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-12-12 | 2013-01-11 | Yes | No | Health Professional | A60P7 | No | No | 2012-12-12 | No | No | Reuse | 2013
A60P7 | 2012-11-27 | Injury | Other | It was reported that approximately two weeks after completion of a da vinci si hysterectomy procedure, it was discovered during a post-surgical check-up that the patient developed erythema at two of the port site incisions. The affected port sites were located at the upper left and upper right quadrants. Reportedly, a pk dissector instrument was installed in the upper left quadrant port site and a monopolar curved shears (mcs) instrument was installed in the upper right quadrant during the surgical procedure. The intended procedure was completed successfully. | On (b)(4) 2012, isi contacted the isi clinical sales representative (csr) who reported the event. The csr indicated that he was present during the surgical procedure and that he did not observe arcing from any of the cautery instruments installed during the surgical procedure. The csr also indicated that no malfunction of the site's da vinci si system, instruments or accessories occurred during the surgical procedure. The csr indicated that the surgical procedure took 45 minutes to complete and the patient was positioned in the lithotomy position and that the system was grounded properly. The csr indicated that it is unclear as to when the port site injury occurred and that the erythema experienced by the patient was mild and caused the patient some discomfort. The csr indicated that no intervention was required to treat the patient. The csr indicated that he has no additional information concerning the reported event. Additionally, the sterile reprocessing point-of-contact has been contacted to with further information. It was stated that the hospital follows isi's cleaning and sterilization protocol step by step, and that additional questions would be answered as soon as approval is obtained from the risk management office. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2935934 | 2911723 | 2808239 | 2955842201300146 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-12-12 | 2013-01-11 | Yes | No | Health Professional | A60P7 | No | No | 2012-12-12 | No | No | Reuse | 2013
A60P7 |  | Injury | Other | It was reported that approximately two weeks after completion of a da vinci si hysterectomy procedure, it was discovered during a post-surgical check-up that the patient developed erythema at two of the port site incisions. The affected port sites were located at the upper left and upper right quadrants. Reportedly, a pk dissector instrument was installed in the upper left quadrant port site and a monopolar curved shears (mcs) instrument was installed in the upper right quadrant during the surgical procedure. | On (b)(4) 2012, isi contacted the isi clinical sales representative (csr) who reported the event. The csr indicated that he was present during the surgical procedure and that he did not observe arcing from any of the cautery instruments installed during the surgical procedure. The csr also indicated that no malfunction of the site's da vinci si system, instruments or accessories occurred during the surgical procedure. The csr indicated that the surgical procedure took 60 minutes to complete and the patient was positioned in the lithotomy position and that the system was grounded properly. The csr indicated that it is unclear as to when the port site injury occurred and that the erythema experienced by the patient caused some irritation around the port site. The csr indicated that no intervention was required to treat the patient. The csr indicated that he has no additional information concerning the reported event. Additionally, the sterile reprocessing point-of-contact has been contacted for further information. It was stated that the hospital follows isi's cleaning and sterilization protocol step by step, and that additional questions would be answered as soon as approval is obtained from the risk management office. A follow-up mdr will be submitted if additional information is received. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2935950 | 2911743 | 2808258 | 2955842201300147 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-12-12 | 2013-01-11 | Yes | No | Health Professional | A60P7 | No | No | 2012-12-12 | No | No | Reuse | 2013
42806503 | 2012-12-13 | Malfunction | Other | Event Description: --/-- It was initially reported that as the gallbladder was being removed during a da vinci si cholecystectomy procedure, the surgeon pushed the pouch through the single site port, instead of the 10 mm trocar first. According to the reporter of the complaint, this was the surgeon's first single site procedure. When the pouch was pushed through, it allegedly went through the side of the port, tearing the patient's skin a little. Furthermore, the pouch also caused a plastic piece to break off the port, which according to the reporter the operating staff was unable to find. | Manufacturer Narrative: --/-- Isi contacted the hospital's robotics coordinator for further clarification regarding the reported incident. It was indicated that no patient injury had occurred. The incident happened while the surgeon attempted to introduce the 10 mm catch bag to the single site lumen port, which then tore apart. A piece of the port broke off and fell inside the patient; contrary to what has been initially reported, according to the robotics coordinator the fallen piece was successfully retrieved. The surgery was completed using the single-site port methodology with the da vinci surgical system. The patient was released from the hospital the day after the surgery. The single-site port used during the case was not returned for evaluation. According to the robotics coordinator, as of the date of the reported incident, the customer switched from using a 5 mm port back to a 10 mm port to insert the catch bag. There have been no recurrence if the same event. The single site instruments and accessories instructions for use (ifu) specifically state: warning: be sure to read and understand all information, particularly caution and warning information, found in the applicable user manuals before using these products. Failure to properly follow all instructions, including those in the da vinci si surgical system user manual, and instructions supplied with accessory devices like generators, may lead to injury and result in improper functioning of the device. --/-- Additional Manufacturer Narrative: --/-- It was indicated that no patient injury had occurred. | SINGLE SITE PORT 85MM | Unkown | 2936231 | 2911990 | 2808506 | 2955842201300148 | 1 | NAY | Manufacturer | Company RepresentativeUser facility | Other | FollowupInitial | 2012-12-14 | 2013-01-11 | Yes | Yes | Health Professional | 42806503 | No | No | 2012-12-14 | Device Not Retu | Yes |  | 2013
42034301 | 2012-12-03 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tips of the thoracic grasper instrument were not working properly. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering could not confirm the customer reported complaint of tips not working properly, however, there were findings of deep scratches on the main tube. Black tube insulation is damaged at the distal end. Insulation is gouged in various places and have pieces of insulation lifted up above the snake wrist and some missing measured roughly about 0. 167 x 0. 108. Metal shaft is exposed as a result. Evidence not conclusive, but damage is likely due to mishandling. The instrument has 17 uses left. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | THORACIC GRASPER | ENDOSCOPIC INSTRUMENT | 2936253 | 2912028 | 2808544 | 2955842201300151 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-11 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
IS2000 A51P7 | 2012-12-13 | Injury |  | It was reported that after completion of general surgical procedure using traditional laparoscopic techniques and the da vinci s surgical system, the surgeon was concerned that he may have hit the patient's bowel. | (b)(4) 2013, isi contacted the isi clinical sales representative who was present during the planned surgical procedure. The csr indicated that the surgeon utilized traditional laparoscopic instrumentation to perform dissection during the initial portion of the surgical procedure and that the da vinci s surgical procedure was utilized to complete the surgical procedure. The csr indicated that she observed that after completion of the surgical procedure, the system was undocked and moved away from the patient and she observed that the surgeon was performing an inspecting the surgical area. The csr discussed the issue with the proctor, and the proctor indicated to her that the surgeon may have hit the patient's bowel and that if a bowel injury had occurred; it was likely introduced prior to use of the da vinci surgical system. The csr indicated that she was present during the entire surgical procedure and she did not observe any malfunction of the site's da vinci s surgical system, instruments or accessories. Based on the limited information provided, it is indeterminable as to how damage was introduced to the patient's bowel. Isi has contacted the site to obtained additional information concerning the reported event; however, no additional information has been received. A follow-up medwatch report will be submitted if additional information is received. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2936257 | 2912030 | 2808546 | 2955842201300150 | 1 | NAY | Manufacturer | Company Representative | Other | Initial | 2012-12-13 | 2013-01-11 | Yes | Yes | Health Professional | IS2000 A51P7 | No | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, a broken wire was noted on the maryland bipolar forceps instrument. There were no instrument collisions. The planned procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported complaint of a broken wire with the following clarification: the drive cable is broken. The conductor wires are intact, but one pitch down cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in clevis. The clevis does not exhibit any damage or wear marks. Other cables at the wrist are not damaged. Additional observation not reported by the site is a bent grip. One grip is bent, causing side to side misalignment of the grips. There is a. 035 offset at the tips. Bent grip has a slight separation of the yaw pulley and pulley cover at the glue joint, indicating overloading at the tip. Electrical continuity passed. There were 7 uses left. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2936249 | 2912041 | 2808556 | 2955842201300149 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-11 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-11-29 | Malfunction |  | It was reported that the pk dissecting forceps instrument's tips were not closing properly during central processing. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering could not confirm the customer reported complaint of tips not closing properly; however, findings revealed a frayed pitch cable. The instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. Electrical continuity passed. The instrument has 1 use remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS | ENDOSCOPIC INSTRUMENT | 2936258 | 2912045 | 2794906 | 2955842201300152 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-11 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2012-12-12 | Malfunction |  | Event Description: --/-- This mdr was created to address the one of the two separate cases reported by the surgeon on mdr 2955842-2013-00120, 'based on additional information provided to the field service engineer, the surgeon claimed that the bipolar instrument had fired on its own in two separate cases. The surgeon did not report these incidents to intuitive surgical at the time of the occurrences. '. | Manufacturer Narrative: --/-- On (b)(4) 2013, isi received a user facility medwatch form from the fda (b)(4). According to the information reported to the fda by the physician, 'the surgeon was using the davinci system for a hysterectomy. The precise bipolar instrument was inserted and when she grasped tissue the current activated without the foot pedal being pressed. The audible power sound was heard and the grasped tissue which fortunately was a fallopian tube was cauterized. The entire crew in the operating room witnessed this event. She then grasped the uterus in several locations and the instrument again transmitted energy with the grasping of tissue without use of the foot pedal. The device was removed and the pk device inserted, tested and the case proceeded without incident. Exam of the bowel and surrounding tissue disclosed no injury and the patient was discharged the next morning. This only occurred 2 days ago however so potential for delayed injury still exists. A similar event had occurred 2 months ago using a fenestrated bipolar device on the same system and generator where upon insertion of the instrument energy was immediately delivered prior to foot pedal activation. The device was disconnected, reinserted and the problem did not recur on the second attempt. ' as stated on the initial mdr, isi contacted initial reporter biomed, at the memorial health system hospital. He indicated that he was not aware of any patient injuries related to the incidents reported. A review of the system logs showed no system malfunction occurring during the procedure on (b)(4) 2012. Several attempts have been made to gather additional information regarding this event without success. As of the date of this report, no additional information has been provided from the site. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The field service engineer (fse) visited the site to verify that the customer was using the correct electrosurgical unit (esu), valley lab force fx, with the correct energy activation cable (blue valley lab cable). The fse confirmed that the equipment was compatible. Based on the information provided to the fse by the operating room surgical staff, this surgeon is the only surgeon who had experienced this issue. Furthermore, it was stated that the surgeon did not realize she was pressing the foot pedal, activating energy to the bipolar instrument. The fse explained to the customer the importance of calling isi's technical support in such instances. The fse also performed a complete system verification and electrical safety test during the preventive maintenance visit. Together with the hospital's biomed personnel, the engineer verified both bipolar and mono-polar functionality with the customer's instruments. The fse was unable to duplicate the customer reported failure. A review of the system logs showed no system malfunction occurred during the procedure on (b)(6) 2012. The fse tested the system setup and found it to be working to manufacturer's specifications with no defect or malfunction. On (b)(4) 2013, isi contacted initial reporter biomed, at (b)(6). He indicated that he was not aware of any patient injuries related to these, incidents reported. | DA VINCI S SURGICAL SYSTEM | Reuse | 2936337 | 2912118 | 2805302 | 2955842201300153 | 1 | NAY | Manufacturer | OtherHealth Professional | Physician | FollowupInitial | 2013-01-15 | 2013-01-11 | No | Yes | Health Professional | IS3000 A60P7 | No | Yes | 2013-01-15 | Device Not Retu | No |  | 2013
IS3000 A60P7 |  | Malfunction |  | Event Description: --/-- This mdr was created to address the one of the two separate cases reported by the surgeon on mdr 2955842-2013-00120, 'based on additional information provided to the field service engineer, the surgeon claimed that the bipolar instrument had fired on its own in two separate cases. The surgeon did not report these incidents to intuitive surgical at the time of the occurrences. '. | Manufacturer Narrative: --/-- On (b)(4) 2013, isi received a user facility medwatch form from the fda (b)(4). According to the information reported to the fda by the physician, 'the surgeon was using the davinci system for a hysterectomy. The precise bipolar instrument was inserted and when she grasped tissue the current activated without the foot pedal being pressed. The audible power sound was heard and the grasped tissue which fortunately was a fallopian tube was cauterized. The entire crew in the operating room witnessed this event. She then grasped the uterus in several locations and the instrument again transmitted energy with the grasping of tissue without use of the foot pedal. The device was removed and the pk device inserted, tested and the case proceeded without incident. Exam of the bowel and surrounding tissue disclosed no injury and the patient was discharged the next morning. This only occurred 2 days ago however, so potential for delayed injury still exists. A similar event had occurred 2 months ago using a fenestrated bipolar device on the same system and generator where upon insertion of the instrument energy was immediately delivered prior to foot pedal activation. The device was disconnected, reinserted and the problem did not recur on the second attempt. ' as stated on the initial mdr, isi contacted initial reporter biomed, at the memorial health system hospital. He indicated that he was not aware of any patient injuries related to the incidents reported. Several attempts have been made to gather additional information regarding this event without success. As of the date of this report, no additional information has been provided from the site. A follow-up medwatch report will be submitted if the instrument is returned or if additional information is received. --/-- Additional Manufacturer Narrative: --/-- The field service engineer (fse) visited the site to verify that the customer was using the correct electrosurgical unit (esu), valley lab force fx, with the correct energy activation cable (blue valley lab cable). The fse confirmed that the equipment was compatible. Based on the information provided to the fse by the operating room surgical staff, this surgeon is the only surgeon who had experienced this issue. Furthermore, it was stated that the surgeon did not realize she was pressing the foot pedal, activating energy to the bipolar instrument. The fse explained to the customer the importance of calling isi's technical support in such instances. The fse also performed a complete system verification and electrical safety test during the preventive maintenance visit. Together with the hospital's biomed personnel, the engineer verified both bipolar and mono-polar functionality with the customer's instruments. The fse was unable to duplicate the customer reported failure. The fse tested the system setup and found it to be working to manufacturer's specifications with no defect or malfunction. On (b)(4) 2013, isi contacted initial reporter biomed, at (b)(6). He indicated that he was not aware of any patient injuries related to these, incidents reported. | DA VINCI S SURGICAL SYSTEM | Reuse | 2936326 | 2912129 | 2808645 | 2955842201300154 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | FollowupInitial | 2013-01-15 | 2013-01-11 | No | Yes | Health Professional | IS3000 A60P7 | No | Yes | 2013-01-15 | Device Not Retu | No |  | 2013
IS3000 A60P6 | 2012-11-22 | Malfunction |  | It was reported that after a da vinci si surgical procedure, the surgical staff reported seeing shavings from the shaft of a endowrist instrument, when it was being removed from the cannula. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of cannula likely causing shaft damage on the instrument was confirmed. Cannula was found with dents on the distal tip of the cannula. Gage pin cannot pass through the cannula tip. No other physical damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2936556 | 2912359 | 2808875 | 2955842201300156 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-13 | 2013-01-11 | No | Yes | Health Professional | IS3000 A60P6 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-13 | Malfunction |  | It was observed during central processing that the pk dissecting forceps instrument had a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of instrument cable broken was confirmed. The pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still remained in the clevis. Clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2936557 | 2912360 | 2808876 | 2955842201300160 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-13 | 2013-01-11 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-10 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the surgical staff reported seeing frayed wires on the pk dissecting forceps instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer initially reported frayed wire, however, for clarification the damaged instrument component was a broken pitch cable. Based on this clarification, the customer reported complaint is confirmed. The pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Electrical continuity passed. Additional finding: scratches on maintube. Distal end of main tube has various scratch marks with light material removal. Scratches are short in length and not aligned with the tube axis, suggesting the may have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2936560 | 2912362 | 2805533 | 2955842201300161 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-13 | 2013-01-11 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-13 | Malfunction |  | It was observed during central processing that the large needle driver instrument had a wire frayed at the middle of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The customer initially reported frayed wire, however, for clarification the damaged instrument component was a broken grip cable. Based on this clarification, the customer reported complaint is confirmed. One grip close cable is broken at the distal idlers. Idler pulley spins freely and was not damaged. Cable segment sticks out at wrist. Other cables at wrist are not damaged. Additional finding: indentation and scratches on pulley. There are scratches and indentation on both distal pulleys. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2936559 | 2912363 | 2808879 | 2955842201300162 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-13 | 2013-01-11 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
IS3000 A60P6 | 2012-12-13 | Other | Other | It was reported that during a da vinci si hysterectomy procedure, the surgical staff reported the arm of the da vinci system hit the patient in the face. It was unknown if the patient was injured during this time. The surgeon made the decision to convert the procedure to traditional open surgical techniques to complete the procedure. | Several attempts have been made to gather additional information regarding this event from this site without success. The investigation is ongoing, a follow-up mdr will be generated if additional details are received. A review of system logs for the date of the procedure showed that system did not have any errors during the procedure. On (b)(6) 2013, the initial reporter indicated that there were no injuries to the patient as a result of the collision. The open surgery was completed successfully without complication. | DA VINCI SI SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2936561 | 2912364 | 2808880 | 2955842201300155 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-13 | 2013-01-11 | Yes | Yes | Health Professional | IS3000 A60P6 | No | No | 2012-12-13 | No | No | Reuse | 2013
42805206 | 2012-12-06 | Malfunction |  | It was reported that 5mm monopolar cautery instrument had misaligned threads right out of the box. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of threads are misaligned cannot be confirmed. Instrument was returned with the monopolar adapter piece broken off at the distal end. This piece contains the threads that mate with the disposable cautery tip. Monopolar adapter broke off at the interface with the metal wrist adapter. Conductor wire is also broken as a result of the adapter breakage. Evidence not conclusive, but damage is likely due to mishandling/misuse. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2936562 | 2912365 | 2808881 | 2955842201300157 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-13 | 2013-01-11 | No | Yes | Health Professional | 42805206 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-04 | Malfunction |  | It was reported that prior to starting a da vinci si hysterectomy procedure, the surgical staff reported seeing jaws of the fenestrated bipolar forceps instrument were misaligned. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint misaligned is confirmed. One grip is bent, causing side to side misalignment of grips. There is a. 062 offset at the tips, indicating overloading at the tip. Electrical continuity passed. Additional finding: frayed pitch cable. Instrument was found with frayed pitch cable at distal clevis hub. No damage was found at the clevis. No other cable damage was found. Last finding scratches on maintube. Distal end of main tube has various scratch marks with light material removal. Suggesting the may have been caused by instrument collisions or rough handling. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2936563 | 2912366 | 2808882 | 2955842201300159 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-13 | 2013-01-11 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-13 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff reported seeing wires popped out of the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of wires are popped is confirmed. Both pitch cables are broken at the distal clevis hub. Cable segments that contains the crimp are still installed in clevis. Clevis does not exhibit any damage or wear marks. Other cables at wrist are not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2936564 | 2912367 | 2808883 | 2955842201300164 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-14 | 2013-01-11 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-14 | Malfunction |  | It was reported that prior to starting a da vinci si hysterectomy procedure, the surgical staff reported seeing a frayed cable at the distal end of the tenaculum forceps instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Customer initially reported frayed cable at distal end. However for clarification, the nonconformance was a broken pitch cable. Based on this, the complaint was confirmed. The pitch cable was broken at the distal clevis hub. Cable segment that contains the crimp was missing in clevis. Clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2936565 | 2912368 | 2808884 | 2955842201300163 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-11 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-04 | Malfunction |  | It was reported that during a davinci si bladder augmentation procedure, the customer noted that the wire at tip of the progras forceps instrument broke. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2936573 | 2912377 | 2805554 | 2955842201300165 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-13 | 2013-01-11 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-10-30 | Malfunction |  | Event Description: --/-- It was reported that during a davinci s prostatectomy procedure, the customer noted a broken wire on the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley spins freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional finding not reported by the site was scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | LARGE NEEDLE DRIVER | Reuse | 2931944 | 2907490 | 2804006 | 2955842201300114 | 1 | NAY | Manufacturer | DistributorUser facility | Other | FollowupInitial | 2013-01-17 | 2013-01-10 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-17 | Device Not Retu | No |  | 2013
42030901 | 2012-12-12 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a broken cable on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley spins freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Additional observation not reported by the customer was scratches on distal pulleys. There were scratches found on the surface of both sides of the distal pulleys on the instrument. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2932559 | 2908142 | 2804658 | 2955842201300117 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-08-31 | Malfunction |  | On (b)(4) 2012, intuitive surgical received uf/importer report number (b)(4) from the fda. The event details are provided below: during a da vinci hysterectomy, monopolar curved scissor was being used. End of plastic tip cover accessory burned/melted and arcing of cautery occurred - spark hit uterus. Instrument removed, tip cover accessory put on by physician still saw arcing occur up instrument. New monopolar instrument opened and used. One monopolar had 1 life left and one had 2 lives left. No harm to patient. Medwatch report number (b)(4) was submitted concerning first arcing event. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. On (b)(4) 2012, isi received a report from the site concerning the monopolar curved scissors (mcs) instrument, from part number (b)(4), lot number m10120730 517, as indicated in the site's uf/importer report number (b)(4). Medwatch report (b)(4) was submitted to the fda on (b)(4) 2012 concerning the first arcing event against the mcs tip cover accessory since the customer reported that issue concerned the tip cover accessory. The site also indicated in their uf/importer report that the reported event occurred on (b)(6) 2012; however, at the time of the site's initial report of this event to isi on (b)(4) 2012, it was indicated by the initial reporter that the incident occurred on (b)(6) 2012 and there was no indication from the initial reporter that arcing occurred twice during the surgical procedure. Isi reviewed that site's system log and found that mcs instruments from lot numbers m10120730 517 and m12120614-221 were used during a surgical procedure performed on (b)(6) 2012. The system log also showed that the mcs instrument from lot number m10120730 517 had 2 uses remaining and the mcs instrument from m12120614-221 had 1 use remaining, which would be consistent with the site's complaint description as indicated in the site's uf/importer report. In addition, both of the mcs instruments were manufactured after (b)(4) 2012. Based on isi's investigation findings, isi has made the decision to report this event using the mcs instrument from lot number m12120614-221 as the suspect device. Isi has made several attempts to verify additional details concerning the reported event. However, no additional information has been provided. If additional information becomes available, a follow-up medwatch report will be submitted to the fda. | MONOPOLAR CURVED SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2933015 | 2908542 | 2805058 | 2955842201300118 | 1 | NAY | Manufacturer | Other | RISK MANAGER | Initial | 2012-12-11 | 2013-01-10 | No | Yes | Health Professional | 42017910 | No | No | 2012-12-11 | No | No | Reuse | 2013
41032204 | 2012-12-11 | Malfunction |  | It was reported that during a davinci si sigmoid colectomy procedure, the customer noted frayed cables and they also experienced difficulty to remove the vessel sealing instrument from the trocar. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with broken grip cables located at the snake wrist. The top wrist disc was also dislodged from the snake wrist. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | VESSEL SEALING | ENDOSCOPIC INSTRUMENT | 2933047 | 2908617 | 2805133 | 2955842201300119 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 41032204 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
IS3000 A60P7 | 2012-12-12 | Malfunction |  | Event Description: --/-- It was initially reported that during a davinci si hysterectomy procedure, the bipolar energy device started to fire on its own. According to the information provided to the clinical sales representative, the customer later alleged that while they were setting up the da vinci robotic system for a procedure, the bipolar instrument started to fire on its own. There was no patient in the room and no one was operating the system at the time of the occurrence. The surgeon decided to convert the case to laparoscopic surgery to finish the planned surgery. She did not call in to isi technical support service for assistance. Based on additional information provided to the field service engineer, the surgeon claimed that the bipolar instrument had fired on its own in two separate cases. The surgeon did not report these incidents to intuitive surgical at the time of the occurrences. | Manufacturer Narrative: --/-- (b)(4). --/-- Additional Manufacturer Narrative: --/-- The field service engineer (fse) visited the site to verify that the customer was using the correct electrosurgical unit (esu), valley lab force fx, with the correct energy activation cable (blue valley lab cable). The fse confirmed that the equipment was compatible. Based on the information provided to the fse by the operating room surgical staff, this surgeon is the only surgeon who had experienced this issue. Furthermore, it was stated that the surgeon did not realize she was pressing the foot pedal, activating energy to the bipolar instrument. The case was converted to laparoscopic surgery to successfully complete the procedure. The fse explained to the customer the importance of calling isi's technical support in such instances. The fse also performed a complete system verification and electrical safety test during the preventive maintenance visit. Together with the hospital's biomed personnel, the engineer verified both bipolar and mono-polar functionality with the customer's instruments. The fse was unable to duplicate the customer reported failure. A review of the system logs showed no system malfunction occurring during the procedure on (b)(6) 2012. The fse tested the system setup and found it to be working to manufacturer's specifications with no defect or malfunction. On (b)(4) 2013, isi contacted initial reporter (b)(6). He indicated that he was not aware of any patient injuries related to these, incidents reported. | DA VINCI S SURGICAL SYSTEM | Reuse | 2933305 | 2908786 | 2805302 | 2955842201300120 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | BIOMEDICAL ENGINEER | InitialFollowup | 2013-12-12 | 2013-01-10 | No | Yes | Health Professional | IS3000 A60P7 | No | Yes | 2013-01-15 | Device Not Retu | No |  | 2013
42018304 | 2012-11-28 | Malfunction |  | It was reported that the round plastic piece from the joint on the permanent cautery hook instrument came out and the wiring was loose. No other information was provided. Late submission of this medwatch report is due to a reporting oversight by the responsible complaint handler. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with the conductor cap detached from the distal end. The conductor wire is exposed. The yaw pulley boss feature that mates with the cap is broken off. The instrument passed electrical continuity testing. Engineering evaluation also found that the distal clevis ear on conductor wire side is broken off, resulting in dislodgment of yaw pulley. The broken piece, which was not returned, is roughly. 285 x. 220 in size. The conductor cap likely detached as a result of the clevis ear breakage. The proximal clevis is cracked in the axial direction halfway between the ears on the same side as the distal clevis breakage. The location of the crack is consistent with overloading at the tip with the wrist pitched. Engineering concluded that the breakage is likely due to overloading at the tip or other mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2933357 | 2908891 | 2788457 | 2955842201300121 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-28 | 2013-01-10 | No | Yes | Health Professional | 42018304 | Device Returned To Manufacturer | No | 2012-12-28 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-14 | Malfunction |  | It was reported prior to starting a da vinci si procedure the tenaculum forceps instrument was noted to have a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a frayed grip cable. The grip cable is frayed at the distal idler pulley. The frayed strands stick out at the wrist. The other cables at wrist are not damaged. Additionally, scratches on the maintube were observed on the distal end, though not reported. Engineering concluded that the damage was likely due to mishandling. The instruments and accessories user manual specifically states: general precautions and warnings o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2933407 | 2908964 | 2805480 | 2955842201300123 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-10 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-11-12 | Malfunction |  | It was reported prior to starting a da vinci s surgical procedure, the maryland bipolar instrument could not be recognized. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. The instrument was checked for recognition on an in-house is3000 system. The is3000 system successfully recognized instrument, showing 3 uses left. Pogo pins do not stick and are not contaminated. Dallas chip programmer (dcp) verification of instrument passed. Engineering was unable to replicate recognition issue. Additional observation not reported by site is a broken cable. The pitch up cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Other cables at wrist are not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2933424 | 2908980 | 2805496 | 2955842201300122 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-11 | 2013-01-10 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-11 | Device Not Retu | No | Reuse | 2013
42018406 | 2012-12-05 | Malfunction |  | It was observed during central processing that the permanent cautery spatula instrument had a chipped shaft. It is unknown when the chipped piece fell off of the instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2933447 | 2909003 | 2805519 | 2955842201300128 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42018406 | No | Yes | 2012-12-12 | No | No | Reuse | 2013
42020505 | 2012-11-14 | Malfunction |  | It was reported that prior to starting a da vinci si hysterectomy procedure, the surgical staff reported that the fenestrated bipolar forceps was not recognized by the system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of instrument not recognized by the system was not confirmed. Instrument was found with frayed pitch cable at distal clevis hub. No damage was found at the clevis. Frayed segment is approximately 0. 03 in length. No other cable damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2933448 | 2909004 | 2805520 | 2955842201300127 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-07 | 2013-01-10 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-07 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-04 | Malfunction |  | Event Description: --/-- It was reported that during a da vinci si hysterectomy procedure, the surgical staff reported seeing a broken cable on the pk dissecting forceps instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. The customer initially reported broken wire, however, for clarification the damaged instrument component was a frayed pitch cable. Based on this clarification, the customer reported complaint is confirmed. Instrument was found with frayed pitch cable at distal clevis hub. No damage was found at the clevis. No other cable damage was found. Electrical continuity passed. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. The customer initially reported broken wire, however, for clarification the damaged instrument component was a frayed pitch cable. Based on this clarification, the customer reported complaint is confirmed. Instrument was found with frayed pitch cable at distal clevis hub. No damage was found at the clevis. No other cable damage was found. Electrical continuity passed. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | Reuse | 2933449 | 2909005 | 2805521 | 2955842201300126 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-07 | 2013-01-10 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-07 | Device Not Retu | No |  | 2013
42027804 | 2012-12-12 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, during set up, the surgical staff reported seeing a broken cable on the grasping retractor instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of broken cable is confirmed with the following clarification: cable is frayed. One grip open cable is frayed at the grip hub. Frayed strands stick out at the hub. Other cables at wrist are not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | GRAPTOR RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2933470 | 2909006 | 2805531 | 2955842201300130 | 1 | NAY | Manufacturer |  | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-11 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff reported seeing a broken cable on the maryland bipolar forceps instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2933456 | 2909007 | 2805523 | 2955842201300133 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42017207 | No | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-10 | Malfunction |  | It was observed during central processing that the pk dissecting forceps instrument had a frayed wire at the wrist. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint wire frayed at wrist is confirmed. Instrument was found with frayed pitch cable at distal idler. No damage found on pulley and clevis. Electrical continuity passed. Additional finding: scratches on maintube. Distal end of main tube has various scratch marks with light material removal. Scratches are short in length and not aligned with the tube axis, suggesting the may have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2933472 | 2909008 | 2805524 | 2955842201300134 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-13 | 2013-01-10 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-12 | Malfunction |  | It was reported prior to starting a da vinci si surgical procedure that the maryland bipolar forceps instrument was noted to have broken wires. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity failed. Additionally, a broken conductor wire was observed. One conductor wire is broken at the yaw pulley exit with a missing piece of yaw pulley measuring roughly about 0. 050 x 0. 069. The wire is detached from its connection at the grip. The yaw pulley shows no signs of arcing. Also, bent tips were observed. One grip is bent, causing side to side misalignment of the grips. There is a. 032 offset at the tips. Engineering concluded that damage was likely due to mishandling. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. Based on this, no additional follow up is required. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2933459 | 2909011 | 2805527 | 2955842201300135 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-11-05 | Malfunction |  | Event Description: --/-- It was reported during a da vinci si hysterectomy with bilateral salpingo-oophorectomy the fenestrated bipolar forceps instrument broke at the wrist. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint engineering evaluation found that the instrument cable was broken. The pitch cable was broken at the distal clevis hub. The cable segment that contained the crimp still remained in the clevis. The clevis did not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | Reuse | 2933443 | 2909012 | 2805528 | 2955842201300124 | 1 | NAY | Manufacturer | ForeignDistributor | Other | FollowupInitial | 2013-01-30 | 2013-01-10 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-30 | Device Not Retu | No |  | 2013
42030901 | 2012-12-14 | Malfunction |  | It was reported during a da vinci si hysterectomy procedure that while suturing the last suture the megasuturecut needle driver instrument wire broke. Upon visual inspection for fragments, none were found. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. Additionally it was observed there were scratches on the surface of the distal pulley. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2933444 | 2909013 | 2805529 | 2955842201300125 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-10 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
42027804 | 2012-12-12 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, during set up, the surgical staff reported seeing a broken cable on grasping retractor instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of instrument cable broken was confirmed. The pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still remained in the clevis. Clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | GRAPTOR RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2933455 | 2909015 | 2805531 | 2955842201300129 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42027804 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-12 | Malfunction |  | It was reported that prior to starting a da vinci si cholecystectomy procedure, the surgical staff reported seeing a wire sticking out of the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of wires sticking out at end is confirmed. The pitch up cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Pitch down cable is frayed at the distal clevis hub. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2933457 | 2909016 | 2805532 | 2955842201300132 | 1 | NAY | Manufacturer |  | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-12 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, during set up, the surgical staff reported seeing a broken cable on the pk dissecting forceps instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint broken cable is confirmed. The pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Electrical continuity passed. Additional finding is scratches on both distal pulleys. Both distal pulleys have scratches on the surface. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2933471 | 2909017 | 2805533 | 2955842201300131 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-05 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the tenaculum forceps instrument wires broke. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed both pitch cables being broken at the distal clevis hub. The cable segments that contain the crimp are still installed in the clevis. The clevis does not exhibit any damage or wear marks. The other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS | ENDOSCOPIC INSTRUMENT | 2933458 | 2909018 | 2805534 | 2955842201300136 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci si surgical procedure the prograsp forceps instrument was noted to have a broken cable. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable being broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The cable is broken inside the housing by the clamping pulley. Additionally, scratches on the maintube were observed. The distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. Based on this, no additional follow up is required. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS | ENDOSCOPIC INSTRUMENT | 2933461 | 2909019 | 2805535 | 2955842201300137 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42014304 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci si surgical procedure the maryland dissector instrument was noted to have a broken cable. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed one drive cable being broken at the distal end of the snake wrist. The cable segment sticks out from the snake wrist and the wrist is unable to properly straighten due to the cable breakage. The pitch/yaw motion is no longer smooth. The other wrist cables are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND DISSECTOR | ENDOSCOPIC INSTRUMENT | 2933475 | 2909033 | 2805549 | 2955842201300138 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42014304 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-04 | Malfunction |  | It was reported that during a da vinci si bladder augmentation procedure the wire on the tip of the prograsp forceps instrument broke. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken pitch cable. The pitch cable was broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. Additionally, insulation damage to the main tube was observed, though not reported by the site. Material is missing. The surface of the main tube appeared to be scraped off. Engineering concluded that damage was likely due to mishandling. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. Based on this, no additional follow up is required. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS | ENDOSCOPIC INSTRUMENT | 2933465 | 2909038 | 2805554 | 2955842201300139 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-13 | 2013-01-10 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-10-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the maryland bipolar forceps instrument was not holding tissue. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken pitch cable and scratches and indentation on the distal pulleys. The pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity passed. Additionally, scratches and indentations on the distal pulleys were observed, though not reported by the site. Both of the pulleys have scratches on the surface and indentations on the edges. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. Based on this, no additional follow up is required. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2933481 | 2909039 | 2805555 | 2955842201300140 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-08 | 2013-01-10 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-08 | Device Not Retu | No | Reuse | 2013
42004906 | 2012-12-10 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci si surgical procedure the cadiere forceps instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering observed scratches on the main tube not reported by site. The distal end of the main tube has various scratch marks with light material removal. Engineering concluded that damage was likely due to mishandling. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- According to additional information received from the reporter of the complaint on (b)(4) 2013, no patient has returned with post-surgical complications. He indicated having no knowledge of missing main tube insulation. No further information was provided. | CADIERE FORCEPS | Reuse | 2933486 | 2909046 | 2805562 | 2955842201300141 | 1 | NAY | Manufacturer | User facility | Other | FollowupInitial | 2013-01-11 | 2013-01-10 | No | Yes | Health Professional | 42004906 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No |  | 2013
42017207 | 2012-12-06 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci si procedure the maryland bipolar forceps instrument was not recognized by the system. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- According to additional information received from the reporter of the complaint on (b)(4) 2013, no pieces fell into the patient and no injuries had occurred. It was also stated that the instruments are always inspected prior to start of a procedure. --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering observed bent bipolar pins and scratches on the maintube. Both pins are bent to the side and loose. The distal end of the main tube has various scratch marks with light material removal. Electrical continuity passed. Engineering concluded that damage was likely due to mishandling. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS | Reuse | 2933487 | 2909048 | 2792389 | 2955842201300142 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2013-01-11 | 2013-01-10 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No |  | 2013
42034301 | 2012-11-05 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure the thoracic grasper instrument was not recognized by the system. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the main tube insulation appears to have scratches and gouge marks with material removed. Engineering concluded that damage was likely due to mishandling. On (b)(4) 2013 the initial reported of this complaint was contacted regarding failure analysis findings. It was indicated that no fragments were observed falling into the patient and no adverse events had occurred. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | THORACIC GRASPER | ENDOSCOPIC INSTRUMENT | 2933510 | 2909051 | 2805567 | 2955842201300143 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-12 | 2013-01-10 | No | Yes | Health Professional | 42034301 | Device Returned To Manufacturer | Yes | 2012-12-12 | Device Not Retu | No | Reuse | 2013
UNKNOWN | 2010-11-03 | Injury | OtherRequired Intervention | Event Description: --/-- According to allegations made in a lawsuit, it was reported that the patient sustained an infection and vesicovaginal fistula post a da vinci hysterectomy and gall bladder procedure. The patient continues to suffer from cramping and bladder problems. Throughout this time period, the patient has reportedly been unable to maintain normal intimate relationships and has suffered emotional distress. The patient had to have added medical tests and additional surgery and has suffered pain, loss of function, emotional distress, and permanent injury. | Manufacturer Narrative: --/-- Intuitive surgical contacted the site to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. A follow-up mdr will be submitted if pertinent additional information is received. --/-- Additional Manufacturer Narrative: --/-- There was no patient or hospital name/contact information provided for the reported event. Therefore, a follow up on the alleged patient's condition and event description could not be made. Should additional information become available, a follow up mdr will be sent. | DA VINCI SURGICAL SYSTEM | Unkown | 2933497 | 2909052 | 2805568 | 2955842201300115 | 1 | NAY | Manufacturer | Other | Other | InitialFollowup | 2013-01-18 | 2013-01-10 | Yes | Yes | Health Professional | UNKNOWN | No | No | 2013-08-07 | No | No |  | 2013
IS2000 A51P8 | 2010-12-17 | Injury | OtherRequired Intervention | According to allegations made in a lawsuit, the patient experienced intra-abdominal bleeding, thermal burning, severe infection with intra-abdominal abscesses, and septic shock after a da vinci hysterectomy procedure. The patient continues to suffer from chronic abdominal pain and severe bowel issues. The patient has been unable to maintain intimate relationships and suffers emotional distress. Allegedly, the patient underwent multiple additional medical tests/procedures and surgery. | There was no patient or hospital name/contact information provided for the reported event. Therefore, a follow up on the alleged patient's condition and event description could not be made. Should additional information become available, a follow up mdr will be sent. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC ELECTROSURGICAL INS | 2933515 | 2909077 | 2805593 | 2955842201300116 | 1 | NAY | Manufacturer | Other | Health Professional | InitialFollowup | 2012-12-14 | 2013-01-10 | Yes | Yes | Health Professional | IS2000 A51P8 | No | Yes | 2013-05-28 | No | No | Reuse | 2013
42030901 | 2012-11-28 | Malfunction |  | It was reported that during a davinci si surgical procedure, the mega suture cut needle driver instrument was not recognized by the davinci robot system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to replicate the customer reported failure mode of instrument not being recognized by the davinci robot system. Instrument was checked for recognition using an in-house davinci robot system. The instrument was successfully recognized, showing 1 use left. The pogo pins did not stick and were not contaminated. A finding not reported by the customer was a broken grip cable. One grip close cable was found to be broken at the distal idlers. The idler pulley spins freely and it was not damaged. The cable segment sticks out at the wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2927912 | 2902710 | 2797131 | 2955842201300100 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-11 | 2013-01-09 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-11 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-11 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, during set up the customer noted exposed wires on the tenaculum forceps instruments, the instrument was not used in the case. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip open cable was found to be frayed at the grip hub. The frayed strands stick out at the wrist. Other cables at wrist were not damaged. The grip hub also exhibited scratch marks adjacent to the fraying. Evidence not conclusive, but engineering concluded that the fraying and scratch marks were likely caused by mishandling / misuse. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS | ENDOSCOPIC INSTRUMENT | 2927888 | 2902717 | 2799233 | 2955842201300099 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-11 | 2013-01-09 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | Yes | 2012-12-11 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-10 | Malfunction |  | Event Description: --/-- It was reported that after a da vinci si cystectomy procedure, a broken cable at the distal end of the tenaculum forceps instrument was discovered. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering found the instrument had a broken cable at the distal end. The instrument was found with a frayed pitch cable at the distal clevis hub. No damage was found at the clevis. The frayed segment was approximately 0. 02 in length. No other cable damage was found. --/-- Additional Manufacturer Narrative: --/-- The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | TENACULUM FORCEPS | Reuse | 2928820 | 2904106 | 2800622 | 2955842201300101 | 1 | NAY | Manufacturer | User facility | Other | InitialFollowup | 2012-02-22 | 2013-01-09 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-02-22 | Device Not Retu | No |  | 2013
42020505 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the cable on the fenestrated bipolar forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one conductor wire is broken at the yaw pulley exit. The wire is detached from its connection at the grip. The instrument failed electrical continuity testing. The yaw pulley exhibits no signs of arcing. Engineering evaluation also found that the pitch cable at distal idler is frayed. The pulley and clevis exhibit no damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2929214 | 2904536 | 2801052 | 2955842201300102 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-13 | 2013-01-09 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-13 | Device Not Retu | No | Reuse | 2013
42017403 | 2012-12-14 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the tip on the harmonic curved shears instrument was observed to be missing. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the harmonic curved shears instrument will only have a tip if there is a harmonic curved shear insert accessory installed. Engineering evaluation also found that one of the housing snap tabs is broken. The housing can be removed as a result of damage. The tube is also slightly and it does not align straight with a ruler and the tube does not rotate 360 degrees. Engineering has concluded that the damage to the housing and the tube is likely due to mishandling. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings,, o handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | HARMONIC CURVED SHEARS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2929619 | 2904928 | 2801424 | 2955842201300103 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-09 | No | Yes | Health Professional | 42017403 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
40027202 | 2012-12-13 | Malfunction |  | It was reported that during set-up of a da vinci si surgical procedure, the handle on the harmonic ace insert accessory fell off. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode, with the following clarification. The teflon pad broke off the clamp arm completely. The teflon pad was not returned. The insert installed on an in-house harmonic ace curved shears instrument failed a generator test sequence due to debris on the base of the curved bladed. No other physical damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | HARMONIC ACE INSERT ACCESSORY | ENDOSCOPIC ELECTROSURGICAL ACC | 2929629 | 2904938 | 2801454 | 2955842201300104 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2013-01-03 | 2013-01-09 | No | Yes | Health Professional | 40027202 | Device Returned To Manufacturer | No | 2013-01-03 | Device Not Retu | No | Unkown | 2013
42017207 | 2012-12-11 | Malfunction |  | It was reported that during reprocessing of the maryland bipolar forceps instrument it was observed that the wire on the instrument was broken. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp remains in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2929829 | 2905151 | 2801667 | 2955842201300105 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-12 | 2013-01-09 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-14 | Malfunction |  | It was reported prior to starting a da vinci si procedure that the prograsp forceps instrument had frayed wires at the distal end. The instrument was not used. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed frayed wires at the distal end. For clarification, they were broken pitch cables. Both pitch cables were broken at the distal clevis hub. Both cable segments that contain the crimp were still installed in clevis. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2929934 | 2905241 | 2801757 | 2955842201300106 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-09 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
42027802 | 2012-12-14 | Malfunction |  | It was reported prior to starting a da vinci si myomectomy procedure that the grasping retractor had a broken cable. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp, one is missing in clevis. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2930069 | 2905380 | 2801896 | 2955842201300107 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-14 | 2013-01-09 | No | Yes | Health Professional | 42027802 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
40018012 | 2012-11-14 | Malfunction |  | It was reported that during a da vinci si axillary lymph node dissection procedure, the surgical staff reported seeing arcing coming from the 8mm mcs tip cover accessory used on the monopolar curved scissors instrument. Other instruments used in the case were the cadier forceps instrument and the fenestrated bipolar instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. Additional information has been requested from the site, however to this date no further information has been received. If additional information is received, a follow-up mdr will be submitted. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2930300 | 2905612 | 2802128 | 2955842201300108 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-09 | No | Yes | Health Professional | 40018012 | No | No | 2012-12-10 | No | Yes | Reuse | 2013
42017207 | 2012-12-11 | Malfunction |  | It was reported that during central processing the mega suturecut needle driver instrument was noted to have a broken cable. Nothing was reported having fallen into the patient during a surgical procedure. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The pitch down cable is frayed at the distal clevis hub. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2930443 | 2905722 | 2802238 | 2955842201300110 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-11 | 2013-01-09 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-11 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-11 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the mega suturecut needle driver instrument was unresponsive, as there was a loose wire. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken grip cable. One grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2930420 | 2905724 | 2802240 | 2955842201300111 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-11 | 2013-01-09 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-11 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-11 | Malfunction |  | It was reported that during central processing the mega suturecut needle driver instrument was noted to have a broken wire. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the one grip close cable is broken at the distal idler pulley. The idler pulley spins freely and has damage on the surface and the edge of the pulley. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2930440 | 2905732 | 2802248 | 2955842201300109 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-11 | 2013-01-09 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-11 | Device Not Retu | No | Reuse | 2013
40030901 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure the mega suturecut needle driver instrument was noted to have frayed wires. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the damaged instrument component was a broken grip cable. The one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. Additionally, it was found that the grip cable was derailed. The grip close cable is derailed at the distal idler pulley. The cable derailment is likely due to cable losing contact with pulley during the wrist articulation. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2930454 | 2905739 | 2802255 | 2955842201300112 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-12 | 2013-01-09 | No | Yes | Health Professional | 40030901 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-12 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure the large needle driver instrument was noted to have a wire sticking out from the tip. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable was loose at distal clevis hub. Upon housing removal, the pitch cable was found broken at the idler block. Additional observation found clamping pulley corroded. Engineering concluded that damage was likely due to mishandling. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2930455 | 2905743 | 2802259 | 2955842201300113 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-12 | 2013-01-09 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-05 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a broken cable on the mega suture cut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip cable was found to be broken at the distal idlers. The idler pulley spins freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist are not damaged. Additional observation not reported by site was scratches on the distal pulley. There are scratches on the surface of one of the distal pulley. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2926731 | 2901577 | 2798093 | 2955842201300088 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-11 | 2013-01-08 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-10 | Malfunction |  | It was reported that during a davinci si partial nephrectomy procedure, the customer reported that a piece of the wire was sticking out of fenestrated bipolar instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in the clevis. The clevis did not exhibit any damage or wear marks. Instrument passed electrical continuity test. Engineering also observed indentations and scratches on the distal pulleys. Both distal pulleys appeared to have scratches and one distal pulley also exhibited indentations on the pulley's edge. Additional observation not reported by site was tube damage. The distal end of main tube had various scratch marks with light material removal. The scratches were short in length and were not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2927076 | 2901975 | 2798490 | 2955842201300089 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-11 | Device Not Retu | No | Reuse | 2013
420227 | 2012-12-03 | Other | Other | It was reported that during a davinci si hysterectomy procedure, the 'bipolar instrumnet was sticking' the surgeon made the decision to complete the planned surgical procedure using traditional open surgical techniques. | Isi contacted surgeon dr (b)(6) to request additional information regarding the reported complaint. Based on the information provided by dr (b)(6), the bipolar instrument got stuck on to a vessel and tore it open. The sticking occurred after the vein was clamped and cauterized, upon opening the mouth of the bipolar instrument, the instrument was sticking to the vein and it caused bleeding to the patient. The surgeon was able to control the patient's bleeding in the operating room with no further complications. The surgeon made the decision to complete the planned surgical procedure using open traditional surgical techniques. The patient made a full recovery from procedure, with no further complications. The site will not be returning the bipolar instrument at this time. A follow-up medwatch report will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. As of (b)(4) 2013, there have been no reported recurrences of this event. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2927505 | 2902352 | 2798867 | 2955842201300090 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Physician | Initial | 2012-12-11 | 2013-01-08 | Yes | Yes | Health Professional | 420227 | No | Yes | 2012-12-11 | No | No | Unkown | 2013
42020505 | 2012-12-14 | Malfunction |  | It was reported that during reprocessing the fenestrated bipolar forceps instrument were noted to have a frayed cable. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp still remained in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2927570 | 2902356 | 2798996 | 2955842201300091 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-14 | 2013-01-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-14 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-11-26 | Malfunction |  | It was reported that during reprocessing of the prograsp forceps instrument, it was observed that the wire at the tip of the instrument was broken. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a frayed grip cable. The grip cable is frayed at the distal idler pulley and also by the proximal clevis. The frayed strands stick out at the wrist. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches are short in length and not aligned with the tube axis, suggesting that the damage to the tube was caused by instrument collisions or rough handling. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2927581 | 2902382 | 2798898 | 2955842201300092 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-11 | 2013-01-08 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | Yes | 2012-12-11 | Device Not Retu | No | Reuse | 2013
42019408 | 2012-12-03 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the cable on the mega needle drive instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. Engineering evaluation also found scratches and dents on the surfaces of both pulleys. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2927648 | 2902417 | 2798952 | 2955842201300093 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-12 | 2013-01-08 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-12 | Malfunction |  | Event Description: --/-- It was reported that during reprocessing of the fenestrated bipolar forceps instrument, it was observed that the cable on the instrument was broken. There were no missing or fallen pieces reported. | Manufacturer Narrative: --/-- --/-- Additional Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable is frayed at the distal clevis hub. The frayed strands stick out at the wrist. The other cables at the wrist are not damaged. Engineering evaluation also found that the distal end of the main tube has various scratch marks with light material removal. The scratches are. 100 -. 210 in length and not aligned with the tube axis. Engineering concluded that the damage to the main tube is likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | Reuse | 2927651 | 2902451 | 2798947 | 2955842201300094 | 1 | NAY | Manufacturer | User facilityHealth Professional | Other | FollowupInitial | 2013-01-08 | 2013-01-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No |  | 2013
42020505 | 2012-12-12 | Malfunction |  | It was reported that during reprocessing of the fenestrated bipolar forceps instrument, it was observed that the wire on the instrument was sticking out. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still remained in the clevis. The clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2927654 | 2902480 | 2798996 | 2955842201300095 | 1 | NAY | Manufacturer |  | Other | Initial | 2012-12-12 | 2013-01-08 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-11 | Malfunction |  | It was reported that during reprocessing of the large needle driver instrument, it was observed that the wire on the instrument was broken. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that both of the pitch cables are frayed at the distal clevis hub. The frayed strands stick out at the wrist. The other cables at the wrist are not damaged. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2927664 | 2902495 | 2799011 | 2955842201300096 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-12 | 2013-01-08 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-07 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a 'loose or broken cable' on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2927908 | 2902712 | 2794949 | 2955842201300097 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-11 | 2013-01-08 | No | Yes | Health Professional | 42022703 | No | Yes | 2012-12-11 | No | No | Reuse | 2013
42030901 | 2012-12-11 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted frayed wires on the mega suture cut needle driver instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. Frayed segment is approximately 0. 03 in length. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2927887 | 2902715 | 2799231 | 2955842201300098 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-11 | 2013-01-08 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-11 | Device Not Retu | No | Reuse | 2013
42018406 | 2012-12-04 | Malfunction |  | It was reported that during a davinci si prostatectomy procedure, the customer noted a broken cable on the permanent cautery spatula instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2923430 | 2898206 | 2794722 | 2955842201300062 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-07 | 2013-01-07 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-12-07 | Device Not Retu | No | Reuse | 2013
40018012 | 2012-12-10 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, 'the mcs tip cover started sliding off from monopolar curved scissors instrument. Nothing fell in the patient and no patient injury. '. | The instrument was returned and evaluated. Engineering was unable to confirm the reported failure mode of 'mcs tip cover started sliding off from the monopolar curved scissors instrument'. The mcs tip cover was visually inspected and did not exhibit tears or arcing signs. The mcs tip cover was installed on the in-house testing monopolar curved scissors (mcs) instrument, the mcs instrument was placed on an in-house davinci robot system and it was driven without any problems. The mcs tip cover did not slide off of the mcs instrument during normal wrist articulation. Additional internal testing was able to recreate the mcs tip cover slippage by over installing the mcs tip cover past the hard stop on the mcs instrument tube extension. During removal of mcs instrument from the cannula with the mcs tip cover over installed, the mcs tip cover interfered with the cannula edge causing the mcs tip cover to slide off. Engineering was unable to recreate the mcs tip cover slide off without any external forces applied to mcs tip cover. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2923505 | 2898293 | 2794809 | 2955842201300064 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 40018012 | Device Returned To Manufacturer | No | 2012-12-10 | No | Yes | Unkown | 2013
IS2000 A5P8 | 2012-12-06 | Injury |  | It was reported that while undocking after a da vinci s hysterectomy procedure, the surgeon noticed charring at one of the port incision sites. The surgeon excised a small area of the affected port site, suture closed the incision and applied a dry dressing. The excised sample was sent to the hospital's pathology department to be analyzed. The robotics coordinator present during the surgical procedure indicated that there was no harm to the patient during the surgical procedure and the system did not fault during the planned surgical procedure. | The investigation conducted by field service engineering (fse) was unable to replicate the issue experienced by the site. Functional testing of the site's da vinci s surgical system found that the system functioned within specification and the system also passed an electrical safety test. On (b)(4) 2012, isi contacted (b)(6) at (b)(6). (b)(6) indicated that an monopolar curved scissors (mcs) instrument was used in the patient side manipulator of the affected port site. (b)(6) indicated that a valley lab force fx electrical surgical unit (esu) was used during the surgical procedure and the setting was 30 coag. (b)(6) indicated that the cannula and mcs instrument were inspected prior to use and there was no damage noted. (b)(6) indicated that arcing during the surgical procedure was not observed. (b)(6) indicated that the ground pad was positioned on the patient's left anterior thigh and inspection of the grounding pad was inspected and there were no issues found. On (b)(6) 2012, isi received an email response from (b)(6). (b)(6) indicated that the surgeon who performed the surgical procedure indicated to him that the pathology report confirmed that the patient sustained a thermal burn at the port site. Per (b)(6), the surgeon indicated that the patient is healing fine and without incident. (b)(6) indicated that it is the surgeon's belief that the thermal burn sustained by the patient was caused by a leak of electrical current from the mcs instrument as it was being activated from the surgeon side console during routine tissue cauterization. The mcs instrument and mcs tip cover accessory were returned to isi for evaluation. Engineering was unable to confirm the customer reported issue. The mcs instrument exhibited slight char marks at the scissor tips; however, no arcing damage to the tube extension nor on the clevis was found. The tip cover accessory used in conjunction with the mcs instrument exhibited no physical damage. The mcs instrument failed cut testing. As of (b)(4) 2012, there have been no reported recurrences of this issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2923536 | 2898303 | 2794819 | 2955842201300063 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-06 | 2013-01-07 | Yes | Yes | Health Professional | IS2000 A5P8 | Yes | Yes | 2012-12-06 | Device Not Retu | No | Reuse | 2013
42020706 | 2012-12-04 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff reported seeing a broken wire on the tenaculum forceps instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of broken wire is confirmed. One grip close cable is broken at the distal idlers. Idler pulley spins freely and does not exhibit damage. Cable segment sticks out at wrist. Other cables at wrist are not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | TENACULUM FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2923596 | 2898360 | 2786326 | 2955842201300075 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-11-14 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgical staff reported the mega suturecut needle driver instrument stopped responding to the surgeon. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint stopped responding is confirmed. The pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2923581 | 2898371 | 2794887 | 2955842201300069 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-05 | 2013-01-07 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-05 | Malfunction |  | It was reported that during a da vinci si herna repair procedure, the surgical staff reported seeing a broken wire at the distal end of the mega suturecut needle driver instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint wires broke is confirmed. One grip close cable is broken at the distal idlers. Idler pulley spins freely and was not damaged. Cable segment sticks out at wrist. Other cables at wrist are not damaged. No other damage found. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2923610 | 2898376 | 2794892 | 2955842201300066 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-05 | 2013-01-07 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | Yes | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42018306 | 2012-12-10 | Malfunction |  | It was observed during central processing that the permanent cautery hook instrument had a wire sticking out. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of wire sticking out is confirmed. Drive cables at wrist are intact and undamaged. A segment of the conductor wire has become dislodged from the conductor cap and is exposed on the outside of the yaw pulley. Wire insulation is not damaged and electrical continuity passes. Additional observation not reported by site is a broken conductor cap. One side of cap has a broken off piece, causing wire to be exposed. Top of the cap exhibits cracking. Evidence not conclusive, but damage is likely due to mishandling/misuse. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2923611 | 2898377 | 2794893 | 2955842201300074 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-07 | Malfunction |  | It was observed during central processing that the pk dissecting forceps instrument had frayed wires. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Customer initially reported frayed wires. However, for clarification, the nonconformance was a frayed pitch cable. Based on this, the complaint was confirmed. Frayed pitch cable was found at distal clevis hub. No damage was found at the clevis. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2923619 | 2898385 | 2794949 | 2955842201300071 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-11-26 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgical staff reported that the pk dissecting forceps would not rotate. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint would not rotate is confirmed. The pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Electrical continuity passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2923624 | 2898390 | 2794906 | 2955842201300068 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-05 | 2013-01-07 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-08-31 | Malfunction |  | On (b)(6) 2012, medwatch uf/importer report (b)(4) was received: event desc: the patient was undergoing a robotic total laparoscopic hysterectomy with bilateral salpingectomy. The wires on the instrument broke. The fragments were retrieved from the patient's abdomen. Device usage problem: device failed (e. G. Broke, couldn't get it to work or stopped working). | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ELECTROSURGICAL INSTRUMENT | 2923626 | 2898412 | 2794928 | 2955842201300073 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42020505 | No | No | 2012-12-10 | No | No | Reuse | 2013
42030901 | 2012-12-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the surgical staff reported seeing a broken cable at the tip of the mega suturecut needle driver instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | Customer initially reported broken cable. However for clarification, the nonconformance was a frayed pitch cable. Based on this, the complaint was confirmed. Frayed grip cable was found at the distal idlers. Frayed segment sticks out at wrist. Idler pulley spins freely. Idler pulley rim shows mechanical indentation damage. Evidence not conclusive, but damage may have been to misuse. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2923628 | 2898414 | 2792318 | 2955842201300065 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-05 | 2013-01-07 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-05 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the surgical staff reported seeing a broken band near the tip of the pk dissecting forceps instrument. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint band near tip is broken is confirmed. The pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Electrical continuity passed. Additional finding: scratches on maintube. Distal end of main tube has various scratch marks with light material removal. Scratches are short in length and not aligned with the tube axis, suggesting the may have been caused by instrument collisions or rough handling. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2923653 | 2898421 | 2794937 | 2955842201300070 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-06 | 2013-01-07 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-06 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-03 | Malfunction |  | It was reported during a da vinci s sacrocolpopexy surgery that the maryland bipolar forceps instrument was broken and unable to be removed from the trocar. Nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed the main tube is broken at the proximal side. Engineering concluded the damage likely due to mishandling. Electrical continuity passed. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2923642 | 2898429 | 2794945 | 2955842201300012 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-03 | 2013-01-07 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | Yes | 2012-12-03 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-04 | Malfunction |  | It was observed during central processing that the pk dissecting forceps instrument had frayed wires. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Complaint of instrument with frayed wires was confirmed. Instrument was found with frayed cable at distal idler pulley. Frayed cable section is approximately 0. 12 in length, with slight damage found on the edge of the pulley. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2923667 | 2898433 | 2794949 | 2955842201300072 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42805206 | 2012-12-05 | Malfunction |  | It was reported that after cholecystectomy procedure, the surgical staff reported the hook of the monopolar cautery instrument broke off while they attempted to remove the hook. The piece was discarded and nothing was reported having fallen into the patient. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | 5MM MONOPOLAR CAUTERY INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2923687 | 2898457 | 2794972 | 2955842201300067 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-05 | 2013-01-07 | No | Yes | Health Professional | 42805206 | No | No | 2012-12-05 | No | No | Unkown | 2013
42030901 | 2012-12-10 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted frayed wires at the tip of the mega suture cut needle driver instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed a broken grip cable at the distal idlers. The cable segment sticks out at wrist, and the idler pulley was able to spin freely. The idler pulley face and the rim exhibited damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2925014 | 2899890 | 2796406 | 2955842201300076 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-10 | Malfunction |  | Event Description: --/-- It was reported that prior to starting a da vinci s hysterectomy surgical procedure, the customer noted a broken cable on the fenestrated bipolar instrument. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with conductor wire insulation frayed and conductor wire exposed. Instrument passed electrical continuity test. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | FENESTRATED BIPOLAR FORCEPS | Reuse | 2925219 | 2900098 | 2796613 | 2955842201300077 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | FollowupInitial | 2013-01-24 | 2013-01-07 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2013-01-24 | Device Not Retu | No |  | 2013
42030901 | 2012-12-10 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted frayed wires at the tip of the mega suture cut needle driver instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the grip cable was found to be frayed at the distal idlers. The frayed grip cable became derailed at the distal idler pulley. The grips were able to open and close, but their movement and functionality could not be precise. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2925293 | 2900159 | 2796675 | 2955842201300078 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42000305 | 2012-12-07 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the wire on the clip applier instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | SMALL CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2925629 | 2900435 | 2796950 | 2955842201300079 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42000305 | No | No | 2012-12-10 | No | No | Unkown | 2013
42000606 | 2012-12-04 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted a frayed cable on the large needle driver instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that a broken grip cable was found at the distal idlers. The cable segment sticks out at wrist. The idler pulley was able to spin freely. The idler pulley rim exhibited damage. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2925678 | 2900473 | 2796989 | 2955842201300080 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | Yes | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-07 | Malfunction |  | It was reported that during a da vinci si sacrocolpopexy procedure, the cables on the megasuturecut needle driver instrument snapped. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers pulley. Idler pulley spins freely but has damages on the edge and on the surface. Cable segment sticks out at wrist. The other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2925724 | 2900500 | 2792360 | 2955842201300082 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-11 | 2013-01-07 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-11 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted a broken wire on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2925725 | 2900502 | 2797068 | 2955842201300083 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-09-10 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, the inner thread of the large needle drive instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode with the following clarification. The instrument was returned with a broken grip cable. The idler pulley spins freely and exhibits damage to the face. The cable segment sticks out at the wrist. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2925707 | 2900505 | 2797020 | 2955842201300084 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-12-11 | 2013-01-07 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2012-12-11 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-11 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, the tips on the fenestrated bipolar forceps instrument were not aligned. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip is bent, causing side to side misalignment of grips. There is a. 080 offset at the tips. The bent grip does not exhibit separation of the yaw pulley and pulley cover at the glue joint. Engineering concluded that the damage is likely due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Engineering evaluation also found that the pitch down cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The other cables at wrist are not damaged. No other damage was found. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2925718 | 2900517 | 2797033 | 2955842201300081 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-19 | 2013-01-07 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | Yes | 2012-12-19 | Device Not Retu | No | Initial | 2013
42009308 | 2012-12-07 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the customer noted a broken wire on the prograsp forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was found with frayed pitch cable at distal clevis hub. No damage was found at the clevis. No other cable damage was found. Additional observation not reported by site was main insulation tube damage. The distal end of main tube had various scratch marks with light material removal. The scratches were short in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling. No other damage was found. Intuitive surgical contacted the initial reporter of this complaint to obtain additional information concerning the reported event; however, no additional information has been provided as of the date of this report. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. The endowrist instruments instructions for use (ifu) specifically states: general precautions and warnings handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2925736 | 2900533 | 2797068 | 2955842201300085 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-11-27 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted a broken wire on the pk dissecting forceps instrument. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2925747 | 2900560 | 2791677 | 2955842201300086 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42022703 | No | No | 2012-12-10 | No | No | Unkown | 2013
42030901 | 2012-11-28 | Malfunction |  | It was reported that during a davinci si surgical procedure, while using the mega suture cut needle driver instrument, the instrument was not being recognized by the davinci robot system. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering was unable to duplicate the customer's reported complaint of instrument not being recognized by the davinci robot system. Engineering observed that one grip close cable was found to be broken at the distal idlers pulley. The idler pulley was able to spin freely; the instrument also exhibited some dents damages on the edge and the surface. The cable segment sticks out at wrist. Other cables at wrist were not damaged. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2925853 | 2900635 | 2797131 | 2955842201300087 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-07 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
420183 | 2012-12-19 | Injury | Required Intervention | The pt is a (b)(6) female admitted to (b)(6) medical center on (b)(6) 2012, for a scheduled surgery. The pt had been diagnosed with a retroperitoneal lymphadenopathy. The planned surgery was a robotic assisted peripancreatic lymph node excision. During the procedure to excise the lymph node the surgeon noticed that the cautery hook was malfunctioning. The malfunction device was removed. Upon removal of the cautery hook, it was noted that part of the plastic casing around the pulley of the hook was broken off. The surgical team immediately began extensive search in the area of the lymph node dissection, underneath the liver and around the porta hepatis, but did not see any evidence of the broken plastic casing. Several attempts were made to locate the missing fragment including creating a hand port in the upper midabdomen. A flat plate kub was taken. The team was unable to locate the missing fragment. The pt remained stable throughout the surgery and at the conclusion of the procedure was taken back to recovery in stable condition. |  | DA VINCI PERMANENT CAUTERY HOOK 8MM | ROBOTIC ARM JHOOK | 2940006 | 2913443 | 2809959 | 2913443 | 1 | NAY | User Facility | Unknown | RISK MANAGER | Initial | 2012-12-31 | 2013-01-07 | Yes | No | Health Professional | 420183 | No | No |  |  |  |  | 2013
42030901 | 2012-12-10 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure the cable of the mega suturecut needle driver instrument broke during the setup. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. Other cables at the wrist are not damaged. The instrument has 7 uses left. Additional findings revealed scratches on the distal pulley. The bottom pulley was in good condition. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2921664 | 2895844 | 2792360 | 2955842201300055 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-04 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42029602 | 2012-12-10 | Malfunction |  | It was reported that the wires of the large suturecut needle driver instrument were exposed at the distal end, which was identified in central processing. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one grip close cable was derailed and frayed at the distal idlers pulley. The idler pulley has difficulty spinning and damages are found on the pulley discs. The grip tips can open and close, but do not move intuitively. The instrument has 7 uses remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2921662 | 2895889 | 2792403 | 2955842201300054 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-04 | No | Yes | Health Professional | 42029602 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-03 | Malfunction |  | It was reported that prior to starting a da vinci s surgical procedure, the wires were loose at the tip of the mega suturecut needle driver instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the instrument was found with a frayed cable at the distal idler pulley. The frayed cable section is approximately 0. 05 in length. The idler pulley was found with indentation on the surface and the edge. The instrument has 3 uses remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2921663 | 2895890 | 2792405 | 2955842201300056 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-04 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
| 2012-12-10 | Malfunction |  | On (b)(6) 2012, nurse (b)(6) reported, during hysterectomy with lnd (4-5 hour procedure) the mcs tip cover was damaged. Fenestrated bipolar instrument was used with the cadiere forceps instrument in the 3rd arm (right side) generator esu setting were at 30/30 according to the additional information provided by nurse (b)(6), the monopolar curved scissors instrument burned a hole through mcs tip cover. | Intuitive surgical contacted the initial reporter of this complaint to request more information concerning the reported event; however, as of the date of this report no response has been received. The mcs tip cover has not been returned for evaluation, the root cause of the damage to the tip cover prior to the reported arcing event is unknown. A follow-up mdr will be submitted if additional information is received. | 8 MM MCS TIP COVER ACCESSORY | ENDOSCOPIC ELECTROSURGICAL INS | 2922382 | 2897172 | 2793688 | 2955842201300057 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-10 | 2013-01-04 | No | Yes | Health Professional |  | No | Yes | 2012-12-10 | No | Yes | Unkown | 2013
IS2000 A51P8 | 2012-12-06 | Other | Other | It was reported that during a da vinci s hysterectomy surgical procedure, the camera was not latched in the endoscopic camera manipulator (ecm) properly when the doctor let go the camera; the scope fell forward and broke the camera cannula mount accessory. The surgeon decided to convert to traditional open surgical techniques to complete the planned procedure. Initial reporter said no patient harm, adverse outcome or injury occurred. | The investigation conducted by the isi field service engineer confirmed that the camera cannula mount accessory had broken. The ecm is the camera arm located on the patient side cart which provides the sterile interface for the 3d endoscope. The camera cannula mount accessory is attached to the ecm and designed to hold the cannula securely in place, outside of the patient cavity. The system was repaired by replacing the affected ecm cannula mount accessory. Isi contacted clinical nurse manager (b)(6), he indicated that the ecm cannula mount broke and they could not hold the endoscope camera in place, the account did not have a cannula mount backup to continue the case with the davinci robot system; consequently, they had to convert the procedure to a traditional open surgery. The procedure was completed successfully. He was not aware of any patient complications. As of (b)(6) 2013, there have been no reported recurrences of the issue at this hospital. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2922708 | 2897376 | 2793892 | 2955842201300058 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-06 | 2013-01-04 | Yes | Yes | Health Professional | IS2000 A51P8 | No | No | 2012-12-06 | No | No | Reuse | 2013
42018305 | 2012-11-02 | Malfunction |  | It was reported during a da vinci si whipple procedure that part of yellow brown tip casing on the permanent cautery hook instrument fell inside the patient and was retrieved. No patient harm, adverse outcome or injury was reported. | Intuitive surgical contacted the initial reporter of this complaint. It was indicated that the fragment was retrieved and there was no harm or injury to the patient. The procedure was completed as planned and the patient has recovered as expected. Instrument was returned and evaluated. Engineering observed the one distal clevis ear broken off at its base. The clevis ear on the side opposite of the conductor wire is broken off, resulting in dislodgment of the yaw pulley. The broken piece was not returned, but is roughly. 285 x. 220 in size. The evidence not conclusive, but breakage is likely due to overloading the tip. Additionally, tube damage was observed though not reported by the site. The distal end of the main tube has various scratch marks with light material removal. The scratches are. 105 -. 215 in length and not aligned with the tube axis. Engineering concluded that damage was likely due to mishandling/misuse. The instruments and accessories user manual specifically states: general precautions and warnings: handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2922748 | 2897390 | 2793906 | 2955842201300059 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-12 | 2013-01-04 | No | Yes | Health Professional | 42018305 | Device Returned To Manufacturer | No | 2012-12-12 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-12-06 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure, the orange part cracked on the monopolar curved scissors instrument. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument and/or accessory have not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation), or if additional information is received. | HOT SHEARS TM MONOPOLARCURVED SCISSORS | ENDOSCOPIC ELECTROSURGICAL INS | 2922902 | 2897558 | 2794074 | 2955842201300060 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-06 | 2013-01-04 | No | Yes | Health Professional | 42017910 | No | No | 2012-12-06 | No | No | Unkown | 2013
420183 | 2012-11-28 | Malfunction |  | It was reported that during a da vinci si prostatectomy procedure, smoke was observed coming from the wrist of the permanent cautery hook instrument and that the instrument was not functioning properly. During replacement of the instrument, an electrical burning odor was also noted. No patient harm, adverse outcome or injury was reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. On (b)(4) 2012, isi contacted (b)(6). (b)(6) indicated that the permanent cautery hook (pch) instrument was inspected prior to use and there was no damage to the instrument noted. (b)(6) indicated that a valley lab electrical surgical unit was used; however, (b)(6) indicated that she does not recall the mode and settings used during the surgical procedure. (b)(6) indicated that the instrument did not make contact with any other instrument during the surgical procedure and there was no excessive biodebris on the instrument. (b)(6) indicated that arcing from the pch instrument was not observed and that it is unknown if a video recording of the planned surgical procedure is available. (b)(6) indicated that no patient harm, adverse outcome or injury occurred and that she does not know if the patient has returned to the hospital due to experiencing any post surgical complications as a result of the reported event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2922988 | 2897648 | 2794164 | 2955842201300061 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-06 | 2013-01-04 | No | Yes | Health Professional | 420183 | No | No | 2012-12-06 | No | No | Unkown | 2013
42020505 | 2012-12-06 | Malfunction |  | It was reported that during cleaning and sterilization, the customer noted broken wires at the distal end of the fenestrated bipolar forceps instrument. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was still installed in clevis. The clevis did not exhibit any damage or wear marks. No other damage was found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2920651 | 2895103 | 2791619 | 2955842201300018 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-06 | 2013-01-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-06 | Device Not Retu | No | Reuse | 2013
42011704 | 2012-12-04 | Malfunction |  | It was reported that during reprocessing of the needle driver instrument, the black coating was observed to be peeling off. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the black tube insulation is damaged at the distal end. The insulation is gouged on one side and has a. 045 x. 045 piece missing and sections lifted up roughly 2. 3 above the snake wrist. There is also a section below the missing piece where the insulation is heavily scratched. Engineering concluded that the damage is likely due to mishandling/misuse. No other damage found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2920738 | 2895128 | 2791644 | 2955842201300020 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Nurse | Initial | 2012-12-04 | 2013-01-03 | No | Yes | Health Professional | 42011704 | Device Returned To Manufacturer | Yes | 2012-12-04 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-11-28 | Malfunction |  | It was reported that post op a da vinci si hysterectomy procedure, the cable on the pk dissecting forceps instrument broke. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable at the distal clevis hub is frayed. There was no damage found at the clevis. The frayed segment is approximately 0. 04 in length. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2920788 | 2895161 | 2791677 | 2955842201300022 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-04 | 2013-01-03 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-04 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-17 | Malfunction |  | Event Description: --/-- It was reported after a da vinci si hysterectomy procedure that the maryland bipolar forceps instrument had cut wires and fragments may have fallen into the patient. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Engineering observed the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- Intuitive surgical contacted the initial reporter of this complaint. It was indicated that no fragments fell into the patient and there was no harm or injury to the patient. The procedure was completed as planned and the patient has recovered as expected without any post-surgical complications. The instrument has not been returned for evaluation. The root cause of the customer reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | Reuse | 2920812 | 2895164 | 2791680 | 2955842201300021 | 1 | NAY | Manufacturer | User facilityOther | Other | InitialFollowup | 2013-01-11 | 2013-01-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No |  | 2013
42000606 | 2012-12-04 | Malfunction |  | It was reported that during reprocessing of the large needle driver instrument, the wire on the instrument observed to be broken. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable at the distal clevis hub is frayed. The frayed segment is approximately 0. 06 in length. The frayed section of the cable showed damage leading into the distal clevis hub. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | LARGE NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2920823 | 2895185 | 2790311 | 2955842201300019 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-04 | 2013-01-03 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2012-12-04 | Device Not Retu | No | Reuse | 2013
42019408 | 2012-12-04 | Malfunction |  | It was reported that prior to starting da vinci s surgical procedure, the cables at the end of the mega needle driver instrument were observed to be shedding. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a frayed grip cable. The frayed grip cable was located at the distal idler pulley. The frayed cable sections are in various lengths sticking out at the wrist. The idler pulley rim exhibited mechanical indentations. Functional testing of the instrument on an isi in-house is2000 system showed that the instrument had 10 lives, indicating that the instrument was never used. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2920969 | 2895301 | 2791817 | 2955842201300024 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-04 | 2013-01-03 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-12-04 | Device Not Retu | No | Initial | 2013
42031802 | 2012-11-29 | Malfunction |  | It was reported that post op a da vinci si surgical procedure, it was noted that the cable on the small grasping retractor instrument was broke. There were no missing or fallen pieces reported. | The instrument has not been returned for evaluation, therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | SMALL GRASPING RETRACTOR INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2920988 | 2895304 | 2791820 | 2955842201300023 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-04 | 2013-01-03 | No | Yes | Health Professional | 42031802 | No | No | 2012-12-04 | Device Not Retu | No | Reuse | 2013
42018406 | 2012-11-30 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, the tip on the permanent cautery spatula instrument crumble and broke loose, while the surgeon was performing lymph node dissection. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a section of the ceramic sleeve broken off. The sleeve is missing pieces at the distal end with an overall footprint measuring roughly. 090 x. 090. Sleeve exhibits various scratch marks below the broken section. Spatula also exhibits light scratch marks in the axial direction, indicating possible attempt at scraping off biodebris, which may have damaged the ceramic. Engineering concluded that the damage was likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. O do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PERMANENT CAUTERY SPATULA INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2921072 | 2895353 | 2791868 | 2955842201300025 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-05 | 2013-01-03 | No | Yes | Health Professional | 42018406 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-03 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, it was observed that the cables on the fenestrated bipolar forceps instrument were broken. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. The instrument also passed electrical continuity testing. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2921126 | 2895388 | 2791903 | 2955842201300027 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-05 | 2013-01-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-03 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, a wire at the distal end of the megasuturecut needle driver instrument was sticking out. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers. The idler pulley spins freely and was not damaged. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2921123 | 2895402 | 2791918 | 2955842201300026 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-05 | 2013-01-03 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42805312 | 2012-12-03 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, the cable on the hem-o-lok medium-large clip applier instrument was observed to be broken. | The instrument was returned and evaluated. Engineering was unable to confirm the customer reported failure mode as the instrument does not contain any drive cables. The instrument uses a push pull rod as the mechanism to open/close grips. The grips do not open/close properly when input is manually rotated. Engineering evaluation also found that the distal end of the main tube is broken at the interface with the tan colored clevis. The broken piece remains installed below the clevis surface. The tube is no longer one rigid piece due to the breakage, which causes grip and clevis assembly to move in and out when input disc is rotated. As a result, the end of the tube no longer provides a hardstop and grips do not open/close properly. Engineering concluded that the damage is likely due to mishandling/misuse. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperatively. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | HEMOLOK MEDIUM LARGE CLIP APPLIER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2921254 | 2895450 | 2791965 | 2955842201300029 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-06 | 2013-01-03 | No | Yes | Health Professional | 42805312 | Device Returned To Manufacturer | No | 2012-12-06 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-12-04 | Malfunction |  | It was reported that during a da vinci s hysterectomy procedure, the cable on the fenestrated bipolar forceps instrument is broken. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still remained in the clevis. Clevis does not exhibit excessive damage. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2921212 | 2895459 | 2791975 | 2955842201300028 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-06 | 2013-01-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-06 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-11-29 | Malfunction |  | It was reported that during a da vinci s prostatectomy procedure, a wire on the maryland bipolar forceps instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one conductor wire is broken at the yaw pulley exit. The wire is detached from its connection at the grip. The instrument failed electrical continuity testing. The yaw pulley shows no signs of arcing. Engineering evaluation also found that the yaw pulley is missing a. 220 x. 060 piece opposite the conductor break, allowing the grip to move within the pulley and have a larger range of motion in yaw. Engineering concluded that it is indeterminable how the piece of yaw pulley broke. No other damage was found. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2921287 | 2895525 | 2792041 | 2955842201300030 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-06 | 2013-01-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-06 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-07 | Malfunction |  | It was reported that during a davinci si hysterectomy procedure, the customer noted exposed cables at the tip of the mega suture cut needle instrument. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that both grip cables were found to be broken at the distal idlers. The idler pulleys spin freely and exhibited damage to the rim and face. The cable segments sticks out at wrist. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2921349 | 2895605 | 2792120 | 2955842201300031 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-07 | 2013-01-03 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-07 | Device Not Retu | No | Reuse | 2013
42000606 | 2012-12-06 | Malfunction |  | Event Description: --/-- It was reported that during a davinci si hysterectomy procedure, the customer noted a broken cable at the distal end of the large needle driver instrument. No patient harm, adverse outcome or injury was reported. | Manufacturer Narrative: --/-- The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable was found to be broken at the distal idlers. The idler pulley was able to spin freely and it was not damaged. The cable segment sticks out at wrist. Other cables at wrist were not damaged. Engineering also observed an additional finding, scratches and indentation marks were found on the edge and surface of the distal pulley. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. --/-- Additional Manufacturer Narrative: --/-- The instrument has not been returned for evaluation; therefore, the root cause of the customer reported failure mode cannot be determined. A follow-up mdr will be submitted if the instrument is returned (post engineering evaluation) or if additional information is received. | LARGE NEEDLE DRIVER | Reuse | 2921387 | 2895621 | 2792137 | 2955842201300032 | 1 | NAY | Manufacturer | User facility | Other | FollowupInitial | 2013-01-11 | 2013-01-03 | No | Yes | Health Professional | 42000606 | Device Returned To Manufacturer | No | 2013-01-11 | Device Not Retu | No |  | 2013
42020706 | 2012-12-04 | Malfunction |  | It was reported that during a davinci si prostatectomy procedure, while using the tenaculum forceps instrument, the customer noted 'broken cable (nothing fell into the patient, no patient harm)' no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch up cable was found to be broken at the distal clevis hub. The cable segment that contains the crimp was missing in the clevis. The clevis did not exhibit any damage or wear marks. No other damage found. The customer reported complaint does not itself constitute a reportable event; however, the reported malfunction if to reoccur could cause or contribute to an adverse event. | TENACULUM FORCEPS | ENDOSCOPIC INSTRUMENT | 2921450 | 2895675 | 2792190 | 2955842201300033 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-07 | 2013-01-03 | No | Yes | Health Professional | 42020706 | Device Returned To Manufacturer | No | 2012-12-07 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-06 | Malfunction |  | It was reported that during a da vinci hysterectomy procedure, one of the pulley cables on the megasuturecut needle driver instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that one grip close cable is broken at the distal idlers pulley. The idler pulley spins freely but has damages on the edge and on the surface. The cable segment sticks out at the wrist. The other cables at the wrist are not damaged. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2921575 | 2895803 | 2792318 | 2955842201300039 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-07 | 2013-01-03 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-07 | Device Not Retu | No | Reuse | 2013
IS2000 A51P7 | 2010-12-01 | Injury | Required InterventionHospitalizationDisability | On (b)(6) 2012, intuitive surgical received a legal summons alleging a patient was burned and the system tore tissues of the patient's rectum, causing a fistula which required subsequent hospitalizations and multiple surgical procedures. It is alleged that the patient remains injured today as a result. It is alleged that the patient is permanently and seriously injured. It is alleged that the patient has suffered severe physical harm and severe emotional distress. The date of the original procedure was (b)(6) 2010. | On (b)(4) 2013 isi contacted dr. (b)(6) the surgeon who had provided the initial patient care. On (b)(6) 2013, dr. (b)(6) informed he would seek legal advice prior to opening dialogue for this complaint. A review of system logs for the da vinci s system used in the case on the event date ((b)(6) 2010) showed no errors during the procedure. The root cause of the reported failure mode cannot be determined. If additional information is received, a follow-up mdr will be submitted to the fda. | DA VINCI S SURGICAL SYSTEM | ENDOSCOPIC INSTRUMENT CONTROL | 2921541 | 2895808 | 2792323 | 2955842201300040 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-12-04 | 2013-01-03 | Yes | Yes | Health Professional | IS2000 A51P7 | No | No | 2012-12-04 | No | No | Reuse | 2013
42030901 | 2012-11-26 | Malfunction |  | It was reported prior to starting a da vinci si procedure that a megasuturecut needle driver instrument had a broken wire. No patient harm, adverse outcome or injury was reported. | Customer initially reported broken wire. However for clarification, the nonconformance was a broken grip cable. Based on this, the complaint was confirmed. Idler pulley spins freely and exhibit damage to the face and rim. Cable segments sticks out at wrist. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGASUTURECUT NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2921542 | 2895809 | 2792324 | 2955842201300034 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-04 | 2013-01-03 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-04 | Device Not Retu | No | Unkown | 2013
42019408 | 2012-12-07 | Malfunction |  | It was reported that during a da vinci s myomectomy procedure, the wires on the mega needle driver instrument broke. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the instrument was returned with a frayed cable at distal idler pulley. The frayed cable section is approximately 0. 14 in length. The grip cable also was found derailed. The distal pulley was found with damages on the surface and the edge. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA NEEDLE DRIVER INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2921581 | 2895810 | 2792325 | 2955842201300041 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-07 | 2013-01-03 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-12-07 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-11-13 | Malfunction |  | It was reported prior to starting the da vinci si procedure the pk dissecting forceps instrument had a broken wire. | Complaint broken wire is confirmed. The pitch cable is broken at the distal clevis hub. Cable segment that contains the crimp is still installed in clevis. Clevis does not exhibit any damage or wear marks. Electrical continuity passed. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PK DISSECTING FORCEPS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2921582 | 2895811 | 2792326 | 2955842201300035 | 1 | NAY | Manufacturer |  | Other | Initial | 2012-12-04 | 2013-01-03 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-04 | Device Not Retu | No | Reuse | 2013
42018306 | 2012-09-04 | Malfunction |  | It was reported that during a da vinci s adrenalectomy procedure, the surgical staff reported seeing smoke coming from the permanent cautery hook instrument. The instrument was replaced with a instrument of the same type to complete the planned surgical procedure. No patient harm, adverse outcome or injury was reported. | Customer initially reported smoke. However for clarification, the nonconformance was charring. Charring is likely a sign of unintended arcing. Based on this, the complaint was confirmed. Charring and burnt marks were observed at the yaw pulley and the exit of the conductor wire. Electrical continuity failed. The customer reported complaint does not itself constitute a mdr reportable event; however the reported malfunction if to recur could cause or contribute to an adverse event. | PERMANENT CAUTERY HOOK INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2921583 | 2895812 | 2792327 | 2955842201300038 | 1 | NAY | Manufacturer | Health ProfessionalUser facility | Other | Initial | 2012-12-04 | 2013-01-03 | No | Yes | Health Professional | 42018306 | Device Returned To Manufacturer | No | 2012-12-04 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-06 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the wire on the prograsp forceps instrument was noted to be broken. There were no missing or fallen pieces reported. The planned surgical procedure was completed and no patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Per the customer reported complaint, engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is missing in the clevis. The clevis does not exhibit any damage or wear marks. The customer reported complaint does not itself constitute a mdr reportable event; however; the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP FORCEPS INSTRUMENT | ENDOSCOPIC INSTRUMENT | 2921592 | 2895819 | 2792334 | 2955842201300042 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-12-07 | 2013-01-03 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-07 | Device Not Retu | No | Reuse | 2013
42017910 | 2012-11-09 | Malfunction |  | It was reported that during a da vinci si cystocele repair procedure, the scissors on the monopolar curved scissors (mcs) instrument were dull. There were no missing or fallen pieces reported. | The instrument was returned and evaluated. Engineering confirmed the customer reported failure mode. The instrument was installed on an is3000 in-house system found that the scissors cut cleanly through (b)(4) latex. The blades are not damaged. The instrument passed electrical continuity testing. Engineering evaluation also found that the printing on the tube extension was scraped off. Engineering has concluded that the damage was likely due to mishandling. Late submission of this report was due to a reporting oversight by the responsible complaint handler. | MONOPOLAR CURVES SCISSORS INSTRUMENT | ENDOSCOPIC ELECTROSURGICAL INS | 2921546 | 2895830 | 2792345 | 2955842201300043 | 1 | NAY | Manufacturer | Other | Other | Initial | 2012-11-27 | 2013-01-03 | No | Yes | Health Professional | 42017910 | Device Returned To Manufacturer | No | 2012-11-27 | Device Not Retu | No | Reuse | 2013
42020505 | 2012-11-28 | Malfunction |  | It was noticed that prior to starting a da vinci si surgical procedure, the tip of the fenestrated bipolar forceps instrument was bent. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the grip, tip and teeth are perfectly in position. However, findings revealed a broken pitch cable and scratches on the main tube. The distal end of the main tube has various scratch marks with light material removal, suggesting it may have been caused by instrument collisions or rough handling. The pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity passed. The instrument has 2 uses remaining. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | FENESTRATED BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2921658 | 2895839 | 2792355 | 2955842201300050 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-06 | 2013-01-03 | No | Yes | Health Professional | 42020505 | Device Returned To Manufacturer | No | 2012-12-06 | Device Not Retu | No | Reuse | 2013
42000305 | 2012-12-07 | Malfunction |  | It was reported that during a da vinci s surgical procedure, the wires of the small clip applier instrument were noted to be broken. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. A visual inspection shows no damage to the drive cables at the distal end. Grips are able to open/close when input discs are manually rotated. Additional observation not reported by site is tube damage. The distal end of the main tube has scratch marks with light material removal. Scratch mark is. 210 in length and axially aligned with the tube axis. There were 65 fires left. No other damage was found. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. Based on this, no additional follow up is required. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | SMALL CLIP APPLIER | ENDOSCOPIC INSTRUMENT | 2921639 | 2895841 | 2792357 | 2955842201300052 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-10 | 2013-01-03 | No | Yes | Health Professional | 42000305 | Device Returned To Manufacturer | No | 2012-12-10 | Device Not Retu | No | Reuse | 2013
42022703 | 2012-12-03 | Malfunction |  | It was reported that during a da vinci si surgical procedure, the pk dissecting forceps instrument cable snapped while in use. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that one conductor wire is broken at the yaw pulley exit. The wire is detached from its connection at the grip. Electrical continuity failed. Yaw pulley shows a broken piece measured roughly about 0. 159 x 0. 039. Additional findings revealed scratches on the main tube. The distal end of main tube has various scratch marks with light material removal. The scratches are short in length and not aligned with the tube axis, suggesting it may have been caused by instrument collisions or rough handling. The instrument has 7 uses left. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. Based on this, no additional follow up is required. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. | PK DISSECTING FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2921617 | 2895862 | 2792377 | 2955842201300044 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-05 | 2013-01-03 | No | Yes | Health Professional | 42022703 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42030901 | 2012-12-05 | Malfunction |  | It was reported that prior to starting a da vinci si surgical procedure, cables at the end of the mega suturecut needle driver instrument were frayed. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed a broken grip cable. The broken grip cable was found at the distal idlers. The cable segment sticks out at the wrist. The idler pulley spins freely. The instrument has zero uses left. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MEGA SUTURECUT NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2921618 | 2895863 | 2792378 | 2955842201300045 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-05 | 2013-01-03 | No | Yes | Health Professional | 42030901 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-11-29 | No Answer Provided |  | It was reported that prior to starting a da vinci s surgical procedure, a loose wire was noted on the maryland bipolar forceps instrument. Reportedly, the instrument was also not reading lives when plugged in. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the pitch cable is broken at the distal clevis hub. The cable segment that contains the crimp is still installed in the clevis. The clevis does not exhibit any damage or wear marks. Electrical continuity failed. There was a broken conductor wire at the yaw pulley exit. The wire is detached from its connection at the grip. The yaw pulley shows no signs of arcing. The instrument has 5 uses remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2921619 | 2895864 | 2792379 | 2955842201300046 | 1 | NAY | Manufacturer |  | Other | Initial | 2012-12-05 | 2013-01-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-05 | Device Not Retu | No | Reuse | 2013
42009308 | 2012-12-06 | Malfunction |  | It was reported that fibers on the prograsp forceps instrument were coming out during central processing. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed scratches on the main tube. Complaint of fibers coming out could not be confirmed. The distal end of the main tube has various scratch marks with light material removal, suggesting it may have been caused by instrument collisions or rough handling. The instrument has 6 uses left. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | PROGRASP TM FORCEPS | ENDOSCOPIC INSTRUMENT | 2921624 | 2895871 | 2792386 | 2955842201300047 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-06 | 2013-01-03 | No | Yes | Health Professional | 42009308 | Device Returned To Manufacturer | No | 2012-12-06 | Device Not Retu | No | Reuse | 2013
42019408 | 2012-10-24 | Malfunction |  | It was reported that during a da vinci si myomectomy procedure the instrument failed midway. The jaws would not close, to a point where force had to be used to extract the instrument. After surgery, the instruments were inspected and tears in the cable were noted on the mega needle driver. No system error(s) had occurred. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that the instrument was found with a frayed cable at the distal idler pulley. Frayed cable section is approximately 0. 14 in length. No excessive damage was found on the pulley. The instrument has 8 uses remaining. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. Late submission of this medwatch report is due to a reporting oversight by the complaint specialist. | MEGA NEEDLE DRIVER | ENDOSCOPIC INSTRUMENT | 2921632 | 2895872 | 2792387 | 2955842201300048 | 1 | NAY | Manufacturer | ForeignUser facility | Other | Initial | 2012-11-20 | 2013-01-03 | No | Yes | Health Professional | 42019408 | Device Returned To Manufacturer | No | 2012-11-20 | Device Not Retu | No | Reuse | 2013
42017207 | 2012-12-06 | Malfunction |  | It was reported that during a da vinci si hysterectomy procedure, broken wires at the distal end of the maryland bipolar forceps instrument were noted. The intended procedure was completed successfully. No patient harm, adverse outcome or injury was reported. | The instrument was returned and evaluated. Engineering observed that both pitch cables are broken at the distal clevis hub. The cable segments that contains the crimp are still installed in the clevis. The clevis does not exhibit any damage or wear marks. Other cables at wrist are not damaged. Additional observation not reported by site is tube damage. The distal end of the main tube has various scratch marks with light material removal. Scratches are. 100 -. 180 in length and not aligned with the tube axis. Evidence was not conclusive, but damage may have been caused by mishandling/misuse. There was one use left. No other damage was found. The instruments and accessories user manual specifically states: general precautions and warnings - handle instruments with care. Avoid mechanical shock or stress that can cause damage to the instruments. Do not use an instrument to clean debris from another instrument intraoperative. This may result in damage to the instruments or other unintended consequences, such as disconnection of the instrument tip. Per the customer, at the time of the reported event, nothing fell into a patient during a surgical procedure. Based on this, no additional follow up is required. The customer reported complaint does not itself constitute a mdr reportable event; however, the reported malfunction if to recur could cause or contribute to an adverse event. | MARYLAND BIPOLAR FORCEPS | ENDOSCOPIC ELECTROSURGICAL INS | 2921633 | 2895874 | 2792389 | 2955842201300049 | 1 | NAY | Manufacturer | User facility | Other | Initial | 2012-12-06 | 2013-01-03 | No | Yes | Health Professional | 42017207 | Device Returned To Manufacturer | No | 2012-12-06 | Device Not Retu | No | Reuse | 2013